## Slide 1: Radiation Protection in Dental Radiology
Protection of Workers and Public in Dental Radiology
L12
Training material developed by the International Atomic Energy Agency
in collaboration with: World Health Organization, FDI World Dental Federation, International Association of Dento-Maxillofacial Radiology, International Organization for Medical Physics, and Image Gently Alliance

## Slide 2: Educational Objectives
Understands the scatter distribution for different radiographic modalities
Identify the safest position of the operator and public relative to the patient and x-ray tube
Understand the need for adequate distance and shielding, and be able to apply these principles in clinical practice
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 3: Overview
Protection from what?
Dose limits / dose constraints
Dose from scatter in dental radiography
Principles of Protection
  Distance
  Time
  Shielding
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 4: Overview
Protection from what?
Dose limits / dose constraints
Dose from scatter in dental radiography
Principles of Protection
  Distance
  Time
  Shielding
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 5: Protection from what?
Protection from (decreasing order of importance):
  Scattered (S) radiation (source = patient)
  Primary (P) radiation (the X ray beam)
    Usually fully captured by image receptor
    Intra-oral radiography: slight discrepancy between beam size and detector size
  Leakage (L) radiation (from X ray tube)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
X
L
S
P

## Slide 6: Overview
Protection from what?
Dose limits / dose constraints
Dose from scatter in dental radiography
Principles of Protection
  Distance
  Time
  Shielding
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 7: Dose limits for workers and public
One of the three principles of radiation protection: dose limitation
  International Basic Safety Standards, Requirement 12 (adopted): “The government or the regulatory body shall establish dose limits for occupational exposure and public exposure, and registrants and licensees shall apply these limits.“
Dose limits established (& periodically revised) by ICRP (and subsequently incorporated into BSS and local legislation)
  See next slide
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 8: Dose limits for workers and public
Type of limit | Occupational | Public
Annual effective dose | 20 mSv 2,3 | 1 mSv4
Annual equivalent dose to: |  | 
Eye lens | 20 mSv2 | 15 mSv
Skin1 | 500 mSv | 50 mSv
Hands and feet | 500 mSv | -
1Averaged over 1 cm2, regardless of the area exposed.
2Averaged over 5 consecutive years (100 mSv in 5 years), with no single year exceeding 50 mSv.
3Additional restrictions apply for pregnant women.
4Similar to the occupational dose, a higher annual dose could be allowed in a single year in special circumstances, providing that the average over 5 years does not exceed 1 mSv/year.
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
Dose limits for occupational and public exposure, according to the current ICRP recommendations and International Basic Safety Standards (2014)

## Slide 9: Personnel monitoring
The need for dose monitoring of workers handling dental X-ray equipment depends on:
  Local legislation
  Type and workload of equipment
  Pregnancy of workers
Different types of badges for monitoring available (film, TLD, OSL, MOSFET, …)
Monitoring should be performed by a certified personnel dosimetry lab
Corrective actions should be performed if any personnel dose approaches dose limits
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 10: Dose limits for workers and public
Apart from dose limits, dose constraints can be applied
  Set separately for each radiation source
  Apply to occupational and public exposure
Not dose limits but investigation levels: exceeding a dose constraint does not represent non-compliance with regulatory requirements, but it could result in follow-up actions.
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 11: Overview
Protection from what?
Dose limits / dose constraints
Dose from scatter in dental radiography
Principles of Protection
  Distance
  Time
  Shielding
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 12: Dose from scatter in dental radiography
Intraoral radiography
  Office of Radiation Safety, Ministry of Health, New Zealand: 0.08-0.5 µGy at 1 m from patient
  (see figure on slide 21)
  Scatter for intra-oral radiography strongly depends on location (see further)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 13: Dose from scatter in dental radiography
Intraoral radiography
  Portable intra-oral machines: local absorbed dose per exposure for operator 0.04-0.6 µGy (Danforth et al. 2009, Hermsen et al. 2008, McGiff et al. 2012, Pittayapat et al. 2010). Use of backscatter shield can reduce local doses 8-fold (McGiff et al. 2012)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 14: Dose from scatter in dental radiography
Panoramic radiography
  Gijbels et al. (2005): 0.04-0.53 µGy at 1m
CBCT
  SEDENTEXCT (2010): maximum scatter at 1m: 4.1-46.8 µGy (median for 10 CBCT models: 6.8 µGy), average scatter 2-8.1 µGy at 1m
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 15: Overview
Protection from what?
Dose limits / dose constraints
Dose from scatter in dental radiography
Principles of Protection
  Distance
  Time
  Shielding
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 16: Protection of Workers and Public
Basic protection principles:
  Increase distance to radiation source
  Apply shielding (room & personal)
  Limit the time spent in the vicinity of a radiation source
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 17: Principles of Protection
Distance
  Inverse square law: beam intensity (I) reduces with the square of the distance (D) to the source
3*D
Source
D
2*D
I
I/4
I/9
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 18: Principles of Protection
Distance
  Intraoral, panoramic, cephalometric radiography:
  Distance >1.5 m → annual dose to worker <1 mSv
  (EC RP 136, 2004)
    >2 m distance from the patient often recommended
    If shielding is used: distance can be shorter
  CBCT: scattered radiation much higher; must be installed in a ‘protected enclosure’ (EC RP 172, 2012)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 19: Principles of Protection
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
Apart from distance, location also matters!
  Scatter not homogenously distributed
  Diagnostic X-ray energy: tendency of scattering towards the forward (toward image receptor) direction
  But: scatter is absorbed by patient as well!
TOWARDS IMAGE RECEPTOR
TOWARD
X RAY TUBE
DISTRIBUTION OF X-RAY SCATTER (60 keV)
0°
90°
180°
270°
Relative amount of Compton scatter at each angle)
R. Pauwels
PATIENT

## Slide 20: Principles of Protection
Scatter doses at 1 m (in µSv) for an intra-oral exposure made at 60 kV, 0.7 s, 200 mm focus-to-skin distance
Reproduced with permission by the Office Of Radiation Safety. Ministry of Health, New Zealand
Highest dose (due to backscatter)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
Lowest dose: most of forward scatter is absorbed by patient

## Slide 21: Principles of Protection
Dose rate (µSv/h) at 2 m from the patient (average from 3 intra-oral units)
Based on data from Tabakov et al. (2000)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
Lowest dose: directly behind X ray tube or behind patient
X ray tube
366
373
344
62
212
102
144
Highest dose in front of patient (note: dose opposite to tube not measured)

## Slide 22: Principles of Protection
For intra-oral exposures (general)
  Distance and shielding: more important than positioning
  Avoid area opposite to X ray tube (high amount of scatter + possible primary radiation
  Avoid area behind X ray tube (tube itself provides shielding from backscatter, but surrounding area is prone to high amounts of scatter)
  When picking a location, keep as much “patient tissue” between yourself and the source of the scatter (i.e. the exposed region in the patient)
    At right angle to tube
    Patient acts as shielding
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 23: Principles of Protection
For intra-oral exposures (specific)
  Tube at side of patient: stand behind patient
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 24: Principles of Protection
For intra-oral exposures (specific)
  Tube diagonally toward patient: stand behind patient (90-135° to X ray beam)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 25: Principles of Protection
For intra-oral exposures (specific)
  Tube in front of patient: stand at either side of patient (90 or 270° to X ray beam)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 26: Principles of Protection
Panoramic radiography: little effect of position (Gijbels et al. 2005; anterior dose not measured)
CBCT: depends on geometric factors (SEDENTEXCT, 2010)
  Sitting/standing devices: scatter distribution +/- homogeneous for 360° rotations, lowest at ‘tube side’ for 180° rotations
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
360° rotation
180° rotation
(X ray tube at back of patient)

## Slide 27: Principles of Protection
CBCT: depends on geometric factors
(SEDENTEXCT, 2010)
  Supine devices: lowest scatter in the direction along the patient’s axis
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 28: Principles of Protection
CBCT
  Absorption of scatter by the machine housing and patient can affect scatter distribution
  Scatter should be measured case-by-case; use of room shielding should always be preferred over picking the location with lowest scatter for CBCT equipment!
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 29: Principles of Protection
Variable patient (ergo: X-ray beam) orientation & machine housing in CBCT
R. Pauwels
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
Sitting
Standing
Supine

## Slide 30: Principles of Protection
Shielding
  Room shielding & personal shielding
  Also protects:
    Patients (when not examined)
    Visitors, public
    Nearby people (e.g. office next to dentist)
  Material & thickness of shielding determines fraction of blocked radiation
  Should be tailored to national requirements, expert advise is required (see IAEA BSS)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 31: Principles of Protection
Shielding materials
  Ok to use concrete or other conventional building materials (lead added in walls when needed; leaded glass can be used in windows)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 32: Principles of Protection
Shielding materials
  Mobile or non-mobile shielding screens
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology
Courtesy of RAY-BAR

## Slide 33: Principles of Protection
Personal shielding
  Lead aprons
    0.25-0.5 mmPb
    Subject to wear (cracks)
      Check for cracks using 2D
      radiograph of the apron
  Thyroid shields
  Leaded glasses/gloves/caps
    Not used in a dental environment
  Image receptor should not be held by worker
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 34: Principles of Protection
Time
  Cfr. Patient dose: dose ~ exposure time
  Exposure time should be adapted to patient (ALARA principle), not to the worker!
  But: time of worker in surrounding to radiation source can be limited e.g. rotation schedule
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 35: Room Design
Special room design considerations for CBCT equipment
Design (i.e. shielding) of room depends on:
  X-ray equipment
    kV, mAs, …; workload; direction(s) of beam
  Organization of room
    Location of operator, location of public
    (e.g. surrounding rooms)
Amount of scattered radiation needs to be measured or estimated to determine shielding requirements
  Case-by-case basis, by qualified expert
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 36: Room Design
Locations to measure scatter (marked as ‘X’)
Example of dental CBCT room
Corridor
Treatment room
Waiting room
CBCT
Adjacent space (e.g. office, living area)
Operator
position
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 37: Room Design
Shielding in CBCT (EC RP 172, 2012)
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 38: References
Danforth RA et al. (2009) Operator exposure to scatter radiation from a portable hand-held dental radiation emitting device (Aribex NOMAD) while making 915 intraoral dental radiographs. J Forensic Sci. 54:415-21.
EC, European Commission (1999) European guidance on diagnostic reference levels (DRLs) for medical exposures. Radiation Protection 109. European Commission, Luxembourg. [https://ec.europa.eu/energy/sites/ener/files/documents/109_en.pdf]
EC, European Commission (2004) European guidelines on radiation protection in dental radiology - The safe use of radiographs in dental practice, radiation protection publication 136 European Commission, Luxembourg. [https://ec.europa.eu/energy/sites/ener/files/documents/136.pdf]
ICRP, International Commission on Radiological Protection (2007) The 2007 Recommendations of the International Commission on Radiological Protection. ICRP publication 103. Ann ICRP 37:1-332
ICRP, International Commission on Radiological Protection (2012) ICRP statement on tissue reactions and early and late effects of radiation in normal tissues and organs - threshold doses for tissue reactions in a radiation protection context. ICRP publication 118. Ann ICRP 41:1-322
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology

## Slide 39: References
Gijbels F et al. (2005) Dosimetry of digital panoramic imaging. Part II: Occupational exposure. Dentomaxillofac Radiol. 34:150-3.
Hermsen KP et al. (2008) Radiation safety for the NOMAD portable X-ray system in a temporary morgue setting. J Forensic Sci. 53:917-21.
McGiff TJ et al. (2012) Maintaining radiation exposures as low as reasonably achievable (ALARA) for dental personnel operating portable hand-held x-ray equipment. Health Phys. 103:S179-85.
Pittayapat P et al. (2010) Image quality assessment and medical physics evaluation of different portable dental X-ray units. Forensic Sci Int. 201:112-7.
SEDENTEXCT. Deliberable D2.2 - Completion of scatter dose measurements around CBCTs and recommendations for protective measures and positioning of CBCT units in dental offices. [http://www.sedentexct.eu/system/files/SEDENTEXCT%20D2.2%20v1%20FINAL_0.pdf]
Tabakov et al. (2000) Instantaneous dose rate of scatter radiation in dental radiography. Phys Med16:27-30
Radiation Protection in Dental Radiology L12 Protection of Workers and Public in Dental Radiology