## Slide 1: Radiation Protection in Dental Radiology
Fundamentals of Intraoral Radiography
L05
Training material developed by the International Atomic Energy Agency
in collaboration with: World Health Organization, FDI World Dental Federation, International Association of Dento-Maxillofacial Radiology, International Organization for Medical Physics, and Image Gently Alliance

## Slide 2: Educational Objectives
Be familiar with the different types of intraoral radiography
Understand the difference between (film-based and digital) image receptors used in intraoral radiography
Recognize and avoid faulty radiographs (position, under/overexposure, film handling and development, etc.)
Know when and how to use handheld intraoral radiography machines
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 3: What is intraoral radiography (IOR)?
IOR: a group name for different types of dental exposures, in which the image receptor is placed in the patient’s mouth:
  Periapical radiography
  Bitewing radiography
  Occlusal radiography
All IOR make use of the same X-ray tube, with different receptor sizes and positioning devices being used depending on the type of IOR
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 4: Overview
Intraoral X ray tubes
Intraoral image receptors
Types of intraoral projections
Causes of faulty intraoral radiographs
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 5: Overview
Intraoral X ray tubes
Intraoral image receptors
Types of intraoral projections
Causes of faulty intraoral radiographs
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 6: Intraoral X ray tubes
Typically wall-mounted (portable: see further) units with an adjustable arm to allow for in-chair exposure
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner

## Slide 7: Circular/rectangular collimation
As image receptors are rectangular, ideally, the beam should be collimated to a rectangular shape to avoid unnecessary exposure to the patient
  Large dose reduction possible through the use of rectangular collimation (Ludlow et al. 2008, Johnson et al. 2014)
Circular collimation reduces the risk of misalignment (‘cone cut’), but only has a marginal effect on the reject frequency (Parrot & Ng 2011)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 8: Technical specifications of IOR tubes
X ray tube potential:
  At least 60 kV
X ray tube filtration
  1.5 mm aluminium-equivalent if up to 70 kV
  2.5 mm aluminium-equivalent (of which 1.5 mm permanent) for units over 70 kV
X ray beam dimensions
  ≤4x5 cm at collimator end if rectangular (≤6 cm diameter if cylindrical)
  Rectangular collimation is recommended
Minimum focus-to-skin distance
  200 mm
See also: IEC 60601-2-65
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 9: Overview
Intraoral X ray tubes
Intraoral image receptors
Types of intraoral projections
Causes of faulty intraoral radiographs
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 10: Image receptors in IOR
Analogue (film)
Digital receptors
  Phosphor plate (PSP)
  Rigid (CCD, CMOS)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner
(top left)
R. Schulze
(others)
See L04 for more details

## Slide 11: Image receptors used in IOR: film
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Film is wrapped in paper/plastic for protection against physical damage & moisture
Lead foil used as backing
  Protection of film from backscatter
  Reduce patient exposure
  Embossed with pattern: if visible on image, indicates backward placement (i.e. R-L is reversed)
K. Horner

## Slide 12: Image receptors used in IOR: film
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Different sizes of film used in IOR
  Periapical:
    Size 0 (small children): 22x35 mm
    Size 1 (anterior teeth): 24x40 mm
    Size 2 (adults): 30.5x40.5 mm
  Bitewing:
    Size 2 for adults, size 1 (0) for small children
    Size 3 (extended bitewing, covers more teeth): 27x54mm
  Occlusal:
    Size 4: 57x76mm

## Slide 13: Image receptors used in IOR: film
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
0
1
2
3
4

## Slide 14: Overview
Intraoral X ray tubes
Intraoral image receptors
Types of intraoral projections
Causes of faulty intraoral radiographs
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 15: Types of intra-oral projections
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Periapical
Bitewing
Occlusal

## Slide 16: Periapical radiography
Aimed at visualizing the root (apex) of the tooth and surrounding structures
Two distinct techniques for placing the film are available:
  Paralleling technique
  Bisecting-angle technique
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 17: Periapical radiography
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
LONG AXIS OF TOOTH
IMAGE
RECEPTOR
X RAY TUBE
PALATE
  Paralleling technique: image receptor is placed parallel to the long axis of the tooth, beam is perpendicular to receptor
LONG AXIS OF TOOTH
IMAGE
RECEPTOR
X RAY TUBE
BISECTOR
  Bisecting-angle technique: image receptor is placed at an angle to the tooth, beam is perpendicular to the plane bisecting the tooth and the receptor

## Slide 18: Periapical radiography
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
  For solid state receptors, a parallel position to the tooth‘s main axis in the maxilla will mostly not be feasible (Roeder et al. 2011); a 'rectangular technique' is a good compromise. The image receptor by means of an aiming device is oriented perpendicular to the central x-ray, whilst no specific requirements apply to the angle between tooth and receptor.
R. Schulze

## Slide 19: Periapical radiography
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
P.Sinpitaksakul
  Series of 6 anterior and 8 posterior periapical radiographs (note different orientation of receptor)

## Slide 20: Bitewing radiography
Aimed at visualizing the crown of the tooth and interproximal surfaces
Orientation of receptor: cfr. paralleling technique
  With slight beam angle e.g. 5-10° (coinciding with occlusal plane angle)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
X RAY TUBE

## Slide 21: Bitewing radiography
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
P.Sinpitaksakul

## Slide 22: Holding / positioning devices
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Periapical
Bitewing
P.Sinpitaksakul
P.Sinpitaksakul

## Slide 23: Occlusal radiography
To visualize floor of mouth or palate
Image receptor in occlusal plane (between upper and lower teeth), tube at steep upward or downward angle
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
65°
90°

## Slide 24: Occlusal radiography
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
P.Sinpitaksakul

## Slide 25: Overview
Intraoral X ray tubes
Intraoral image receptors
Types of intraoral projections
Causes of faulty intraoral radiographs
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 26: Faulty radiographs
There are different possible image quality aberrations in IOR, each of which with a variety of causes (White & Pharoah, 2014)
In IOR, the following issues may cause faulty radiographs:
  Film development issues
  Over- or underexposure
  Image receptor artefacts (e.g. due to damage)
  Improper film or X ray tube positioning
  Other
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 27: Faulty radiographs (film)
Film too bright:
Processing errors
  Underdevelopment (too low temperature or time)
  Depleted/diluted/contaminated developer
  Excessive fixation
Underexposure
  kV, mA and/or exposure time too low
  Source-receptor distance too long
  Film packet placed inside out
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 28: Faulty radiographs (film)
Film too dark:
Processing errors
  Overdevelopment (too high temperature/time)
  Developer concentration too high
  Time in fixer too short
  Exposure to light (any type of IOR)
  Improper storage (e.g. no shielding, temperature too high)
Overexposure
  kV, mA and/or exposure time too high
  Source-receptor distance too short
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 29: Faulty radiographs (film)
Poor image contrast:
Can be caused by both under- and over-development and -exposure (for the latter, esp. too high kVp)
Film fogging (see further)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Simulated image quality aberrations; original image P.Sinpitaksakul

## Slide 30: Faulty radiographs (film)
Fogging:
Exposure to light (incl. improper safelight)
Overdevelopment
Contaminated solutions (e.g. developer ↔ fixer)
Improper storage
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner

## Slide 31: Faulty radiographs (film)
Local darkening:
Fingerprints
Wrapping paper stuck to film
Contact with tank/other film during fixation
Contaminated by developer before processing
Excessive bending
Static discharge to film before processing
Excessive roller pressure (automatic processing)
Dirty rollers (automatic processing)
LED from developer (or other local lighting)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 32: Faulty radiographs (film)
Local darkening:
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner

## Slide 33: Faulty radiographs (film)
Local brightening:
Contamination with fixer before processing
Contact with tank/other film during development
Excessive bending after development
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner

## Slide 34: Faulty radiographs (film)
Staining:
Depleted/contaminated developer/fixer
Insufficient washing
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner

## Slide 35: Faulty radiographs (film)
Partial image:
Film not entirely immersed in developer
Also: misalignment of X-ray beam and image receptor (see further, ‘mispositioning’)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 36: Faulty radiographs (film)
Emulsion peel:
Abrasion of film during processing
Excessive washing time
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 37: Faulty radiographs (digital)
High image noise:
Digital imaging: overall brightness not affected by exposure (↔ film); underexposure will manifest itself as an increase in image noise
Excessive exposure to light can also cause a noisy appearance
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Underexposed
Properly exposed
￩￪ K. Horner
Exposed to light (PSP)

## Slide 38: Faulty radiographs (digital)
Receptor artefacts:
Scratched surface (or other damage) of PSP
Excessive bending can cause permanent damage
Disinfectant solution
Surface contamination (PSP)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner

## Slide 39: Faulty radiographs (digital)
Receptor artefacts:
Damaged CCD/CMOS sensor (e.g. dropped) or cable
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
R. Schulze

## Slide 40: Faulty radiographs (film & digital)
Partial exposure to light
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
K. Horner

## Slide 41: Faulty radiographs (film & digital)
Blurring:
Movement of patient or X ray tube during exposure
Double exposure (or: incomplete erasure of PSP )
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
www.dentalcare.com
Permission granted by Crest + Oral-B
K. Horner

## Slide 42: Faulty radiographs (film & PSP)
Distortion:
Moderate/severe bending film/PSP during placement
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
www.dentalcare.com
Permission granted by Crest + Oral-B

## Slide 43: Faulty radiographs: mispositioning
In IOR, proper alignment between X-ray beam and image receptor is the responsibility of the operator
Proper use of receptor holders and beam aiming devices should avoid excessive reject frequencies
Even if alignment is OK, misplacement of the receptor holder/sensor or beam can result in missing or poorly visualized areas of diagnostic interest, or elongation/foreshortening
Bending: see ‘distortion’ and ‘local brightening’
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 44: Faulty radiographs: mispositioning
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
LONG AXIS OF TOOTH
IMAGE
RECEPTOR
X RAY TUBE
Correct position of tube and image receptor: relation between vertical and horizontal dimensions in radiograph correspond with reality

## Slide 45: Faulty radiographs: mispositioning
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Elongation due to improper position of receptor and/or X ray beam
www.dentalcare.com
Permission granted by Crest + Oral-B

## Slide 46: Faulty radiographs: mispositioning
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Foreshortening due to improper position of receptor and/or X ray beam
www.dentalcare.com
Permission granted
by Crest + Oral-B

## Slide 47: Faulty radiographs: mispositioning
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Poor horizontal alignment (overlap of interproximal surfaces)
Proper horizontal alignment
Poor horizontal alignment
CENTRAL RAY
K. Horner

## Slide 48: Faulty radiographs: mispositioning
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Poor alignment between X ray beam and image receptor (‘cone cutting’)
Receptor
X ray beam
K. Horner

## Slide 49: Faulty radiographs: objects
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Partial denture
Nose ring
Thyroid collar
www.dentalcare.com
Permission granted by Crest + Oral-B

## Slide 50: Faulty radiographs: image processing
Improper use of display settings and filters may lead to misdiagnosis
Exporting images: low bit depth + JPEG compression can deteriorate image quality
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Original image

## Slide 51: Portable (handheld) intra-oral radiography
Conventional IOR: using mounted X ray source, operated from distance
Handheld IOR introduced for applications requiring mobility
Variety of handheld IOR equipment available
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Pittayapat et al. (2012),
with author’s permission

## Slide 52: Portable (handheld) intra-oral radiography
The image quality of a handheld IOR machine (if following good practice; see further) can be the same as that of a mounted unit
However, this does not justify its use in a setting in which mounted units can be used instead!
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography
Pittayapat et al. (2010), reproduced with author’s permission
Portable
Mounted

## Slide 53: Portable (handheld) intra-oral radiography
Appropriate situations for using handheld IOR
(Berkhout et al. 2015)
  Patients under general anesthesia during surgery (if not mounted unit in room) (first choice: semi-mobile device)
  Immobile patients in emergency rooms, surgical suites, patient rooms etc. (first choice: semi-mobile device)
  Nursing homes (and similar): consider whether follow-up treatment can be done on-site; if not, perform imaging at healthcare facility
  Detention centers (physical confinement)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 54: Portable (handheld) intra-oral radiography
Appropriate situations for using handheld IOR
(Berkhout et al. 2015) (cont.)
  Remote areas (e.g. military operations)
  Rural/isolated areas (e.g. developing countries)
  Forensic (e.g. mass disasters)
For (iii.) to (vii.), risk assessment needed to limit occupational/public exposure
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 55: Portable (handheld) intra-oral radiography
Good practice recommendations (Berkhout et al. 2015)
  Complete documentation
  Device should at least have mandatory certification
  (e.g. CE, FDA)
  Device should support receptors holders & beam aiming devices
  Image receptor: E/F-speed film or digital
  Exposure time: can be longer than mounted units (due to low mA for portable), but never >1s (motion artefacts)
  Battery: should not compromise tube output
  Weight: use tripod or other support if available
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 56: Portable (handheld) intra-oral radiography
Good practice recommendations (Berkhout et al. 2015)
  Always use backscatter shielding (≥0.25 mm lead-equivalent, ≥15.2cm⍉, ≤1cm from end of positioning device)
  Licensing/authorization: following national regulations
  Safe storage
  Other national regulations (can vary considerably between countries)
  Patient protection: same as mounted IOR
  Operator protection: risk analysis, monitoring if needed
  Public protection: risk assessment incl. distance/shielding
  Image quality requirements: same as mounted IOR
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 57: Portable (handheld) intra-oral radiography
Good practice recommendations (Berkhout et al. 2015)
  QA program required (incl. reject analysis)
  Other specifications cfr. EC RP 162 (2012) e.g. focus-skin distance ≥200mm, field size ≤40x50mm
  Medical physics expert involvement in acceptance/routine testing
  Proof of training by user
See also: HERCA position statement on use of handheld portable dental x-ray equipment (2014)
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 58: References
Berkhout WE et al. (2015): Justification and good practice in using handheld portable dental X-ray equipment: a position paper prepared by the European Academy of DentoMaxilloFacial Radiology (EADMFR). Dentomaxillofac Radiol. 44:20140343.
Dentalcare: CE course: Intraoral radiography: Principles, Techniques and Error Correction. https://www.dentalcare.com/en-us/professional-education/ce-courses/ce137/toc
EC, European Commission (2012): Radiation Protection no 162, Criteria for acceptability of medical radiological equipment used in diagnostic radiology, nuclear medicine and radiotherapy. http://ec.europa.eu/energy/sites/ener/files/documents/162.pdf
HERCA (2014): Position statement on use of handheld portable dental x-ray equipment. http://www.herca.org/uploaditems/documents/HERCA%20position%20statement%20on%20use%20of%20handheld%20portable%20dental%20x-ray%20equipment.pdf
IAEA (2017). Radiation Protection and Safety in Medical Uses of Ionizing Radiation. SSG-46 (to be published)
Johnson KB et al. (2014) Reducing the risk of intraoral radiographic imaging with collimation and thyroid shielding. Gen Dent.;62:34-40.
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography

## Slide 59: References
Ludlow JB et al (2008). Patient risk related to common dental radiographic examinations: the impact of 2007 International Commission on Radiological Protection recommendations regarding dose calculation. J Am Dent Assoc.;139:1237-43.
Parrott LA & Ng SY (2011). A comparison between bitewing radiographs taken with rectangular and circular collimators in UK military dental practices: a retrospective study. Dentomaxillofac Radiol. ;40:102-9.
Pittayapat P et al (2012). Forensic odontology in the disaster victim identification process. J Forensic Odontostomatol.;30:1-12.
Pittayapat P et al. (2010). Image quality assessment and medical physics evaluation of different portable dental X-ray units. Forensic Sci Int. ;201:112-7.
Roeder F et al. (2011). Spatial relation between a rigid (digital) intraoral X-ray receptor and longitudinal axes of maxillary teeth. Clin Oral Investig.;15:715-9.
Stewart Whitley A, Jefferson G, Holmes K, Sloane C, Anderson C, Hoadley G. (2015). Clark’s Positioning in Radiography. 13th ed. CRC Press: Taylor&Francis Group.
White SC & Pharaoh MJ. (2014). Oral Radiology: Principles and Interpretation. 7thed. St. Louis: CV Mosby Company.
Radiation Protection in Dental Radiology L05 Fundamentals of Intraoral Radiography