## Slide 1: Radiation Protection in Dental Radiology
Quality Assurance in Dental Radiology
L10
Training material developed by the International Atomic Energy Agency
in collaboration with: World Health Organization, FDI World Dental Federation, International Association of Dento-Maxillofacial Radiology, International Organization for Medical Physics, and Image Gently Alliance

## Slide 2: Educational Objectives
Understand the general aspects of QA in radiology, and particular QA and QC aspects in dental radiology
Able to perform QC tests for different dental radiographic equipment (if responsible for such task)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 3: Overview
General principles of QA and QC
QC test protocols per modality
Dosimetry and dose monitoring
Clinical image quality assessment
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 4: Overview
General principles of QA and QC
QC test protocols per modality
Dosimetry and dose monitoring
Clinical image quality assessment
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 5: Quality Assurance (QA) in radiology
“as an organized effort by the staff operating a facility to ensure that the diagnostic images produced are of sufficiently high quality so that they consistently provide adequate diagnostic information at the lowest possible cost and with the least possible exposure of the patient to radiation”
World Health Organization (WHO), 1982
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 6: Quality Assurance (QA) in radiology
International Basic Safety Standards (BSS), 2014
“Registrants and licensees shall establish a comprehensive Quality Assurance program for medical exposures with the participation of appropriate qualified experts in radiophysics taking into account the principles established by the WHO.”
Notes: Let as consider in more details the main requirements of the BSS in respect to the scope and the content of the QA program in diagnostic radiology.

## Slide 7: Elements of a QA program
Acceptance and commission of a radiographic device
Periodic testing (and maintenance) of performance throughout its lifespan
Determination and follow-up of imaging protocols
Estimation of radiation dose to patients, workers and public
Training of workers
Adherence to policies and procedures
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
QC

## Slide 8: Quality control (QC)
Deals with performance of the equipment after installation
Regular timetable
  Acceptance testing
  Commissioning testing
  Routine (periodic) testing
Usually involves a medical physicist
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 9: Acceptance & commissioning tests
Performed at installation of the equipment
Acceptance testing: a check of performance before clinical use, using pass/fail acceptability criteria
(cfr. national/international standards)
Commissioning testing: provides baseline values for image quality and radiation dose
  Compared with future periodic tests
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 10: Routine QC tests
Different timetables (daily, monthly, annually) can be used depending on the nature of the test
Safety system tests (e.g. emergency stop)
X ray tube performance (e.g. nominal kV)
Image quality tests (e.g. resolution, artefacts)
Display performance
Dosimetry
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 11: Overview
General principles of QA and QC
QC test protocols per modality
Dosimetry and dose monitoring
Clinical image quality assessment
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 12: Routine QC tests: safety system tests
Emergency stop button available & in working order
  Or: release of exposure button serves as stop
Audible signal / warning light during exposure
Mechanical obstruction
  e.g. CBCT detector hitting patient: automatic stop?
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 13: Specific QC test protocols
The following sections will provide an overview of specific QC test protocols used in dental radiology for intra-oral, extra-oral and CBCT equipment
These protocols were based on (a combination) of existing guidelines, and discrepancies may exist regarding the tests, periodicity, acceptability criteria, etc.
Please refer to any relevant guidelines for your country/region when developing a QC protocol, ensuring that it fits with conventions and regulations (e.g. IEC standards when available, such as 61224-3-4)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 14: QC test protocols: intra-oral
A series of test are recommended for intra-oral radiography, encompassing the X ray tube (incl. patient and scattered dose), film QC, digital receptor QC, and image display
More details, acceptability criteria, action levels, practical procedures: see AAPM 175 publication (AAPM, 2016) or similar documents applicable to your country/region
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 15: QC test protocols: intra-oral
Stability of the X ray tube head after it is released by the operator
Visual inspection of the tube housing
  Leakage radiation should be measured is damage to the tube’s shielding is suspected
X ray field size (i.e. collimation) and source-skin distance should be in accordance with regulations and relevant recommendations
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 16: QC test protocols: intra-oral
Test of X ray tube output:
  Beam quality: half-value layer
  Kilovoltage and exposure time, compared with nominal values
    See next slides for example
  Reproducibility of exposure
  Linearity of mA and/or mAs
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 17: Routine QC tests: X ray tube
Accuracy of tube voltage (kV)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans
AC
DC
Pre-heating period (excluded from analysis)

## Slide 18: Routine QC tests: X ray tube
Accuracy of exposure time
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans

## Slide 19: QC test protocols: intra-oral
When films are used:
  QC of darkroom and/or film processors
    To avoid film fogging due to exposure to light
  QC of film processing (see also L04/L05)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 20: QC test protocols: intra-oral
When digital image receptors are used:
  Evaluation of uniformity, spatial resolution, contrast resolution
  Not only periodic, but also when damage is suspected (e.g. due to dropping the sensor)
  Evaluation of display performance (AAPM 2005, EFOMP 2015)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 21: Routine QC tests: display performance
Luminance meters, test patterns
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
AAPM TG18-QC multipurpose test pattern for evaluation of display monitors

## Slide 22: QC test protocols: intra-oral
Entrance surface dose
  Compare with local or international Diagnostic Reference Levels (see L11) and achievable doses
Scattered radiation
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 23: Specific QC tools for dental radiography
Intraoral radiography
  Step wedge (different thicknesses, typically of aluminium)
    Find exposure at which a certain darkening is found (which is predefined, e.g. corresponding to the required exposure for a maxillary molar)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans

## Slide 24: Specific QC tools for dental radiography
Intraoral radiography
  Line pair pattern (spatial resolution)
  Contrast-detail pattern
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans

## Slide 25: QC test protocols: extra-oral
General tests (tube leakage, kV, reproducibility, film QC, digital QC) for extra-oral radiography (incl. panoramic radiography) are the same as for intra-oral radiography
  But: different specifications or requirements may apply
    Beam quality: minimum HVL value
    Dose: measurement methodology, DRL value
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 26: QC test protocols: extra-oral
Specific test for panoramic radiography and ‘scanning-based’ cephalometric radiography: vertical/horizontal dimensions of X ray beam
  Should coincide with slit at image receptor
  Panoramic radiography: check alignment (image quality highly sensitive to proper patient alignment)
Broad-beam cephalometric radiography: field size should be compared with nominal values & light field
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 27: Specific QC tools for dental radiography
Panoramic radiography
  Uniformity
    Add filtration to tube, make ‘blank’ exposure
    Central area will be darker (compensation for cervical spine)
    (note: the horizontal lines in the image below are the result of suboptimal detector calibration)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans

## Slide 28: Specific QC tools for dental radiography
Panoramic radiography
  Alignment
    Position test object as a patient
    Evaluate image (circular dots: good alignment, elliptical dots: poor alignment)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans

## Slide 29: Image quality tools for dental radiography
Panoramic radiography
  Beam size
    Width (at secondary collimation, detector side) and height
    Radiochromic film attached to X ray tube
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans

## Slide 30: QC test protocols: CBCT
Several available guidelines regarding QC procedures CBCT
  European Commission, Radiation Protection 172 (SEDENTEXCT guidelines) (EC, 2012)
  UK Health Protection Agency (HPA, 2010)
  German DIN 6868-161 / DIN 6868-15
  Unified protocol for CBCT (not exclusive for dental) (EFOMP-ESTRO-IAEA, 2017)
  AAPM task group 261 (upcoming)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 31: QC test protocols: CBCT
Standard tests: cfr. manufacturer’s specifications, recent guidelines, national regulations
  X ray tube potential, tube leakage, total filtration or HVL, repeatability/reproducibility, beam collimation, slice thickness, display performance, visual inspection of image artefacts, dose to operator
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 32: QC test protocols: CBCT
Specific image quality tests
  Uniformity: stability of grey values
    Conventionally: compare grey value for different regions in the FOV for a homogeneous object (intra-scan uniformity)
    In CBCT, if grey values are to be used as HU (see further), grey values for scans with central/peripheral test object position should be assessed as well (inter-scan uniformity)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
Pauwels et al. (2015), under the British Institute of Radiology's License to Publish

## Slide 33: QC test protocols: CBCT
Specific image quality tests
  Geometrical accuracy: check on raw data (mechanical movement of equipment and reconstructed image (linear/angular measurements)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
R. Pauwels

## Slide 34: QC test protocols: CBCT
Specific image quality tests
  Grey value (GV) stability: reproducibility of grey values for a number of materials over time
  Hounsfield unit (HU) accuracy (only for machines which claim to yield HU): compare GV for different materials with corresponding nominal HU, check stability under varying exposure geometries (FOV, position of test object)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
R. Pauwels

## Slide 35: QC test protocols: CBCT
Specific image quality tests
  Noise: standard deviation of grey values in a homogeneous (section of a) test object
  Can be combined with contrast (contrast-to-noise ratio, CNR) and/or measured using the same region of interest as that used for uniformity
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
R. Pauwels

## Slide 36: QC test protocols: CBCT
Specific image quality tests
  Contrast resolution: contrast-to-noise ratio and/or visual / computer assisted evaluation of contrast-detail
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
R. Pauwels

## Slide 37: QC test protocols: CBCT
Specific image quality tests
  Spatial resolution
  Modulation transfer function (MTF)
    Evaluates decrease in contrast at increasing resolution
    Limit for visual perception at ~10% MTF
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
R. Pauwels

## Slide 38: Image quality tools for dental radiography
CBCT: combined phantoms
  SEDENTEXCT IQ phantom
  DIN phantom (different models)
  …
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
SEDENTEXCT IQ
Example of DIN 6868-161 phantom

## Slide 39: QC test protocols: CBCT
Specific image quality tests
  Spatial resolution
  Point / line / edge spread function
  (are related to MTF)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
R. Pauwels

## Slide 40: Overview
General principles of QA and QC
QC test protocols per modality
Dosimetry and dose monitoring
Clinical image quality assessment
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 41: Routine QC tests: dosimetry
Estimation of patient dose (‘average’ patient)
Using test objects (phantoms)
Different metrics used for different modalities
  ESD, DAP/KAP, CTDI, DLP, …
See L12 for more information on worker dose & protection
Complemented by dose monitoring (see further)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 42: Routine QC tests: dosimetry
Intra-oral radiography: entrance skin dose (cfr. local legislation)
Panoramic/cephalometric radiography: dose-area product (DAP)
CBCT: DAP most commonly used
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 43: Patient dose monitoring
During clinical use
Typically relies on display of dose estimates by the equipment (not direct measurement)
  Straightforward collection for groups of patients
  Accuracy of dose estimates should be verified
Allows for the determination of (and comparison with) diagnostic reference levels (DRLs; see L11)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 44: Overview
General principles of QA and QC
QC test protocols per modality
Dosimetry and dose monitoring
Clinical image quality assessment
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 45: Clinical image quality assessment
Difficult to relate technical parameters (MTF, CNR) to clinical acceptability
  Variations of technical parameters over time can indicate an issue, but imaging performance should be verified based on clinical image quality
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
Pauwels et al. (2014b), under the British Institute of Radiology's License to Publish
Which is (not) acceptable?

## Slide 46: Clinical image quality assessment
Reference image (of excellent or adequate quality, with the latter being more suitable in terms of optimization)
  Provided by manufacturer or acquired during acceptance testing (skull phantom)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
R. Pauwels

## Slide 47: Clinical image quality assessment
Reject analysis
  Frequency of rejects for all images acquired during a certain period
  Possible cause (e.g. operator error, patient movement, excessive artefacts, over- or underexposure) recorded
  Target (e.g. 5%) can be set
  Corrective actions (e.g. removal of metal objects in CT, improved positioning procedures, revision of exposure protocols)
Quality analysis on a sample of clinical images (standardized visual grading using dedicated image quality criteria)
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 48: Possible reasons for rejection
Under/overexposure: could be due to inconsistent tube output or improper selection of exposure settings
  Poor contrast in intra-oral radiography: kV as likely cause
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
H. Bosmans

## Slide 49: Possible reasons for rejection
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
Excessive noise: could be due to inconsistent tube output or improper selection of exposure settings
R. Pauwels

## Slide 50: Possible reasons for rejection
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology
Artefacts: may require detector recalibration, evaluation of film processing, …
H. Bosmans
Exposure to
visible light
Stains due to
developer
Bleaching
(overexposure
to fixator)

## Slide 51: References
AAPM (2016) Report No. 175 - Acceptance Testing and Quality Control of Dental Imaging Equipment. http://www.aapm.org/pubs/reports/RPT_175.pdf
AAPM (2005) Online Report No. 03, Assessment of Display Performance for Medical Imaging Systems. http://aapm.org/pubs/reports/OR_03.pdf
EFOMP( 2015) Quality controls in digital mammography – Protocol of the EFOMP mammo working group. http://www.efomp.org/index.php/scientific-guidance-and-protocols/351- mammo-protocol
European Commission (2012) Radiation Protection No 172: Cone beam CT for dental and maxillofacial radiology. Evidence based guidelines. Luxembourg: Office for Official Publications of the European Communities.
HPA, Health Protection Agency (2010) Recommendations for the design of X-ray facilities and quality assurance of dental Cone Beam CT (Computed tomography) systems.
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology

## Slide 52: References (cont.)
Deutsches Institut für Normung (2013) DIN 6868-161: Image quality assurance in X-ray departments - Part 161: RöV acceptance testing of dental radiographic equipment for digital cone beam computed tomography. Berlin: Deutsches Institut für Normung. Original and English translation available at https://www.beuth.de/en/standard/din-6868-161/164214522.
Deutsches Institut für Normung (2015) DIN 6868-15: Image quality assurance in X-ray departments - Part 15: RöV constancy testing of X-ray installations for dental radiographic equipment for digital cone-beam computed tomography. Berlin: Deutsches Institut für Normung [in German].
Pauwels R et al. (2014) A pragmatic approach to determine the optimal kVp in cone beam CT: balancing contrast-to-noise ratio and radiation dose. Dentomaxillofac Radiol. 43(5):20140059.
Pauwels R et al (2015) CBCT-based bone quality assessment: are Hounsfield units applicable? Dentomaxillofac Radiol. 44(1):20140238. doi: 10.1259/dmfr.20140238.
WHO (1982) Quality assurance of health services : concepts and methodology
Radiation Protection in Dental Radiology L10 Quality Assurance in Dental Radiology