## Slide 1: Radiation Protection in Dental Radiology
General Principles of Film and Digital Radiography
L04
Training material developed by the International Atomic Energy Agency
in collaboration with: World Health Organization, FDI World Dental Federation, International Association of Dento-Maxillofacial Radiology, International Organization for Medical Physics, and Image Gently Alliance

## Slide 2: Educational Objectives
To understand the principle of image formation in film and digital radiography, and fundamental differences between them
To be familiar with image manipulation methods in digital radiography
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 3: Introduction: Film and Digital Radiography
To create a radiographic image, an image receptor needs to be placed opposite to the X ray tube
Both analogue (i.e. film) and indirect/direct digital receptors are used
While the resulting image looks highly similar, there are fundamental differences between the way images are formed using film and digital receptors, as well as differences in terms of image manipulation, handling and storage
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 4: Overview
Radiographic film
  Components
  Film-screen systems
  Film development
  Film image characteristics
Digital image receptors
  PSP
  CCD/CMOS
  Digital image characteristics
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 5: Overview
Radiographic film
  Components
  Film-screen systems
  Film development
  Film image characteristics
Digital image receptors
  PSP
  CCD/CMOS
  Digital image characteristics
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 6: Radiographic film: components
X ray film consists of emulsion (sensitive to X-rays & light) and base (plastic surface coated with emulsion)
Double-emulsion films (figure): less exposure required
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Overcoat/abrasion layer
Overcoat/abrasion layer
Emulsion
Emulsion
Base

## Slide 7: Radiographic film: components
Emulsion:
  Silver halide grains (AgBr crystals)
    Different structure for E/F-speed films vs. D-speed; hence, a lower exposure can be used for the former
  Vehicle matrix (for suspension of crystals)
    Allows chemicals to reach grains during processing (see further)
    Additional layer added as overcoat for protection (e.g. scratching)
Base
  Polyester polyethylene terephtalate (translucent, flexible material, withstands processing solutions)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 8: Radiographic film: components
During exposure (1):
  X ray photons cause silver halide crystals to release electrons (from Br)
  Electrons reach ‘sensitivity site’ (i.e. areas with free Ag ions & trace chemicals) & are trapped
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Ag Br Ag Br
Ag Br Ag Br
Br Ag Br Ag
Br Ag Br Ag
Sensitivity
site
Ag Br Ag Br
Ag Br Ag Br
Br Ag Br Ag
Br Ag Br Ag
Sensitivity
site
Photon
e-

## Slide 9: Radiographic film: components
During exposure (2):
  Negative charge at sensitivity site attracts free Ag ions, which are reduced forming neutral silver, resulting in a ‘latent image’
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Ag Br Ag Br
Ag Br Ag Br
Br Ag Br Ag
Br Ag Br Ag
Sensitivity
site
Ag Br Ag Br
Ag Br Ag Br
Br Ag Br Ag
Br Ag Br Ag
-
Ag+
Ag

## Slide 10: Film-screen systems
Intensifying screens
  Phosphorescent (often rare-earth materials e.g. La, Gd & small amounts of Tm/Nb/Tb)
  Expose visible light when absorbing X-ray energy
  Light sensitivity should be matched with that of film, speed of film-screen stystem should be at least 400 for extraoral radiography (EC 2004)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Protective coating
Phosphor layer
Reflective layer (e.g. TiO2)
Support layer (e.g. plastic)
Cross-section of an intensifying screen

## Slide 11: Film-screen systems
Intensifying screens
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Casette front (plastic / thin metal)
Sponge rubber
Front intensifying screen
Film (emulsion and base)
Back intensifying screen
Sponge rubber
Casette back (thick metal)
  Reduced exposure possible (through amplification of signal), but loss in sharpness (through lateral spread of light)
  Typically provided by manufacturer
  Used in panoramic and other 2D extra-oral imaging; not in intra-oral radiography
K. Horner

## Slide 12: Film development
During processing:
  Film immersed in developer
  Developer rinsed off in water bath
  Film immersed in fixer
  Fixer removed in water bath
  Drying
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 13: Film development
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Before exposure, the film emulsion consist of silver bromide crystals (grey)

## Slide 14: Film development
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
After exposure (before development), the film contains a latent image, consisting of neutral silver atoms at sensitivity sites which were exposed to X rays (green spots)

## Slide 15: Film development
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
The developer converts the neutral silver atoms into solid grains of metallic silver (black)
Darkening of film proportionate to X ray exposure in that area

## Slide 16: Film development
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
The fixer dissolves unexposed/undeveloped silver bromide crystals; only solid silver grains are retained, which form the radiographic image

## Slide 17: Film development
Manual or automatic processing
  Manual:
    Darkroom needed to avoid light contamination
    Developer/solution should be stirred to equalize solutions and temperature
    Development time should be adapted to solution temperature
  Automatic:
    Transport mechanism (rollers)
    Saves time (4-6 min for complete process)
    No darkroom required (shielded compartment)
    More consistent density/contrast
    Rollers can induce artefacts, developer temperature higher: image quality of careful manual development is better
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 18: Film image characteristics
Films are distinguished based on various characteristics
Radiographic density
  Optical density (OD) = log10 (incident light/transmitted light)
  Characteristic curve: OD vs. (log of) exposure
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 19: Film image characteristics
Film contrast (↔ subject contrast)
  Average slope of (diagnostic part of) characteristic curve
Low contrast
High contrast
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 20: Film image characteristics
Radiographic speed
  Amount of exposure required to reach a given OD
  Currently available: D, E/F
  High-speed films result in a lower patient dose
Low speed
High speed
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 21: Film image characteristics
Film latitude
  Range of ‘detectable’ exposures, inversely related with film contrast
High latitude
Low latitude
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 22: Overview
Radiographic film
  Components
  Film-screen systems
  Film development
  Film image characteristics
Digital image receptors
  PSP
  CCD/CMOS
  Digital image characteristics
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 23: Digital image receptors: PSP
Photostimulable phosphor (i.e. indirect digital)
  Stores energy from X rays
  Releases energy (as light) when stimulated by light of different wavelength
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
R. Schulze

## Slide 24: Digital image receptors: PSP
  Barium fluorohalide (crystal lattice)
    Europium creates imperfections in lattice
    Before exposure: all electrons (e-) are in valence band
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Conduction band
Valence band
e-
F center
Eu2+

## Slide 25: Digital image receptors: PSP
  Energy due to exposure to X rays causes electrons to move to the conduction band
  Eu2+ → Eu3+
  Some electrons get trapped in F centers
  Number of trapped electrons is proportionate to exposure
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Conduction band
Valence band
e-
F center
Eu3+
X ray

## Slide 26: Digital image receptors: PSP
  During scanning of the plate, energy from a red laser releases trapped electrons
  When electrons move back to the valence band, energy is released as green light, which is detected by a photomultiplier tube
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Conduction band
Valence band
F center
Eu2+
e-
Laser
Photomultiplier
tube

## Slide 27: Digital image receptors: PSP
  The plates are then flooded with white light, ensuring all electrons are returned to the valence band before re-use
  Light source: dedicated scanner or manual (e.g. place against viewbox)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Conduction band
Valence band
e-
F center
Eu2+
Light source

## Slide 28: Digital image receptors: PSP
Fiberoptics + photomultiplier tube for read-out
  Light → electrical energy
  Red filter to remove light from scanning laser
  Analog-to-digital converter
  Read-out should be as soon as possible after exposure, without exposure to light (incl. red safelight in dark room!)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 29: Digital image receptors: CCD/CMOS
a.k.a. solid-state “sensors” (direct digital)
  Charge generated by X rays, collected in solid semiconductor material
  Image immediately available after exposure
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
R. Schulze

## Slide 30: Digital image receptors: CCD/CMOS
Plastic housing for protection
  Intra-oral sensors tend to be bulky i.e. less comfortable for patient
  Electronic components are increasingly miniaturized
Cable (can complicate positioning & increase vulnerability) or wireless (increases bulk)
Intra-oral sensors:
  different sizes (cfr. film) &
  pixel size (~20-70 µm)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 31: Digital image receptors: CCD/CMOS
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Insulation
n-type silicon
p-type silicon
(Optical fibre)
(Scintillator)
Electrode
1 pixel
CCD
  Silicon wafer insulated from electrodes (which form a pixel matrix)
  Optional: optical fibre (to improve resolution) and scintillator (to improve efficiency)

## Slide 32: Digital image receptors: CCD/CMOS
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Insulation
n-type silicon
p-type silicon
(Optical fibre)
(Scintillator)
e-
+
  Exposure to X rays leads to covalent breaks in the silicon, resulting in electron (e-) / hole (+) pairs (amount = proportionate to exposure)
  Resulting positive / negative charge distribution leads to electrical field
Electrical field
X ray

## Slide 33: Digital image receptors: CCD/CMOS
CMOS
  Also silicon-based, but read-out is different
    CCD: row-by-row
    CMOS: each pixel connected to separate transistor
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 34: Digital image receptors: CT/CBCT
In CT and CBCT, digital image receptors are used
  Arranged in (multiple) 1D arrays in (MD)CT
  Arranged as a 2D flat panel detector in CBCT
  Capable of fast X ray detection and read-out (several times per second)
  See L08 for more details
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 35: Digital image characteristics
A two-dimensional digital image, just like a digital image receptor, consist of a rectangular array (matrix) of square picture elements or ‘pixels’
A 3D image (CT/CBCT) consists of cubical elements or voxels (see L08)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 36: Digital image characteristics
In digital radiography, each pixel is assigned a grey value (or ‘intensity value’)
  According to relative attenuation (see L03)
  Brighter values indicate higher attenuation/densities for conventional purposes (cfr. film)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 37: Digital image characteristics
Unlike in film imaging, the ‘darkness’ of a digital radiograph is not exclusively determined by the amount of exposure
Digital receptors can receive signals between 0 (no X ray hitting a pixel) and a maximum value determined by the receptor’s dynamic range (cfr. latitude of a film)
The final digital radiograph is the result of post-processing, in which grey value are assigned depending on the relative detector signal in order to optimize contrast
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 38: Digital image characteristics
In case of under- or over-exposure, grey values can be rescaled in order to achieve adequate contrast
This also means that overexposure will not be detected on the image until the dynamic range of the receptor is exceeded
  Exposure indicators should be checked to avoid routine overexposure
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 39: Digital image characteristics
Image quality (see also L03) is determined by:
  Spatial resolution
    Pixel dimension (matrix size)
    Blurring
  Contrast
    Bit depth (number of possible grey values per pixel e.g. 12 bit = 4096 grey value possible)
    Characteristic function (signal at receptor transferred to histogram of final image)
    Noise
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 40
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Effects of Pixel size and Bit-depth
For an image of a M x N matrix, k bytes/pixel,
the memory needed to store the image is k x M x N bytes
32x32
64x64
128x128
512x512
1 bit
2 bit
4 bit
8 bit

## Slide 41: Digital image manipulation
Unlike a developed film, which is a ‘final’ product, several modifications can be made to a digital image
  Brightness/contrast adjustment
  Zoom
  Measurement of distance/area/volume
  Image enhancement/filtering
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 42: Digital image manipulation
Window/level
  Cfr. Brightness/contrast (but different principle); used to optimize the visualization/discrimination of tissues/pathology (usually first step when viewing an image)
  Typical CBCT image: ≥12 bit (i.e. ≥4096 grey values), but human eye can only discriminate ~200 grey levels
  Window = width of the grey value range which is displayed
  Level = middle grey value of the Window
  All grey value outside the window: shown as black (<) or white (>)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 43: Digital image manipulation
Grey value histogram of a CBCT image
Grey value
Number of voxels
Range of grey values being displayed
Window
Level
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 44: Digital image manipulation
Pauwels et al. (2015), under the British Institute of Radiology's License to Publish
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 45: Digital image manipulation
Zooming ensures that the perceived image sharpness is determined by the resolution of the image, not the monitor
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 46: Digital image manipulation
Measurements
  Length, area, volume can be measured if the image is calibrated according to distance (pixel/voxel size can be expressed in mm)
    CT/CBCT: inherently calibrated to absolute distance
    2D radiography: pre-calibrated (if imaging geometry is constant) or post-calibrated (using marker of known distance)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 47: Digital image manipulation
Image filters
  e.g. sharpening (to enhance small details) / smoothing (to suppress image noise)
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography
Original
Smoothed
Sharpened

## Slide 48: Overall benefits of digital imaging
Availability
  A digital image can be transmitted electronically to distant locations and can exist simultaneously at multiple locations
Flexibility
  The appearance of a digital image can be modified
Convenience
  A digital image can be stored electronically without occupying physical space
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography

## Slide 49: References
EC, European Commission (2004): European guidelines on radiation protection in dental radiology. The safe use of radiographs in dental practice. Radiation protection nr. 136.
Pauwels et al. (2015). Technical aspects of dental CBCT: state of the art. Dentomaxillofac Radiol. 2015;44(1):20140224.
White SC & Pharaoh MJ (2014). Oral Radiology: Principles and Interpretation. 7th ed. St. Louis: CV Mosby Company.
Radiation Protection in Dental Radiology L04 General Principles of Film and Digital Radiography