## Slide 1: Radiation Protection in Dental Radiology
Fundamentals of Panoramic Radiography
L06
Training material developed by the International Atomic Energy Agency
in collaboration with: World Health Organization, FDI World Dental Federation, International Association of Dento-Maxillofacial Radiology, International Organization for Medical Physics, and Image Gently Alliance

## Slide 2: Educational Objectives
Understand the general principles of image acquisition in panoramic radiography
Recognize and identify causes of image distortion and artefacts
Able to apply proper patient positioning and alignment
Recognize image aberration due to mispositioning
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 3: Overview
Introduction
Panoramic image acquisition
Magnification & distortion
Real, double & ghost images
Patient positioning
Errors in panoramic radiography

## Slide 4: Overview
Introduction
Panoramic image acquisition
Magnification & distortion
Real, double & ghost images
Patient positioning
Errors in panoramic radiography
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 5: What is a panoramic radiograph?
A two-dimensional panoramic view of the upper and lower jaw
Not a ‘one-shot’ 2D radiograph: image is acquired through sequential exposures using a narrow X-ray beam, during (reciprocal) movement of the X-ray tube and image receptor along an arced trajectory
Also known as orthopantomograph (trade name, best avoided)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 6: Indications for panoramic radiography
  Overall evaluation of dentition
  Impacted teeth
  Cysts, tumours, infection in jaw bones
  Dentomaxillofacial trauma
  Developmental disorders of maxillofacial skeleton, abnormal development of teeth
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 7: Technical specifications
Based on a survey of current equipment:
  kV typically between 60 and 85
  mA typically between 1 and 10
    New generation: kV & mA varies during acquisition, higher exposure when tube is at back of patient
  Exposure time: ~5-18 s (modern high-speed modes: as short as 2 s)
  Filtration: usually >2.5 mmAl
  Focal spot: 0.5 mm
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 8: Overview
Introduction
Panoramic image acquisition
Magnification & distortion
Real, double & ghost images
Patient positioning
Errors in panoramic radiography
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 9: Panoramic image acquisition
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
The following slides illustrates the principle of panoramic image acquisition
A more detailed description can be found in White & Pharoah (2014)
Terminology may vary

## Slide 10: Panoramic image acquisition
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Imagine a rotating disc exposed with a stationary narrow X ray beam, and an image receptor moving (at the same speed as objects A-D on the rotating disc) along a narrow slit in a collimator
Receptor
A
B
C
D
Collimator with slit
X ray tube (narrow beam)

## Slide 11: Panoramic image acquisition
B
C
D
A
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
When objects A-D pass along the slit, they will appear on the image receptor sharply (since they are moving at the same speed and in the same direction as the receptor)
Receptor
Collimator with slit
X ray tube (narrow beam)

## Slide 12: Panoramic image acquisition
B
C
D
A
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Objects closer to the X ray tube are moving at slower speeds ( ) or in opposite direction to the receptor ( ) and will therefore appear blurry on the receptor
Receptor
Collimator with slit
X ray tube (narrow beam)

## Slide 13: Panoramic image acquisition
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
The same principle applies if the disc is stationary and, instead, both the X ray tube and collimator-receptor rotate around the center of the disc ( ) (while the receptor is moving along the slit)
A
B
C
D
X ray tube (narrow beam)
Rotational center

## Slide 14: Panoramic image acquisition
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
If a patient is exposed using this set-up, the translational speed of the receptor and the center of rotation can be determined in order to sharply visualize the jaw bones
X ray tube (narrow beam)
Mandible

## Slide 15: Panoramic image acquisition
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Due to the shape of the human jaw bones, the rotational center should not be fixed
Modern panoramic equipment use a continuously moving center of rotation ( ) to ensure optimal sharpness along the jaw arch
Mandible
Trajectory of center of rotation

## Slide 16: Panoramic image acquisition
The result is a panoramic radiograph in which only structures along the dental arch are in focus
  ‘Focal trough’ or ‘image layer’ ( )
  Different rotation paths available for ≠ jaw shapes
  The closer a structure is to the center of the trough ( ), the sharper it will appear
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
D. Schulze

## Slide 17: Panoramic image acquisition
As a result of the unique acquisition process, dose distribution inside the patient is highly inhomogeneous (highest dose typically in the parotid gland region)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
T. Sullivan, S. Tabakov

## Slide 18: Special types of panoramic radiographs
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
To limit radiation exposure, collimated panoramic modes are available e.g. ‘child size’ and ‘posterior region’
K. Horner

## Slide 19: Overview
Introduction
Panoramic image acquisition
Magnification & distortion
Real, double & ghost images
Patient positioning
Errors in panoramic radiography
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 20: Magnification in panoramic radiography
Image distortion is inherent to the panoramic technique: linear/angular measurements not reliable
Horizontal magnification
  Depends on relative distance to focal trough
  Different for anterior & posterior structures (‘snowman principle’)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Middle of trough: no magnification
Anterior: horizontal narrowing
Posterior: horizontal widening

## Slide 21: Magnification in panoramic radiography
Vertical magnification
  Depends on source-object distance (cfr. planar radiography)
  Distance can be constant for certain units
  Slight difference in magnification between edges of focal trough (due to difference in distance to source)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
X ray tube
Image
receptor

## Slide 22: Magnification in panoramic radiography
Vertical distortion
  Due to upward angle of X-ray beam
  Structures closer to tube (i.e. lingually) ( ) projected higher on image receptor than structures closer to receptor ( )
    e.g. cervical spine vs. mandible
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
X ray tube
Image
receptor

## Slide 23: Overview
Introduction
Panoramic image acquisition
Magnification & distortion
Real, double & ghost images
Patient positioning
Errors in panoramic radiography
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 24: Real/double/ghost images
Real image
  Object is between center of rotation ( ) and receptor
  Sharper if closer to focal trough
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak
Ear ring
X ray tube

## Slide 25: Real/double/ghost images
Ghost image
  Object between x-ray source and center of rotation
  Appears on opposite side of image (L↔R) and higher on the image (due to upward beam) than their anatomical location
  Far outside focal trough: appears blurred and magnified
  Mandibular ramus, hyoid bone, cervical spine, earrings, etc.
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak
Ear ring
Image
receptor

## Slide 26: Real/double/ghost images
Double image
  Object is posterior to center of rotation and intercepted by beam twice (area marked by )
  Hyoid bone, epiglottis, cervical spine, palate
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Starting angle of panoramic exposure
End angle of panoramic exposure
Trajectory of
center of rotation

## Slide 27: Real/double/ghost images
Double image
  First projection of points A and B (A is projected first, then B)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Image
Receptor
B
A
A
A
B
A
B
Panoramic image
(exposed section of image shown in grey )

## Slide 28: Real/double/ghost images
Double image
  Second projection of points A and B (B is projected first, then A)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
B
A
B
A
B
Panoramic image
A
B
B
A

## Slide 29: Real/double/ghost images
Double image
  Double images are mirror images, appearing on the same location but opposite sides
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
P. Pittayapat
Hard palate
Cervical spine

## Slide 30: Overview
Introduction
Panoramic image acquisition
Magnification & distortion
Real, double & ghost images
Patient positioning
Errors in panoramic radiography
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 31: Patient positioning
Panoramic radiography is highly sensitive to correct patient positioning
  Mandibular and maxillary structures should be positioned in the focal trough
  Improper head orientation and tongue position can lead to shading and other artefacts
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 32: Patient positioning
Metal objects such as earring should be removed to avoid interference with image interpretation
Possible collision of the tube/receptor with the patient’s shoulder should be checked
  “Test run” (no exposure) can be used as verification
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 33: Patient positioning
Patient bites in bite block or uses chin rest
  Corresponds with anterior part of focal trough to ensure coverage of anterior teeth (focal trough is thinnest in this region)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 34: Patient positioning
Head alignment done in three different planes (with help of alignment lights and mirrors)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Midsagittal plane: centered
Occlusal/Frankfort plane: parallel to floor
A-P plane: machine-specific alignment
Neck straight to avoid excessive overlap cervical spine

## Slide 35: Patient positioning
Patient instructions before acquisition
  No movement (~15-20s; fast modes available when needed)
  Close lips, place and hold tongue against palate
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 36: Patient positioning
Essential image characteristics (cfr. Clark’s Positioning in Radiography)
  Edge to edge incisors
  No removable metallic foreign bodies
  No motion artefacts
  Tongue is against roof of mouth
  Minimization of spine shadow
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 37: Patient positioning
Essential image characteristics (cfr. Clark’s Positioning in Radiography) (cont.)
  No A-P positioning errors
  No mid-sagittal plane positioning errors
  No occlusal plane positioning errors
  Correct position of spinal column
  Good density and adequate contrast between enamel and dentine
  Name/date/left-right markers visible
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 38: Overview
Introduction
Panoramic image acquisition
Magnification & distortion
Real, double & ghost images
Patient positioning
Errors in panoramic radiography
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography

## Slide 39: Patient positioning
Proper patient positioning
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
Center of focal trough in anterior region
D. Schulze

## Slide 40: Patient positioning
Head too much forward; ‘minification’
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
K. Horner

## Slide 41: Patient positioning
Head too much backward; magnification
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
K. Horner

## Slide 42: Patient positioning
Chin too high; upward overangulation
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
K. Horner

## Slide 43: Patient positioning
Chin too low; downward overangulation
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
K. Horner

## Slide 44: Patient positioning
Rotation & shifting; left-right distortion
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak

## Slide 45: Patient positioning
Neck extension (e.g. patient standing too much back); cervical spine overlap
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak

## Slide 46: Patient positioning
Collision with shoulder
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak

## Slide 47: Patient positioning
Patient movement (slight)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak

## Slide 48: Patient positioning
Patient movement (severe)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
K. Horner

## Slide 49: Patient positioning
Improper tongue placement
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak

## Slide 50: Patient positioning
Lead apron artefact (see also L11)
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography
O. Silkosessak

## Slide 51: References
Stewart Whitley A, Jefferson G, Holmes K, Sloane C, Anderson C, Hoadley G. (2015). Clark’s Positioning in Radiography. 13th ed. CRC Press: Taylor&Francis Group.
White SC & Pharaoh MJ. (2014). Oral Radiology: Principles and Interpretation. 7th ed. St. Louis: CV Mosby Company.
Radiation Protection in Dental Radiology L06 Fundamentals of Panoramic Radiography