## Slide 1: Radiation Protection in Dental Radiology
Optimization of Protection of Patients in Dental Radiology
L11
Training material developed by the International Atomic Energy Agency
in collaboration with: World Health Organization, FDI World Dental Federation, International Association of Dento-Maxillofacial Radiology, International Organization for Medical Physics, and Image Gently Alliance

## Slide 2: Educational Objectives
Be familiar with the (equipment and patient) factors affecting patient dose
Be able to apply various optimization strategies, and estimate the amount of dose that can be saved in a given exposure condition
Be able to judge the potential benefit and drawback of using patient shielding in a given situation
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 3: Overview
The optimization principle
Exposure parameters and their effect
Diagnostic Reference Level (DRL)
Patient shielding
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 4: Overview
The optimization principle
Exposure parameters and their effect
Diagnostic Reference Level (DRL)
Patient shielding
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 5: The optimization principle
Underexposure
(high noise)
Optimized exposure
(adequate image quality)
Overexposure
(high image quality, but high dose)
The principle of optimization of protection: the likelihood of incurring exposures, the number of people exposed, and the magnitude of their individual doses should all be kept as low as reasonably achievable, taking into account economic and societal factors (ICRP, 2007)

## Slide 6: The optimization principle
Factors to consider in optimization
  Design considerations for equipment
  Operational considerations
  Calibration
  Dosimetry of patients
  Diagnostic Reference Levels (DRLs)
  Quality assurance
  Dose constraints

## Slide 7: Overview
The optimization principle
Exposure parameters and their effect
Diagnostic Reference Level (DRL)
Patient shielding
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 8: Exposure parameters
Several exposure parameters affect dose
  Beam energy (kV, filtration)
  Tube current (mA)
  Exposure time (s)
  Geometric factors (beam size, source-skin distance, …)
  (see L03 for more information)
For most parameters, image quality is affected as well: Always consider image quality & radiation dose simultaneously
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 9: Exposure parameters: kV
kV ↑ → Dose ↑ (if all other parameters are constant)
Absorbed dose in PMMA phantom, CBCT
Pauwels et al. 2014a, under the British Institute of
Radiology’s License to Publish
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 10: Exposure parameters: filtration
Attenuation of an X ray beam depends on X ray energy
  When X ray traverse a medium, low-energy X rays are preferentially absorbed
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Low energy X rays
Medium energy X rays
High energy X rays
Exposed
Transmitted

## Slide 11: Exposure parameters: filtration
Apart from the inherent filtration of the X ray tube, varying amount of filtration can be added
  For some equipment, the amount of filtration can be selected or additional filtration can be manually added
  Since low-energy X rays contribute to the radiation dose but not to the image, filtering them out can reduce dose
  Excessive filtration can reduce the tube output below that needed for a diagnostic image → mAs can be increased to compensate for this
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 12: Exposure parameters: kV
‘Optimal’ beam energy?
  Depends on modality
  Too low: too many low-energy photons + high mAs required for adequate image quality
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
R. Pauwels
Low kV, high mAs
Mid kV, mid mAs
High kV, low mAs
All scans were acquired at the same radiation dose

## Slide 13: Exposure parameters: kV
‘Optimal’ beam energy?
  Depends on modality
  Energy too high: contrast ↓ (see slide notes)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
High contrast
Medium contrast
Low contrast
R. Pauwels
Notes: The attenuation of X rays by a material is determined by this material’s linear attenuation coefficient (µ), which depends on its composition and density as well as X ray energy.
The linear attenuation coefficient is the product of the mass attenuation coefficient µm and the physical density ρ.
The graph in this slide shows the ratio between the µ of bone and the µ of soft tissue as a function of X ray energy. The higher this ratio, the higher the contrast between these tissues in a radiographic image.
At very high X ray energies, there is virtually no difference in µm between bone and soft tissue, and the difference in attenuation between these tissues is only determined by their relative difference in physical density.

## Slide 14: Exposure parameters: filtration
‘Optimal’ beam energy?
  In dental radiography: often fixed kV used for a given equipment, exposure is varied via adjustment of mA and/or exposure time
  If kV is variable (e.g. some CBCT equipment): optimal kV should be determined with the involvement of a medical physicist, considering the balance between image quality & dose
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 15: Exposure parameters: kV
‘Optimal’ beam energy: example
  CBCT with kV selectable between 60 and 90 kV
  Radiation dose was determined at a constant image quality in function of effective head size, kV and mA
  For each head size, 90 kV resulted in the lowest radiation dose → optimal kV for adults and children
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Pauwels et al. 2017

## Slide 16: Exposure parameters: mAs
mA * s = mAs (e.g. 5 mA, 10 s → 50 mAs)
mAs ↑ → X ray flux ↑ (if other parameters are constant)
  Linear relation
  No effect of mAs on X ray energy
mAs ↑ → Dose ↑ (also linear)
R. Pauwels
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Theoretical extrapolation of effective dose for a CBCT unit, based on an experimental measurement at 5 mA (dark grey bar)

## Slide 17: Exposure parameters: mAs
General effect on image quality: mAs ↑ → noise↓
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
mAs ↑
R. Pauwels

## Slide 18: Exposure parameters: mAs
Specific effects on image quality
  In CBCT: aliasing artefacts may occur if the exposure time (i.e. number of projections) is too low
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
R. Pauwels

## Slide 19: Exposure parameters: mAs
Specific effects on image quality
  In film radiography: optical density vs. exposure (see L04)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 20: Exposure parameters: mAs
‘Optimal’ mAs?
  Machine-specific
    e.g. different CBCT models, film vs. digital receptor, …
  Depends on clinical application (i.e. image quality requirements)
  Depends on patient size (esp. in CBCT)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 21: Exposure parameters: mAs
Age vs. mAs in CBCT at a constant noise level
(Percentiles refer to growth reference charts for head circumference)
Pauwels et al. (2017)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 22: Exposure parameters: mAs
‘After installation of an equipment, dedicated exposure protocols should be established
  Involvement of medical physicist recommended
  Avoid ‘routine’ exposure; optimization is patient-specific!
  Protocols should be reviewed and further optimized periodically (‘radiological review’; IAEA BSS Requirement 42)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 23: Exposure parameters: geometric
Intraoral radiography:
  Source-patient distance (SPD) can be varied → lower peak skin dose at increased SPD, but same absorbed dose (see further: inverse square law)
  The use of rectangular collimation can lead to a dose reduction of ~40-50% compared with circular collimation (Aps & Scott 2014, Stenström et al. 1987, Rehnmark-Larsson et al. 1982)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 24: Exposure parameters: geometric
CBCT & CT: collimation along craniocaudal axis possible for both, along AP/LR axis only for CBCT (MDCT: always covers full diameter of head)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Mandible
Maxilla
Maxillofacial
Full head
R. Pauwels & R. Pralier
X ray beam coverage in CT
X ray beam coverage in CBCT

## Slide 25: Exposure parameters: geometric
Panoramic radiography: different collimations (e.g. adult vs. child, TMJ only, …); collimation can reduce brain dose by >50% (Davis et al. 2015)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
K. Horner

## Slide 26: Exposure parameters: geometric
Effect of beam size on patient dose
  Larger beam size → larger volume of patient exposed
  At constant source-patient distance: X ray flux (amount of X rays per unit area) not affected by beam size
  Therefore, larger beam size → increased effective dose
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 27: Exposure parameters: geometric
Example 1: CBCT, small FOV (1 projection shown)
P
T
Sub
B
Direct exposure to parotid gland (P)
Small area of skin (S) exposed
Scatter exposure to submandibular and sublingual glands (Sub)
Negligible scatter exposure to brain (B) and thyroid gland (T)
S
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
X ray
beam
R. Pauwels & R. Pralier

## Slide 28: Exposure parameters: geometric
Example 2: CBCT, large FOV (1 projection shown)
P
T
Sub
B
Direct exposure + increased scatter to parotid gland (P)
Direct exposure to submandibular and sublingual glands (Sub)
Larger area of skin (S) exposed
(also: bone marrow, bone surface)
Increased scatter exposure to brain (B) and thyroid gland (T)
Increased effective dose
S
S
S
R. Pauwels & R. Pralier
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
X ray
beam

## Slide 29: Exposure parameters: geometric
Note: when adults and children are exposed using the same FOV, the dose for children will be higher!
  e.g. a FOV which covers upper & lower jaw for a large adult (left) would cover the entire maxillofacial region for a small child (right)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
R. Pauwels & R. Pralier

## Slide 30: Exposure parameters: geometric
Effective dose vs. FOV size in CBCT
Pauwels et al. 2014b, under the British Institute of
Radiology’s License to Publish
BJ, both jaws; LJ, lower jaw; LM, lower molar; MF, maxillofacial; UC, upper canine; UF, upper frontal; UJ, upper jaw.
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 31: Exposure parameters: geometric
CBCT: effective dose distribution vs. FOV height
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Based on
published data

## Slide 32: Exposure parameters: geometric
CBCT: horizontal and vertical stitching of multiple scans can be used to yield a larger FOV
May lead to higher patient dose (as there is a small overlapping region between the scans)
Horizontal stitching in particular is discouraged
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
R. Pauwels

## Slide 33: Exposure parameters: geometric
CBCT
  Effective dose & absorbed dose (for non-whole-body organs) also dependent on position of FOV
    Mandibular scan: higher effective dose (mainly due to ↑ dose thyroid, salivary glands, oesophagus, lymph nodes) than maxillary scan (see Figure slide nr. 14)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 34: Exposure parameters: geometric
Example 1: Maxillary scan
P
T
Sub
B
Direct exposure to parotid gland (P), scatter exposure to submandibular and sublingual glands (Sub)
Small area of skin (S) exposed
Negligible scatter exposure to brain (B) and thyroid gland (T)
S
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
R. Pauwels
R. Pralier

## Slide 35: Exposure parameters: geometric
Example 2: Mandibular scan (same FOV, mAs, etc.)
P
T
Sub
B
Direct exposure to submandibular and sublingual glands (Sub), direct+scatter exposure to parotid gland (P) → overall, increased dose to salivary glands vs. maxillary scan
Same area of skin (S) exposed
Increased scatter exposure to thyroid gland (T), oesophagus and lymph nodes
Increased effective dose
S
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
R. Pauwels
R. Pralier

## Slide 36: Exposure parameters: geometric
Panoramic radiography
  Dose distribution is highly localized due to rotational geometry
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
T. Sullivan, S. Tabakov

## Slide 37: Overview
The optimization principle
Exposure parameters and their effect
Diagnostic Reference Level (DRL)
Patient shielding
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 38: Diagnostic reference levels (DRLs)
International BSS:
  “DRL is a level used in medical imaging to indicate whether, in routine conditions, the dose to the patient in a specified radiological procedure for medical imaging is unusually high or unusually low for that procedure.”
Based on dose audits in a defined region
  Country, province, …
Patient dose estimation for all available equipment (usually for a given clinical indication, esp. in CT)
  e.g. Patient dose for [radiographic modality] for [clinical application], for all available equipment in [region]

## Slide 39: Diagnostic reference levels (DRLs)
Modality-specific measurable dose indices (not effective dose!)
Standardized, consistent measurement methodology
DRL = third quartile value of dose distribution
  75% of the equipment operates below the DRL
  Used to indicate the need for an investigation. No legal value (not be used as dose limit!), but equipment operating above DRL should be evaluated for possible actions to reduce dose
Value of DRL can change (ideally lower) over time; periodic re-evaluation needed

## Slide 40: Diagnostic reference levels (DRLs)
R. Pauwels; illustrative example, not based on real data
Graph showing radiation doses of a series of radiographic units, ordered from lowest to highest dose.
The dose below which 75% of the investigated units operate is the DRL

## Slide 41: DRLs in dental radiology
Panoramic radiography (post-2000 publications):
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Index | Region | Protocol | Detector | DRL | Reference
ESD (mGy) | Greece | adult | All | 4.1 | Manousaridis et al. (2015)
 |  |  | DR | 3.5 | 
 |  |  | CR | 4.2 | 
 |  |  | Film | 3.7 | 
 |  | small adult / female | All | 3.3 | 
 |  | child | All | 2.2 | 
 | Spain |  | All | 0.66 | González et al. (2001)
DWP (mGy mm) | Ireland | adult | N/A | 60 | Walker & van der Putten (2012)
 | South Korea |  | All | 60.1 | Lee et al. (2010)
 | UK | adult | All | 74 | HPA (2010)
 |  | child | All | 49 | 
DAP (mGy cm2) | South Korea | child | All | 95.9 | Kim et al. (2014)
 |  | adult | All | 120.3 | Han et al. (2012)
 | Greece | male | All | 117 | Tierris et al. (2004)
 |  | female | All | 97 | 
 |  | child | All | 77 | 
 | UK | adult | All | 93 | HPA (2010)
 |  | child | All | 67 | 
 | Germany | large adult | All | 101 | Poppe et al. (2007b)
 |  | adult male | All | 87 | 
 |  | adult female | All | 84 | 
 |  | child | All | 75 | 

## Slide 42: DRLs in dental radiology
Intra-oral radiography (post-2004 publications):
Radiation Protection in Dental Radiology L08 Protection of Patients in Dental Radiology
Index | Region | Protocol | DRL | Reference
DAP (mGy cm2) | Germany | Maxillary molar | 61.5 | Poppe et al. (2007a)
 |  | Maxillary premolar | 44.1 | 
 |  | Maxillary canine | 36.2 | 
 |  | Maxillary incisor | 37.9 | 
 |  | Mandibular molar | 41.2 | 
 |  | Mandibular premolar | 31.8 | 
 |  | Mandibular canine | 31.8 | 
 |  | Mandibular incisor | 26.2 | 
 |  | Bitewing (front) | 46.1 | 
 |  | Bitewing (back) | 47.3 | 
 |  | Occlusal (maxilla) | 87 | 
 |  | Occlusal (mandible) | 71.2 | 
 | South Korea | Maxillary molar | 55.5 | Han et al. (2012)
 |  | Maxillary premolar | 46 | 
 |  | Maxillary incisor | 36.5 | 
 |  | Mandibular molar | 87.4 | Kim et al. (2012)
ESD (mGy) | Greece | Maxillary molar | 1.2 (digital) | Manousaridis et al. (2013)
 |  |  | 3.7 (film) | 
 |  | Incisor | 0.65 (digital) | 
 |  |  | 2.35 (film) | 
 | Ireland |  | 2.3 (2.5 proposed) | Fanning (2008)
 | Ireland | Mandibular molar | 2.4 | Walker & van der Putten (2012)
 | South Korea | Mandibular molar | 3.1 | Kim et al. (2012)
 | Spain | Maxillary molar | 2.8 | Alcaraz et al. (2015)
 | UK | Adult | 2.3 | Hart et al. (2009)
 |  | Child | 1.5 | 
 | USA | Bitewing | 2.3 | Gray et al. (2005)

## Slide 43: DRLs in dental radiology
Lateral cephalogram
Index | Region | Protocol | DRL | Reference
DAP (mGy cm2) | Germany | adult | 33 | Looe et al. (2007)
 |  | child | 26 | 
 | South Korea | adult | 146 | Kim et al. (2014)
 |  | child | 121 | Han et al. (2012)
 | UK | adult | 41 | Holroyd (2011)
 |  | child | 25 | 
CBCT
  Tentative DRLs proposed, wide dose range (despite the use of similar technology) indicates dire need for optimization
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 44: DRLs in dental radiology
DRLs can reduce over time as a result of improvements in imaging technology, e.g.
  D-speed film vs. E/F-speed film & digital
  Improved detectors and reconstruction in CBCT
When old DRLs are considered, the applicability of these values in current practice should be questioned, and alternative DRLs should be looked up (e.g. 2005 DRL for Country 1 vs. 2015 DRL for Country 2: which is more suitable for Country 1?)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 45: Overview
The optimization principle
Exposure parameters and their effect
Diagnostic Reference Level (DRL)
Patient shielding
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 46: Patient shielding
The ‘paradox’ of shielding:
  Shielding materials block radiation entering the body
    Primary radiation & small amount of scattered radiation coming from other sources than the patient
    Patient-generated scatter: will travel inside the body: shielding not effective
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 47: Patient shielding
The ‘paradox’ of shielding:
  Therefore, shielding will be most effective when blocking primary radiation (i.e. being placed within the field of view)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Shielding is effective
(e.g. reduced thyroid dose)
Shielding is ineffective

## Slide 48: Patient shielding
The ‘paradox’ of shielding:
  However, shielding inside the field of view should be used with great care and is sometimes counter-advised
    Blocking areas of diagnostic interest
    Metal artefacts (CBCT, CT)
    Affects tube output for systems with automatic exposure control
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 49: Patient shielding
Lateral cephalogram:
  Cranial shielding (or wedge-shaped shielding)
  Thyroid shielding (fixed or collar)
  Possible reduction in effective dose: ~40% (thyroid shield), ~60% (thyroid + cranial shield) (Hoogeveen et al. 2015)
  Avoid overlap of shielding with landmarks
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
(Illustration of lateral cephalogram with cranial and thyroid shielding (white regions)

## Slide 50: Patient shielding
Intraoral radiography:
  Use of lead foil: 32-59% reduction in organ dose
  (Nejaim et al. 2015)
  Thyroid shielding: >25% reduction in thyroid dose
  (Johnson et al. 2014)
Nejaim et al. (2015) (With permission by Y. Nejaim)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 51: Patient shielding
Panoramic radiography:
  Thyroid collar:
    Thyroid (absorbed) dose reduced by 10-23%
    (Han et al. 2013)
    Effect of collar on thyroid dose <5%
    (Rottke et al. 2013b)
    (Note: reduction of absorbed dose to oesophagus expected to follow similar trend as thyroid)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 52: Patient shielding
Panoramic radiography:
  Thyroid collar may overlap with diagnostic region (due to upward angle of X ray beam): use with caution
  EC RP 136 (2004): “Thyroid shielding is inappropriate for panoramic radiography as it may interfere with the primary beam”
  Lead apron normally not used
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 53: Patient shielding
CBCT
  Thyroid collar:
    35-44% dose reduction to thyroid
    (Qu et al. 2012,Tsiklakis et al. 2005, Goren et al. 2013, Hidalgo et al. 2015)
    (shielding ≥ 0.25 mmPb)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 54: Patient shielding
Issues regarding use of thyroid collar in CBCT:
  Blocking of the primary beam by thyroid shield can lead to severe metal artefacts → do not use if region of interest is at the vertical level of the shielding (use scout image to verify)
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology
Hidalgo et al. 2015 (with permission)

## Slide 55: Patient shielding
Issues regarding use of thyroid collar in CBCT:
  For systems using automatic exposure control (currently rare in CBCT, but expected to become more widespread): thyroid collar in primary beam would lead to unnecessary increase in tube output
    Pre-scan AEC (scout-based): apply collar after scout
    In-scan AEC (real-time): ensure collar outside primary beam
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 56: References
Al-Okshi A et al. (2013) Using GafChromic film to estimate the effective dose from dental cone beam CT and panoramic radiography. Dentomaxillofac Radiol. 42:20120343.
Al-Okshi A et al. (2015) Effective dose of cone beam CT (CBCT) of the facial skeleton: a systematic review. Br J Radiol. 88:20140658.
Alcaraz M et al. (2015) Dose reference levels in Spanish intraoral dental radiology: stabilisation of the incorporation of digital systems in dental clinical practices. Radiat Prot Dosimetry. 172(4):422-427
Aps JK & Scott JM. (2014) Oblique lateral radiographs and bitewings; estimation of organ doses in head and neck region with Monte Carlo calculations. Dentomaxillofac Radiol. 43:20130419.
Avendanio BN et al. (1996) Effective dose and risk assessment from detailed narrow beam radiography. Oral Surg Oral Med Oral Pathol Oral Radiol Endod. 82:713-9.
Bornstein MM et al. (2014) Cone beam computed tomography in implant dentistry: a systematic review focusing on guidelines, indications, and radiation dose risks. Int J Oral Maxillofac Implants. 29 Suppl:55-77.
Cederberg RA et al. (1997) Effect of the geometry of the intraoral position-indicating device on effective dose. Oral Surg Oral Med Oral Pathol Oral Radiol Endod. 84:101-9.
Danforth RA & Clark DE. (2000) Effective dose from radiation absorbed during a panoramic examination with a new generation machine. Oral Surg Oral Med Oral Pathol Oral Radiol Endod. 89:236-43.
Davis AT et al. (2015) The reduction of dose in paediatric panoramic radiography: the impact of collimator height and programme selection. Dentomaxillofac Radiol. 44:20140223.
Dula KR et al. (1997) Hypothetical mortality risk associated with spiral tomography of the maxilla and mandible prior to endosseous implant treatment. Eur J Oral Sciences. 105:123-9.
Dula KR et al. (2001) The radiographic assessment of implant patients: decision-making criteria. Int J Oral Maxillofac Implants 16:80-9.
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 57: References (cont.)
EC, European Commission (1999) European guidance on diagnostic reference levels (DRLs) for medical exposures. Radiation Protection 109, European Commission, Luxembourg. [http://www.fanc.fgov.be/GED/00000000/3700/3756.pdf]
EC, European Commission (2004) European guidelines on radiation protection in dental radiology - The safe use of radiographs in dental practice, radiation protection publication 136. European Commission, Luxembourg. [https://ec.europa.eu/energy/sites/ener/files/documents/136.pdf]
Fanning B. (2008) The establishment of a diagnostic reference level for intra-oral dental radiography in Co. Wicklow. J Ir Dent Assoc. 54:181-3.
Frederiksen NL et al. (1995) Effective dose and risk assessment from computed tomography of the maxillofacial complex. Dentomaxillofac Radiol. 24:55-8.
Frederiksen NL et al. (1994) Effective dose and risk assessment from film tomography used for dental implant diagnostics. Dentomaxillofac Radiol 23:123-7.
Gijbels F et al. (2001) Diagnostic yield of conventional and digital cephalometric images: a human cadaver study. Dentomaxillofac Radiol. 30:101-5.
Gijbels F et al. (2002) A comparison of the effective dose from scanography with periapical radiography. Dentomaxillofac Radiol. 31:159-63.
González L et al. (2001) Reference doses in dental radiodiagnostic facilities. Br J Radiol. 74:153-6.
Gori C et al. (2000) Dose evaluation and quality criteria in dental radiology. Radiat Prot Dosimetry 90:225-7.
Gray JE et al. (2005) Reference values for diagnostic radiology: application and impact. Radiology. 235:354-8.
Han GS et al. (2013) Shielding effect of thyroid collar for digital panoramic radiography. Dentomaxillofac Radiol. 42:20130265.
Han S et al. (2012) Dose area product measurement for diagnostic reference levels and analysis of patient dose in dental radiography. Radiat Prot Dosimetry 150:523-31.
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 58: References (cont.)
Hart D et al. (2009) National reference doses for common radiographic, fluoroscopic and dental X-ray examinations in the UK. Br J Radiol. 82:1-12.
Hatzioannou K et al. (2005) Quality control and diagnostic reference levels in intraoral dental radiographic facilities. Dentomaxillofac Radiol. 34:304-7.
Hidalgo A et al. (2015) Effectiveness of thyroid gland shielding in dental CBCT using a paediatric anthropomorphic phantom. Dentomaxillofac Radiol. 44:20140285.
Hirsch E et al. (2008) Dosimetry of the cone beam computed tomography Veraviewepocs 3D compared with the 3D Accuitomo in different fields of view. Dentomaxillofac Radiol. 37:268–73.
Holroyd JR. (2011) National reference doses for dental cephalometric radiography. Br J Radiol. 84:1121-4.
Hoogeveen RC et al. (2015) Dose reduction in orthodontic lateral cephalography: dosimetric evaluation of a novel cephalographic thyroid protector (CTP) and anatomical cranial collimation (ACC). Dentomaxillofac Radiol. 44:20140260.
HPA (2010). HPA-CRCE-034: Doses to patients from radiographic and fluoroscopic x-ray imaging procedures in the UK (2010 review).
ICRP, International Commission on Radiological Protection (2007) The 2007 Recommendations of the International Commission on Radiological Protection. ICRP Publication 103. Oxford: Pergamon.
Johnson B et al. (2014) Intraoral radiographic imaging risk reduction with collimation and thyroid shielding. Gen Dent. 62:34-40.
Kim EK et al. (2012) Diagnostic reference levels in intraoral dental radiography in Korea. Imaging Sci Dent. 42:237-42.
Kim YH et al. (2014) Diagnostic reference levels for panoramic and lateral cephalometric radiography of Korean children. Health Phys. 107:111-6.
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 59: References (cont.)
Lee JS et al. (2010) Reference dose levels for dental panoramic radiography in Gwangju, South Korea. Radiat Prot Dosimetry 142:184-90.
Lee GS et al. (2013) Effective dose from direct and indirect digital panoramic units. Imaging Sci Dent. 43:77-84.
Lecomber AR et al. ((2000) Optimisation of patient doses in programmable dental panoramic radiography. Dentomaxillofac Radiol. 29:107-12.
Lecomber AR et al. (1998) Conference Proceeding: Dose and risk in Dental Radiography, Luxembourg 1997. Reference Doses and Quality in Medical Imaging: What the referring practitioner and directing medical staff should know. Radiat Prot Dosimetry 80:23-25.
Looe HK et al. (2007) Dose-area product measurements and determination of conversion coefficients for the estimation of effective dose in dental lateral cephalometric radiology. Radiat Prot Dosimetry. 124:181-6.
Ludlow JB et al. (2008a) Patient risk related to common dental radiographic examinations: the impact of 2007 International Commission on Radiological Protection recommendations regarding dose calculation. J Am Dent Assoc. 139:1237-43.
Ludlow JB et al. (2008b)Comparative dosimetry of dental CBCT devices and 64-slice CT for oral and maxillofacial radiology. Oral Surg Oral Med Oral Pathol Oral Radiol Endod. 106:106–14.
Ludlow JB et al. (2015) Effective dose of dental CBCT-a meta analysis of published data and additional data for nine CBCT units. Dentomaxillofac Radiol. 44:20140197.
Maillie HD & Gilda JE. (1993) Radiation-induced cancer risk in radiographic cephalometry. Oral Surg Oral Med Oral Pathol. 75:631-7.
Manousaridis G et al. (2013) Establishment of diagnostic reference levels for dental intraoral radiography. Radiat Prot Dosimetry. 156:455-7.
Manousaridis G et al. (2015) Establishment of diagnostic reference levels for dental panoramic radiography in Greece. Radiat Prot Dosimetry. 165:111-4.
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 60: References (cont.)
National Research Council of the National Academies (2006) Health Risks from Exposure to Low Levels of Ionizing Radiation: BEIR Vi. Washington, DC: The National Academies Press.
Nejaim Y et al. (2015) Efficacy of lead foil for reducing doses in the head and neck: a simulation study using digital intraoral systems. Dentomaxillofac Radiol. 44:20150065.
Pasler FA & Visser H. (1999) Zahnmedizinische Radiologies, Vol. 5. Georg Thieme, Auflage.
Pauwels R et al. and The SEDENTEXCT Project Consortium (2012) Effective dose range for dental cone beam computed tomography scanners. Eur J Radiol. 81:267–71.
Pauwels R et al. (2014a) A pragmatic approach to determine the optimal kVp in cone beam CT: balancing contrast-to-noise ratio and radiation dose. Dentomaxillofac Radiol. 43:20140059.
Pauwels R et al. (2014б) Effective radiation dose and eye lens dose in dental cone beam CT: effect of field of view and angle of rotation. Br J Radiol. 87:20130654.
Pauwels R et al. (2015) Technical aspects of dental CBCT: state of the art. Dentomaxillofac Radiol. 44:20140224.
Pauwels R et al. (2017) Determination of size-specific exposure settings in dental cone-beam CT. Eur Radiol. 27:279-285.
Poppe B et al. (2007а) Radiation exposure and dose evaluation in intraoral dental radiology. Radiat Prot Dosimetry 123:262-7.
Poppe B et al. (2007б) Dose-area product measurements in panoramic dental radiology. Radiat Prot Dosimetry 123:131-4.
Qu XM et al. (2012) Dose reduction of cone beam CT scanning for the entire oral and maxillofacial regions with thyroid collars. Dentomaxillofac Radiol. 41:373-8.
Rehnmark-Larsson S et al. (1982) Radiation absorbed doses at radiographic examination of third molars. Swed Dent J. 6:189-201.
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology

## Slide 61: References (cont.)
Rottke D et al. (2013а) Effective dose span of ten different cone beam CT devices. Dentomaxillofac Radiol. 42:20120417.
Rottke D et al. (2013б) Influence of lead apron shielding on absorbed doses from panoramic radiography. Dentomaxillofac Radiol. 42(10):20130302. doi: 10.1259/dmfr.20130302. Epub 2013 Oct 30.
Scaf G et al. (1997) Dosimetry and cost of imaging osseointegrated implants with film-based and computed tomography. Oral Surg Oral Med Oral Pathol Oral Radiol Endod. 83:41-8.
Stenström B et al. (1987) Effective dose equivalent from intraoral radiography. Swed Dent J. 11:71-7.
Tierris CE et al. (2004) Dose area product reference levels in dental panoramic radiology. Radiat Prot Dosimetry 111:283-7.
Tsiklakis K et al. (2005) Dose reduction in maxillofacial imaging using low dose Cone Beam CT. Eur J Radiol. 56:413-7.
Velders XL et al. (1991) Risk assessment from bitewing radiography. Dentomaxillofac Radiol. 20:209-13.
Visser HT et al. (2001) Dose reduction by direct-digital cephalometric radiography. Angle Orthodontist 71:159-63.
Walker C & van der Putten W. (2012) Patient dosimetry and a novel approach to establishing Diagnostic Reference Levels in dental radiology. Phys Med. 28:7-12.
White SC (1992) Assessment of radiation risk from dental radiography. Dentomaxillofac Radiol 21:118-26.
Widmann G et al. (2015) Ultralow-dose CT of the craniofacial bone for navigated surgery using adaptive statistical iterative reconstruction and model-based iterative reconstruction: 2D and 3D image quality. AJR Am J Roentgenol. 204:563-9.
Radiation Protection in Dental Radiology L11 Optimization of Protection of Patients in Dental Radiology