## Slide 1: Radiation Protection in Dental Radiology
X ray Production and Interaction
Image Formation and Image Quality
L03
Training material developed by the International Atomic Energy Agency
in collaboration with: World Health Organization, FDI World Dental Federation, International Association of Dento-Maxillofacial Radiology, International Organization for Medical Physics, and Image Gently Alliance

## Slide 2: Educational Objectives
To understand the function of the different components of the X ray tube
To understand the effect of kV, mAs and filtration on the quantity and energy of X rays
To understand the effect of geometric exposure parameters
To understand the basic principles of image formation in X ray imaging
To be familiar with the different essential image quality characteristics
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 3: Overview
Components of an X ray tube
Exposure parameters
Interactions of X rays with matter
Image formation in radiography
Image quality parameters
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 4: Overview
Components of an X ray tube
Exposure parameters
Interactions of X rays with matter
Image formation in radiography
Image quality parameters
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 5: X ray tube
All dental radiographic equipment uses an X ray tube with similar basic properties
Components of an X ray tube
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
Pauwels et al. (2015), under the British Institute of Radiology’s License to Publish

## Slide 6: Components of an X ray rube
Filament (at cathode)
Metal with high melting point
  Tungsten (symbol: W)
  (melting point 3370◦C)
Electrical current (few amperes):
    Causes heating of filament
    Electrons are released (thermionic effect)
Electrons are focused towards focal spot at anode
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
Pauwels et al. (2015), under the British Institute of Radiology’s License to Publish

## Slide 7: Components of an X ray rube
Focal spot (at anode)
Where electrons are absorbed, X rays are produced
  Bremsstrahlung (caused by deceleration of electrons)
  Characteristic radiation (caused by local ionization)
As small as possible to reduce ‘penumbra’ effect (see further)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
Pauwels et al. (2015), under the British Institute of Radiology’s License to Publish

## Slide 8: Overview
Components of an X ray tube
Exposure parameters
Interactions of X rays with matter
Image formation in radiography
Image quality parameters
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 9: Exposure parameters
Several exposure parameters are associated with the X ray tube
  Beam energy (kV, filtration)
  Tube current (mA)
  Exposure time (s)
  Geometric factors (beam size, source-skin distance, …)
Usually, both image quality and radiation dose are affected by these parameters, and their proper selection is a crucial aspect of dose optimization (see L11)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 10: Exposure parameters
X rays exiting the tube exhibit a spectrum of energies
The shape and area-under-curve of an X ray spectrum is determined by kV, total filtration, mA, exposure time and anode material
Typical X ray spectrum
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 11: Exposure parameters: kV
kV ↑ → X ray energy ↑ (both avg. and max. energy)
      X ray flux ↑
Max. energy (in keV)
= Tube voltage (in kV)
Energy levels of characteristic peaks not affected by kV (fixed for a given anode material)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 12: Exposure parameters: filtration
Attenuation of an X ray beam depends on X ray energy (see further)
  When X-ray traverse a medium, low-energy X rays are preferentially absorbed
The amount of matter that an X ray beam needs to traverse before exiting the tube is referred to as filtration
Unfiltered X ray beam
Filtered X ray beam
Filtration
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 13: Exposure parameters: filtration
Apart from the inherent filtration of the X ray tube, varying amount of filtration can be added
Mainly low-energy X rays interact with the filter material (see further).
These X rays do not contribute to the image but contribute to the radiation dose. It is therefore beneficial to filter them out.
Part of the high-energy X rays are filtered as well. Increased mAs (see next slide) is needed to compensate for this.
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 14: Exposure parameters: mAs
Tube current (mA) * exposure time (s) = mAs
(e.g. 5 mA, 10 s → 50 mAs)
mAs ↑ → X ray flux (i.e. nr. of X rays) ↑
The mAs does not affect the average and maximum energy of the X ray beam. The X ray flux (area-under-curve of the X ray spectrum) is linearly related to the mAs. As a result, radiation dose is also linearly related to mAs.
50 mAs
100 mAs
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 15: Exposure parameters: geometric
Source-skin distance
  Fixed for a given panoramic, CT or CBCT device
  Variable distances used in intra-oral & cephalometric radiography
Dose to a given point on the skin reduces with increased source-skin distance
  Inverse square relation between X ray intensity (I) and distance (D)
  But: total absorbed dose to skin not affected by SSD
Source
D
2*D
I
I/4
I/9
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 16: Exposure parameters: geometric
Penumbra
  Unsharpness at edges in the projected image
  Due to finite size of focal spot (≠ point source) + other geometric factors
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
SOD: source-object distance
ODD: object-detector distance
Pauwels et al. (2015), under the British Institute of Radiology’s License to Publish

## Slide 17: Exposure parameters
A longer SOD reduces penumbra and therefore increases image sharpness
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
Pauwels et al. (2015), under the British Institute of Radiology’s License to Publish

## Slide 18: Exposure parameters
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
On the other hand, geometric magnification is higher at a shorter SOD (but: larger detector required)
x1.3

## Slide 19: Exposure parameters
A smaller focal spot reduces penumbra and therefore increases image sharpness
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
Pauwels et al. (2015), under the British Institute of Radiology’s License to Publish

## Slide 20: Overview
Components of an X ray tube
Exposure parameters
Interactions of X rays with matter
Image formation in radiography
Image quality parameters
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 21: Interactions of X rays
When traversing a medium, X rays will interact with atoms in that medium (i.e. attenuation)
Depending on the energy of the X ray and the atomic number of the interacting atom, different effects may occur
Seibert & Boone
(2005), © by the Society of Nuclear Medicine and Molecular Imaging, Inc.
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 22: Interactions of X rays: Rayleigh scatter
Incoming X ray changes path (slightly) after interacting with electron
Does not lead to excitation/ionization
  But degrades image quality if it reaches image receptor
Mainly occurs for low-energy X ray photons
Seibert & Boone (2005), © by the Society of Nuclear Medicine and Molecular Imaging, Inc.
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 23: Interactions of X rays: absorption
“Photoelectric” absorption: energy from X ray photon transferred to electron, which is released
  Atom is ionized, photoelectron can interact further
Probability ↓↓↓ with higher X ray energy (~1/E³)
Probability ↑↑↑ with higher atomic number Z (~Z³)
Seibert & Boone (2005), © by the Society of Nuclear Medicine and Molecular Imaging, Inc.
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 24: Interactions of X rays: Compton scatter
Dominant effect at diagnostic X ray energies
Leads to scattered electron (i.e. ionization) + photon at an angle to the incident X ray photon
  Latter may degrade image quality if reaching image receptor
Probability ↓ with higher X ray energy (~1/E)
~ independent of atomic number
Seibert & Boone (2005), © by the Society of Nuclear Medicine and Molecular Imaging, Inc.
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 25: Interactions of X rays: total attenuation
Through a combination of these three interactions, an X ray beam is attenuated (i.e. reduces in intensity) when traversing a patient
In medical radiography, only a small fraction (<10%) of the X rays reaches the detector
Interaction probability for 30-120 keV X rays traversing 14 cm of PMMA. Only a small fraction of (mainly high-energy) X rays is not attenuated
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 26: Interactions of X rays
Lambert-Beer law:
  For homogeneous material:
  (N = number of X rays, µ = linear attenuation coefficient, l = thickness)
  For non-homogeneous material:
Attenuation of X rays in a homogeneous material
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 27: Overview
Components of an X ray tube
Exposure parameters
Interactions of X rays with matter
Image formation in radiography
Image quality parameters
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 28: Image formation in radiography
In radiography, image formation is done by a receptor (i.e. detector) placed at the opposite side of the X ray tube, which captures non-attenuated (primary or scattered) X rays
  Stationary tube/receptor (intra/extra-oral projectional)
  Moving tube/receptor (panoramic radiography & CT)
  Receptor = film or digital (see L04)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 29: Image formation in radiography
The receptor will capture a projection of the object under investigation
(Simplified representation)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 30: Image formation in radiography
Based on the intensity of X rays hitting the receptor at a given point, the image at that point will be:
  Darker: more X rays hitting receptor (lower attenuation)
  Brighter: less X rays hitting receptor (higher attenuation)
  This is a convention resulting from the image formation process in film radiography (see L04); in digital radiography the signal is inversely proportionate to attenuation, implying that the image is inverted at some point during processing
P.Sinpitaksakul
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
X RAY TUBE
Low signal
Low signal
High signal

## Slide 31: Overview
Components of an X ray tube
Exposure parameters
Interactions of X rays with matter
Image formation in radiography
Image quality parameters
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 32: Image quality parameters
Any radiographic image can be characterized by 5 fundamental image quality parameters:
  Spatial resolution (sharpness)
  Contrast resolution
  Noise
  Artefacts
  Geometric accuracy
(see also L10)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 33: Image quality: spatial resolution
Determined by:
  Size of image elements (i.e. pixel size, voxel size)
  Blurring (focal spot, detector, patient motion, viewing conditions, …)
Characterized as:
  Blurring of a (small) dot: point spread function (PSF)
  Amplitude vs. frequency of signal: modulation transfer function (MTF)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 34: Image quality: contrast resolution
Difference in intensity (usually: grey value) between adjacent regions in an image
Contrast between tissues on a radiograph is the result of differences in attenuation. It is dependent on:
  Differences in effective atomic number, density, thickness (‘subject contrast’)
  X ray energy (kV, filtration)
  Image receptor
  Use of contrast agents
  Viewing conditions
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 35: Image quality: noise
X ray imaging is a stochastic process
  Statistical nature, with random component being noise
  CBCT image of a large water container. In a noise-less image, each grey value inside the value would be the same
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 36: Image noise
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
Noise interferes with the ability to detect contrast
Noise is connected to exposure
(higher exposure → lower noise)
Low noise
Medium noise
High noise

## Slide 37: Image quality: artefacts
Features in the image which do not correspond with physical reality
  Imprint on film
  Metal artefacts in CT
  Motion artefacts
  …
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality
K. Horner
K. Horner
R. Pauwels

## Slide 38: Image quality: geometric accuracy
Ability of the image to represent spatial relationships
  Absolute distances (e.g. CT) or relative distances (e.g. intra-oral)
Partially determined by spatial resolution
Image can be deformed due to:
  Receptor bending
  Improper calibration
  Improper positioning
  (see modality-specific modules L05-L08)
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 39: Image quality
More information on parameters affecting image quality and measurement of image quality can be found in:
  L10. Quality Assurance in Dental Radiology
  L11. Optimization of Protection of Patients in Dental Radiology
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality

## Slide 40: References
Pauwels R, Araki K, Siewerdsen JH, Thongvigitmanee SS (2015): Technical aspects of dental CBCT: state of the art. Dentomaxillofac Radiol., 44(1):20140224.
Seibert JA, Boone JM (2005): X-ray imaging physics for nuclear medicine technologists. Part 2: X-ray interactions and image formation. J Nucl Med Technol., Mar;33(1):3-18. http://tech.snmjournals.org/content/33/1/3.long
Radiation Protection in Dental Radiology L03 X Ray Production and Interaction, Image Formation and Image Quality