## Slide 1
Presentation Part I
RADIOTHERAPY
METHODOLOGY FOR DOSE ESTIMATES IN NORMAL OPERATION.
Internacional Atomic Energy Agency

## Slide 2: OBJECTIVE
Methodology for dose estimation in normal operating teletherapy conditions.
Example of dose estimation in normal operating teletherapy conditions.

## Slide 3
To estimate doses during normal operating conditions it is necessary to identify the people potentially exposed and the exposure conditions during the daily routines.
INTRODUCTION
It is required to do estimations for:
  Exposed workers.
  Members of the public.

## Slide 4
EXPOSED WORKERS IN THE PRACTICE OF TELETHERAPY
Exposed Worker | Assigned tasks | Dose
Co-60 Unit operator technician | Operating the equipment from the control panel | Y/N
 | Positioning the patient at the equipment for treatment. | Y/N
LINAC unit operator technician | Operating the equipment from the control panel | Y/N
Medical Physicist | LINAC unit calibration | Y/N
 | Co-60 unit calibration | Y/N
 | LINAC quality control | Y/N
 | Co-60 quality control | Y/N

## Slide 5
EXPOSED WORKERS IN THE PRACTICE OF COBALT THERAPY
Exposed Worker | Assigned tasks | Dose
Co-60 Unit operator technician | Operating the equipment from the control panel | Y/N
 | Positioning the patient at the equipment for treatment. | Y/N
LINAC unit operator technician | Operating the equipment from the control panel | Y/N
Medical Physicist | LINAC unit calibration | Y/N
 | Co-60 unit calibration | Y/N
 | LINAC quality control | Y/N
 | Co-60 quality control | Y/N

## Slide 6
The estimation of the dose during the unit operation, from the control panel, depends on the control panel location:
    Control panel located on a Primary barrier.
    Control panel located on a Secondary barrier.
EXPOSED WORKERS IN THE PRACTICE OF COBALT THERAPY
Co-60 unit operator.
Primary barrier
Secondary barrier

## Slide 7
The instantaneous dose rate (IDR) at the primary barrier can be estimated by the equation:
DR0 : dose rate at the isocentre of the equipment.
B: barrier transmission factor. (*)
d: distance at the calculus point.
S: thickness of the barrier.
TVL: Tenth Value Layers.
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT OPERATOR
Control panel located on a Primary barrier
(*)

## Slide 8
The average dose rate that the operator receives in a week can be estimated from the instantaneous doses rate (IDR):
DR0: Dose rate at the isocenter of the equipment.
W: Weekly work load,
U: Use factor
T: Occupation factor.
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT OPERATOR
Control panel located on a Primary barrier
The annual dose received by a control panel operator is:
D1= Rw * Nw
Nw: Number of working weeks in a year.

## Slide 9
The instantaneous doses rate (IDR) at the secondary barrier can be estimated by the equations:
Bp y Bf: transmission factors of disperse and leak radiation
DR0 -dose rate at the isocenter of the unit
dsca -distance between the source of radiation and the patient, in meters
dsec -distance from the patient to the interest point, in meters
df –distance from the head of the unit to the interest point, in meters
α -dispersion fraction defined at the distance dsca. Look at the Bibliography. (α(30°)= 3,18E-03 (15MV)= 6,00E-03 (Co-60 1.25 MeV))
F –Radiation field area affecting the patient, in cm2. It is assumed a typical treatment radiation field size of 15x15 cm.
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT OPERATOR
2. Control panel located on a Secondary barrier

## Slide 10
The average dose rate that the operator receives in a week can be estimated from the instantaneous total doses rate (IDRT):
DR0: Dose rate at the isocenter of the equipment
W: Weekly work load
U: Use factor (for the Secondary barrier, use factor =1)
T: occupation factor.
The annual dose received by a control panel operator is:
D1= Rw * Nw Nw: Number of worked weeks in a year.
The instantaneous total dose rate is:
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT OPERATOR
2. Control panel located on a Secondary barrier

## Slide 11
EXPOSED WORKERS IN THE PRACTICE OF COBALT THERAPY
Exposed Worker | Assigned tasks | Dose
Co-60 Unit operator technician | Operating the equipment from the control panel | Y/N
 | Positioning the patient at the equipment for treatment. | Y/N
LINAC unit operator technician | Operating the equipment from the control panel | Y/N
Medical Physicist | LINAC unit calibration | Y/N
 | Co-60 unit calibration | Y/N
 | LINAC quality control | Y/N
 | Co-60 quality control | Y/N

## Slide 12
Considerations for dose estimation
According to the standard IEC 60601-2-1 it is assumed that the dose rate at 1 m from the head of the unit is 0.02 mGy/h
N0 patients a day are treated, assuming that each work shift has 2 technicians and each technician positions half of the patients.
Each patient receives an average of 3 radiation fields in each treatment sessions.
In each positioning field the technician spends 2 minutes.
Technicians works 5 days a week, 50 weeks a year.
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT OPERATOR
Positioning of the patient at the unit for treatment .

## Slide 13
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT OPERATOR
Positioning of the patient at the unit for treatment.
D2= No/2 pat/days * 3 pos/pat * 0.0333 h/pos * 0.02 mSv/h * 50 weeks/year * 5 days/week (mSv/year)
D2: Annual dose due to the positioning of the patients

## Slide 14
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT OPERATOR
Annual total dose (Dta) that the operator receives in normal operating conditions.
Dta = D1 + D2
+
D1 D2
Conclusion: Dta must be less than the dose constrain (P)

## Slide 15
EXPOSED WORKERS IN THE PRACTICE OF LINAC THERAPY
Exposed Worker | Assigned tasks | Dose
Co-60 Unit operator technician | Operating the equipment from the control panel | Y/N
 | Positioning the patient at the equipment for treatment. | Y/N
LINAC unit operator technician | Operating the equipment from the control panel | Y/N
Medical Physicist | LINAC unit calibration | Y/N
 | Co-60 unit calibration | Y/N
 | LINAC quality control | Y/N
 | Co-60 quality control | Y/N

## Slide 16
EXAMPLE OF DOSE ESTIMATIONS LINAC UNIT OPERATOR
Annual total dose (Dta) that the operator receives in normal operating conditions.
Dta = D1
+
D1 D2 = 0
Conclusion: Dta must be less than the dose constrain (P)

## Slide 17
MEMBERS OF THE PUBLIC EXPOSED IN THE PRACTICE OF CO-60 THERAPY
Members of the Public | Activity | Dose
Patient´s escort | Waiting during treatment | Y/N
 | Assistance of the elderly people and children. | Y/N
Hospital workers and other patients | Offices , bathrooms ,halls where the hospital workers are present. | Y/N

## Slide 18
EXAMPLE OF DOSE ESTIMATIONS Co-60 UNIT MEMBERS OF THE PUBLIC
Primary barrier:
Secondary barrier:
To estimate the doses received by the members of the public we use the same equations that are used to estimate the control panel operator doses. Distances, use and occupation factors should be realistic.

## Slide 19
<number>