## Slide 1
Presentation
RISK ESTIMATIONS.
“RISK MATRIX METHOD”
RADIOTHERAPY AND INDUSTRIAL GAMMAGRAPHY
International Atomic Energy Agency

## Slide 2
To show the theoretical elements that support the "Risk Matrix" method and explain the logical sequence of steps that must be followed in the practical application of this method.
OBJECTIVE

## Slide 3
Logic sequence of occurrence of accidents and its relation with risk equations.
Criteria to evaluate different variable’s levels of the risk equation.
Steps for the practical application of the "Risk Matrix" method. First screening.
Second screening procedure.
SCOPE

## Slide 4: LOGIC SEQUENCE OF ACCIDENTS
Human error or equipment failure (f)
Defenses or security barriers (p)
Accidental exposition
Consequences (C)
R = f * p * C
Notes: <number>

## Slide 5: RISK MATRIX
f
H
High
M
Medium
L
Low
VL
Very low
P
H
High
M
Medium
L
Low
VL
Very low
C
VH
Very high
H
High
M
Medium
L
Low
VL
very low
VH
Very high
VH
Very high
Notes: <number>

## Slide 6
CRITERIA FOR BUILDING A RISK MATRIX
General variable combinations logic:
The first two variables are multiplied. The result is multiplied by the third variable.
The multiplication of same level variables gives, as a result, the same level. Example: Low*Low=Low.
The multiplication of different contiguous level variables gives, as a result, the most conservative level. Example: Medium*Low= Medium.
The multiplication of different non contiguous level variables gives always two possible solutions, but the chosen variable is the one with the highest p variable. Example: Take fL*PL*CVH combination. First result: fL*PL = L. When multiplying this result with CVH, there are two intermediates, the M and the H. In this case, giving more importance to the probability level, the result would be RM.

## Slide 7: RISK MATRIX.
R
VH
H
M
R
R
L
R
Notes: <number>

## Slide 8
CRITERIA TO EVALUATE FREQUENCY LEVELS
f = pE * Nt
Every human error has its own probability (pE).This probability is a function of the human behavior. The occurrence frequency of the initiating events motivated by human errors will be expressed in events/year. It depends of the human error probability and the number of times that the activity is performed in a given year (Nt), according to the following equation:
Notes: <number>

## Slide 9
CRITERIA TO EVALUATE FREQUENCY LEVELS
Every equipment failure occurs with its own probability (n).This failure rate is a function of the characteristics of the failed component. The occurrence frequency of the initiating events motivated by equipment failure is expressed in events/ year. It depends on the failure rate and the component working time in a year (T) according the following equation:
Notes: <number>

## Slide 10
CRITERIA TO EVALUATE FREQUENCY LEVELS
f
H
High: The initiating event occurs frequently, more than 50 events /year.
M
Medium: The initiating event occurs occasionally, greater or equal than 1 and equal or less than 50 events/year.
L
Low: Unusual or rare occurrence of the initiating event, less than 1 event/year and greater or equal than 5 events per 100 years.
VL
Very Low: It is very rare that the initiating event occurs, less than 5 events per 100 years. There is no information the event ever occurred.
Notes: <number>

## Slide 11
CRITERIA TO EVALUATE LEVELS OF CONSEQUENCES
Accidents can affect patients, workers and members of the public. Any human error or equipment failure can affect at the same time to one or more of those involved in the process.
.
Patient: It has consequences for the patient because causes overdose.
Worker: It has consequences for the worker because it receives anomalous exposure.
Public: It has consequences for the public because if the source is undetected at the patient’s body , this patient goes freely home causing anomalous exposure to the members of the public.
Example: unplugging the source cable while it is inside the patient, at the end of the treatment with HDR Brachytherapy.
Notes: <number>

## Slide 12
CRITERIA TO EVALUATE LEVELS OF CONSEQUENCES
C
VH
Very high :Death or disability damage to various patients (systematic exposure). It is assumed that the magnitude of error in the dose is higher than 25%, regardless the prescribed dose.
High: Death or disability damage to one patient affected by the whole or a great part of the treatment (programmatic exposure) (the magnitude of error in the dose is higher than the prescribed dose).It also includes those expositions that affect multiple patients with dose errors between 10% and 25%, regardless the prescribed dose.
Medium: There is no risk to the patient's life. Only one of the patients treated is exposed during the session.
Low: No effects whatsoever are produced on the patients. The level of defenses has decreased.
H
M
L
FOR PATIENTS
Notes: <number>

## Slide 13
CRITERIA TO EVALUATE LEVELS OF CONSEQUENCES
C
VH
H
M
L
WORKERS AND MEMBERS OF THE PUBLIC
Very high: Serious consequences producing very severe determinists effects that might become fatal or produce permanent disability.
High: Produce determinist effects, but do not represent danger to human life and do not produce permanent damage.
Medium: Produce anomalous exposition below the determinist effects threshold. It is manifested as an increase of probability of the stochastic effects.
Low: no effects are produced on the workers or public. The level of defenses has decreased..
Notes: <number>

## Slide 14: CRITERIA TO EVALUATE THE PROBABILITY OF BARRIER FAILURE
Barrier group failure probability (p)
p = p1 * p2 * pn
p1
p2
p3
pn
Simplified method.
p1 = p2 = pn
Notes: <number>

## Slide 15
CRITERIA TO EVALUATE THE PROBABILITY OF BARRIER FAILURE.
P
H
High: most likely and expected accidental sequence (no safety barrier)
Medium: failure of defenses is accepted if the barriers are not applied correctly.(one or two barriers)
Low: there are enough defenses but it is accepted the last failure case.(three barriers)
Very low: accidental sequences virtually impossible. There are enough deepest barriers (more than three barriers)
M
L
VL

## Slide 16: STEPS FOR PRACTICAL APPLICATION
Step 1: Determination of the list of starting events
Step 2: Frequency estimation of the IE. Classification according the established levels.
One starting event is analyzed
Step 3: Evaluating consequences of the IE. Classification according the established levels.
Step 4: Analysis of the existing barriers for the IE. Diferentiate barriers, frequency and consequences reducers. Classification according the established levels.
Step 5: Obtain the risk level directly from the matrix
Analyze the following IE
Notes: <number>

## Slide 17: STEPS FOR PRACTICAL APPLICATION
  The list of initiating events (IE) can be realized by using risk analysis techniques, or
  Adapting the generic lists of IE elaborated for similar installations.
Step 1: Determination of the list of initiating events (IE)
Diagnose type of Treat. Def. of Volum. Simulation Planning Treatment Next
Notes: <number>

## Slide 18
STEPS FOR PRACTICAL APPLICATION
Step 2: Frequency estimation of the IE. Classification according the established levels.
f = pE * Nt = 0,0016 events /year (< 0,05 events/year)
Example of IE: Error in the determination of the absorbed dose in reference conditions (Telecobaltteharapy)
pE
Human error , in a non-monotonous activity, technical complex activity that is realized following procedures, activity realized in pre operational conditions ,no influenced by the pressure to deliver the treatment. A human error probability of 8.0E-03 is accepted( 8 errors per 1000 times the job is performed)
This task is performed during the initial assembly of the source and it is repeated every 5 years , when the source is changed. It is accepted a change rate of 1/5 times a year.
Nt
fVL
fL
Notes: <number>

## Slide 19
STEPS FOR PRACTICAL APPLICATION
Step 3: Evaluating consequences of the IE. Classification according the established levels.
Example of IE: Error in the determination of the absorbed dose in reference conditions (Telecobaltteharapy)
Question: What consequences can cause this IE supposing there is no barrier to avoid the accident occurrence?
Answer: It will affect multiply patients (systematic error)
The dose administrated to the patient differed more than 25% of the prescribed dose by the physician. Probably it might cause the patient death or disability damage to a lot of patients.
CVH
Notes: <number>

## Slide 20: ANALYSIS OF DEFENSES. DEFENCE IN DEEP
Human error or equipment failure (f)
Safety barriers or defenses (p)
Accidental exposure
Consequences (C)
R = f * p * C
Frequency reducers
Consequences reducers
Direct barrier
Notes: <number>

## Slide 21
STEPS FOR PRACTICAL APPLICATION
Step 4: Analysis of the existent barriers for the initiating events (IE). Differentiate barriers ,frequency and consequences reducers. Classification according the established levels.
Example of IE: Error in the determination of the absorbed dose in reference conditions (Telecobaltteharapy)
In this case the answer should be:
FREQUENCY REDUCERS
Physicists capacitation through services test.
International acknowledged protocols to do the tests.
Notes: <number>

## Slide 22
STEPS FOR PRACTICAL APPLICATION
Step 4: Analysis of the existent barriers for the IE. Differentiate barriers ,frequency and consequences reducers. Classification according the established levels.
Example of IE : Error in the determination of the absorbed dose in reference conditions (Telecobaltteharapy)
In this case the answer should be:
DIRECT BARRIERS
Redundant and independent verification of calibration results (by another Physicist and other dosimetry system).
Commissioning of the TPS. Test Case planning and comparison of results with direct measurements.
Notes: <number>

## Slide 23
STEPS FOR PRACTICAL APPLICATION
Step 4: Analysis of the existent barriers for the IE. Differentiate barriers ,frequency and consequences reducers. Classification according the established levels.
Example of IE : Error in the determination of the absorbed dose in reference conditions (Telecobaltteharapy)
In this case the answer should be:
CONSEQUENCES REDUCERS
QA of the Hospital. Monthly testing reference dose constant.
QA of the Hospital. Annual testing reference dose constant . Intercomparing exercises (OIEA- OMS).
Daily observation of the patient by the operator technician.
Weekly follow up procedures of the patient by his physician.
Periodic external audit. (Determination of the absorbed dose using reference conditions test ).
Notes: <number>

## Slide 24
APPLICATION EXAMPLE
Step 5: Obtain the risk level directly from the matrix
R
=
f
*
P
*
C
Frequency
web off
Consequences
Defences
Risk
Process stage : Acceptance and start of service
fL
CVH
Physicist capacitation
PM
?
Example of IE: Error in the determination of the absorbed dose in reference conditions (Telecobaltteharapy)
Notes: <number>

## Slide 25
USING THE RISK MATRIX TO OBTAIN THE RESULTING RISK.
Analyzing all the initiating events, a first screening can be estimated in order to establish priority according to the risk criteria.
Notes: <number>

## Slide 26
RISK MANAGMENT CRITERIA
RVH, is considered unacceptable in medical practice (Eminent risk).
RH is considered unacceptable in the long term. Necessary measures must be taken in order to reduce the risk.
Tolerable
region
Broad accepted region
Negligible risk
RVH ; RH
RM
RL
Notes: <number>

## Slide 27
RISK MANAGEMENT CRITERIA
The risk matrix method is a conservative method because in its application one assumes several conservative hypothesis, which are:
All the direct barriers have the same probability of failure and the barriers robustness is not taken into account.
The influence of the frequency reducers diminishing the IE frequency and the resulting risk has not being considered .
The influence of consequences reducers diminishing the consequences and the resulting risk has not being considered .
A second screening is justified in order to show more realistic results.
Notes: <number>

## Slide 28
SECOND SCREENING PROCEDURE
A1- Are sufficiently robust the existing barriers to assign a lower failure probability that could allow to classify risk to a lower level?
A2- Are sufficiently robust the frequency reducers or the existing consequences reducers?
A3- Is it possible to introduce new barriers, or frequency or consequences reducers?
A4- Conclusion. What additional measures can be proposed to diminish global risk?
Notes: <number>

## Slide 29
SECOND SCREENING
A1- Are sufficiently robust the existing barriers to assign a lower failure probability that could allow to classify risk to a lower level?
No | Type of Barrier | Robustness expressed in points
1 | Type 1 barriers : Interlocks | 32
2 | Type 2 barriers: Alarms | 16
3 | Type 3 barriers: work procedure performed by different persons. | 8
4 | Type 4 barriers: work procedure performed by the same person but in different stages or moments. | 4
Notes: <number>

## Slide 30
SECOND SCREENING
A1- Are sufficiently robust the existing barriers to assign a lower failure probability that allow us to classify risk to a lower level?
To failure probability pM: (2 Barriers)
A group of barriers is considerate robust if: p1*p2 ≥ 32 points. This allows to reclassify the probability from pM to pL.
A group of barriers is considerate very robust if :p1*p2 > 64 points. This allows to reclassify the probability from pM to pL.
To failure probability pL: (3 Barriers)
A group of barriers is considerate robust if :p1*p2*p3 > 64 points. This allows to reduce the probability from pL to pVL.
Notes: <number>

## Slide 31
SECOND SCREENING
A1- Are sufficiently robust the existing barriers to assign a lower failure probability that could allow to classify risk to a lower level?
No | Type of barrier | Robustness expressed in points
1 | “Redundant and independent verification of calibration results (by another Physicist and other dosimetry system)” Type 3 Barriers | 8
2 | “Commissioning of the TPS. Tests Case planning and comparison of results with direct measurements” Type 4 Barriers | 4
Example: How to evaluate the robustness of the existing barriers?
(p1*p2) = 32. Meets the criteria of two robust barrier (p1*p2 ≥ 32 points). PM PL
Notes: <number>

## Slide 32
A2- Are sufficiently robust the frequency reducers or the existing consequences reducers?
SECOND SCREENING
Overview of Frequency Reducers | Robustness | Weight
Interlocks and Technological Improvements | Very robust | 32
Signals and Alarms | Robust | 16
Protocols, procedures and moderate workload | Normal | 8
Training | Soft | 4
CRITERIA FOR ASSESSING THE ROBUSTNESS OF THE FREQUENCY REDUCER ASSEMBLY. METHODOLOGY OF THE RISK MATRIX.
Robustness of Frequency Reducers
Notes: <number>

## Slide 33
A2- Are sufficiently robust the frequency reducers or the existing consequences reducers?
SECOND SCREENING
If the multiplication of the robustness of the Frequency reducers is greater than or equal to 32 Points (RF1*RF2*RF3*…*RFn ≥ 32), It is possible to reduce a level of Frequency, ie: for example, from FH to FM.
If the multiplication of the robustness of the Frequency reducers is greater than 64 Points (RF1*RF2*RF3*…*RFn > 64), It is possible to reduce two Frequency levels, ie: for example, from FH to FL.
Note: In both cases it is not allowed to reach the very low frequency (FVL) level, in the case of events derived from human errors.
For frequency reducers
Notes: <number>

## Slide 34
A2- Are sufficiently robust the frequency reducers or the existing consequences reducers?
SECOND SCREENING
Example: how is the robustness of the frequency reducers evaluated?
There are only 2 frequency reducers. The robustness of the Frequency reducers is greater than or equal to 32 Points (RF1 * RF2 = 32), it is possible to reduce a Frequency level. It is not possible to reduce the frequency from the FL level to the FVL level if an initiating event derived from Human Errors is treated.
No | Type of Barrier | Robustness expressed in points
1 | Internationally acknowledged protocols to do the tests. Frequency Reducer Type 3 | 8
2 | Physicists capacitation. Frequency Reducer type 4 | 4
Notes: <number>

## Slide 35
A2- Are sufficiently robust the frequency reducers or the existing consequences reducers?
SECOND SCREENING
If the multiplication of the consequences reducers robustness is greater or equal to 64 Points (RC1*RC2*RC3*…*RCn > 64) it is possible to reduce consequence level, given example from CVH goes to CH.
Note: Regarding the consequences reducers, no case can be reduce medium consequences to low consequences because by definition never low consequences can be reached from medium consequences sequences.
Consequences reducers
Notes: <number>

## Slide 36
A2- Are sufficiently robust the frequency reducers or the existing consequences reducers?
SECOND SCREENING
TABLE 6. CRITERIA TO EVALUATE ROBUSTNESS OF THE CONSEQUENCES reducer GROUP .RISK MATRIX METHODOLOGY-
Robustness of the consequences reducers
GENERAL DESCRIPTION OF THE CONSEQUENCES REDUCERS | Robustness | Weight
interlocks | very Robust | 32
Alarms | Robust | 16
Protocols and procedures | Normal | 8
Emergency plans | Soft | 4
Quality controls (annual and monthly) | Theoric | 1

## Slide 37
A2- Are sufficiently robust the frequency reducers or the existing consequences reducers?
SECOND SCREENING
Example: How to evaluate the reducers robustness of consequences?
No | Type of barrier | Robustness Expressed in pointsn
1 | “weekly medical revision of the patient” consequence reducer type 3 | 8
2 | “annual and monthly quality controls” consequence reducer type 5 | 1
(RC1* RC2 = 8) The robustness of the consequences reducers is not greater than 64 points, therefore it is not possible to reduce the consequence level.
Notes: <number>

## Slide 38
A3- Is it possible to introduce new barriers, or frequency or consequences reducers?
This objective of the analysis is to propose new safety measures in order to reduce the accidental sequence risk.
The introduction of new barriers and reducers influences in the independent variables of the risk equation. To propose each one of this measures, the risk matrix reevaluates the robustness of the barriers and reducers group with the criteria exposed in questions A1 and A2
SECOND SCREENING
Notes: <number>

## Slide 39
A4- Conclusion. What additional measures can be proposed to diminish global risk?
The main objective is to propose a strategy to reduce risk in each accidental sequence .
The answer to these questions allows to define, about to which variable of the risk equation we must act, to reduce the risks to an acceptable level of safety, with the lowest cost.
SECOND SCREENING
Notes: <number>

## Slide 40
<number>