## Slide 1
LESSON 6:
WORKPLACE MONITORING
OF AIR ACITIVITY

## Slide 2
Contents
Designing the air activity monitoring programme
Airborne particulate monitoring techniques and equipment
Type tests, Calibration, verification and checking

## Slide 3
Air Activity
Air Activity can present in different forms, as:
  Particulates : dust, corrosion products, fission products, actinides…
  Vapors (HTO, ..)
  Gas : noble gases (Xe,Kr,Ar), Iodine, …

## Slide 4
Designing the Air Monitoring Programme
Define Objectives
Identify Purpose
Confirm Monitoring Type
Establish Fingerprint
The processes identified in Lessons 1&2 apply also to airborne contamination monitoring.
This lesson will therefore limit to discussion on the specific aspects of the WPM programme relevant to airborne contamination monitoring.

## Slide 5
Principal Objectives of Airbone WPM
Additional specific objectives of airborne monitoring are to:
Control the intake of radioactive substances by detecting and preventing its entry.
Prompt warning to workers to evacuate in case of loss of containment.
Select and recommend the appropriate personal protective equipment.
Define Objectives
Identify Purpose
Confirm Monitoring Type
Establish Fingerprint

## Slide 6
The Type of Measurement depends on:
  The aim of the measurement
  The radionuclide of interest
  The type or emission; ,  or 
  The physical form of the airborne contamination, e.g. Aerosols
Define Objectives
Identify Purpose
Confirm Monitoring Type
Establish Fingerprint

## Slide 7
Extent of Air Sampling
The quantity of radionuclides handled.
The limit on intake of the nuclides.
The release fraction of isotopes based on its physical and chemical form.
Type of confinement.
  The extent of air sampling is based on:

## Slide 8
Fingerprinting of Radionuclides
In order to have an effective monitoring programme for air activity contamination, it is necessary
  To know the radionuclides and their chemical and physical forms expected at the workplace.
  The probability and abundance of radionuclides.
This is called Fingerprinting. There are various methods available to determine the fingerprinting of radio-nuclides (see lesson 1)
Define Objectives
Identify Purpose
Confirm Monitoring Type
Establish Fingerprint

## Slide 9
Fingerprinting in different Facilities
Facility / Operation | Radionuclides
Mining, milling of Uranium Thorium processing Uranium enrichment Radioactive laboratories | 238U,234U, 222Rn and progeny 232Th, 228Th, 220Rn and progeny UF6 is the main hazard 125I, 32P,35S 14C, 3H
Define Objectives
Identify Purpose
Confirm Monitoring Type
Establish Fingerprint

## Slide 10
Fingerprinting in different Facilities
Facility / Operation | Radionuclides
Fuel Fabrication Reactor operation Fuel reprocessing | Natural / enriched U, Plutonium Fission products e.g. Xe and Kr, isotopes of iodine; [131Xe, 85Kr ,131I, 137Cs] Activation products e.g.3H, 41Ar, 16N, 24Na etc. Corrosion products (60Co, 108Ag, 54Mn, …) Fission products such as 95Zr, 95Nb,144Ce, 106Ru, 90Sr, 137Cs and actinides like U, Pu, Am etc.
Define Objectives
Identify Purpose
Confirm Monitoring Type
Establish Fingerprint

## Slide 11
Fingerprinting in different Facilities
Facility / Operation | Radionuclides
Radioactive Waste management Processing of radioisotopes, medical and research applications of radioisotopes Accelerators | Various fission and activation products depending on the facilities being served Depends on the radioisotopes being handled [125I, 32P, 14C, 3H] Activation products [7Be, 22Na, 26A1 , 54Mn, 59Fe, 58Co, and 60Co ]

## Slide 12
Fingerprinting
Produce a full list of radionuclides potentially present and their relative activity ratios,
Identify which radionuclides in the mix are the most dosimetrically demanding,
  i.e. which result in the highest dose

## Slide 13
Derived Air Concentration
Derived air concentration (DAC)
The concentration of airborne activity (in Bq/m3) that would result in the limit on intake of Ij,inh,L by a worker exposed continuously at that level for one year.
DAC= Limit on Intake (Bq)
2400 m3

## Slide 14
Example of DAC Calculation
DAC = I j,inh,L / (2000 * 1.2)
Assume airborne 137Cs with a 5 μm AMAD.
e(g)inh = 6.7 E-9 Sv/Bq
Annual dose limit = 20 mSv = 0.02 Sv
I j,inh,L = 0.02 / 6.7 E-9 = 3 E+6 Bq
DAC = 3E+6/ (2000*1.2) = 1.3 E+3 Bq/m3

## Slide 15
Use of DAC-h
The measured airborne activity concentration, expressed as a fraction of the DAC, multiplied by the exposure time in hours gives an estimate of intake expressed in DAC-h.
Example: 1 week at the 0.1 DAC would be 4 DAC-h, or an intake of 4/2400 = 0.002 Ij,inh,L.
2400 DAC-h corresponds to an intake of Ij,inh,L.

## Slide 16
Methods of Air Monitoring
Continuous and real time monitoring
  The continuous measurement is intended to trigger an alarm in the event of an untimely release of radioactive material in order to limit as much as possible the intake by workers and real time monitoring.
Sampling with offline laboratory analysis
  Continuous, low, medium or high volume sampling
  Spot or grab sampling
    This is used when open sources are handled occasionally. Also for specific operation that generates localised aerosol.

## Slide 17
Extent of Air Sampling
Fingerprints identify the sources of airborne activity
To determine whether sampling or continuous monitoring required, quantify the potential release in Bq/m3
Potential annual intake for a worker should be calculated based on the expected concentration, the occupancy of the area over the course of a year and the limit on intake calculated from the fingerprint at the facility or in the area
The potential intake should then be compared with the following table

## Slide 18
When do we need Airborne Monitoring?
Annual intake as a fraction of Limit on intake (ALI) | Recommendation
< 0.02 | Air sampling is generally not necessary. Routine surveys should be used to confirm contamination levels remain low.
 0.02 ALI< 1.0 | Continuous air sampling is recommended if activity concentrations may exceed 0.1 DAC averaged over 40 h or longer. Continuous monitoring recommended where activity concentrations may exceed 1 DAC averaged over 40 h
 1.0 | Continuous air monitoring with alarm capability is recommended.

## Slide 19
When do we need real time monitoring?
Real time monitoring is mandatory where there is a need to alert potentially exposed individuals to unexpected increases in airborne radioactivity levels.

## Slide 20
Examples of Air Monitoring Situations
When gaseous or volatile materials are handled in quantity
When the handling of any radioactive material in such operations results in frequent and substantial contamination of workplace.
During the processing of moderate to highly toxic radioactive materials.
During the handling of unsealed therapeutic radionuclides in hospitals.
During hot cell operations, reactor operations and handling of critical assemblies.

## Slide 21
A Monitoring Programme should specify:
The quantities to be measured.
Location and the number of points of measurement.
Measurement frequency and duration of sampling.
The lower and upper limits of detection.
The most appropriate sampling and measurement methods and procedures.
Action levels and the steps to be taken if the levels are exceeded.

## Slide 22
AIRBORNE PARTICULATE MONITORING EQUIPMENT AND TECHNIQUES

## Slide 23
Alpha and Beta Airborne
Contamination Monitoring
  Real Time and Delayed Measurement
1
  Air Sampling Circuit
2
  Aerosol Monitoring Techniques
5
  Collection media
3
  Parameters Influencing Detection Efficiency
6
  Measurement Assembly
4
  Uncertainties in the Estimation of Airborne Contamination
7

## Slide 24
Real Time and Delayed Measurement
  Monitoring continuously is known as real time monitoring.
  Either with a static filter which is regularly changed or with moving filter or tape
  In each case the filter is being measured by a detector
  Delayed measurement involves sequential sampling and analysis at a later time
  Sampling has better detection sensitivity for lower levels of airborne activity due to larger volume of air and more sophisticated forms of analysis of filter medium
The two methods can be used simultaneously in some cases

## Slide 25
Static Air Sampling (Low, Medium
or High-Volume)
Static air sampling (low, medium or high-volume)
High-volume portable
air sampler (0.5-1.0 m3/min)
High-volume portable
air sampler
Low-volume air sampler
(0.04 to 0.1 m3/min)
Notes: High-volume air sampler: flow rate is about 0.5-1.0 m3/min

## Slide 26
Continuous Airborne
Monitoring Equipment
This consists (from upstream to downstream) of:
  A sampling head
  A pipe to transport the aerosols to the collection filter
  A device to measure the pressure drop of the filter (not necessary when a mass flowmeter is used)
  A flowrate measurement device
1
2
3
4
  A sampling pump.
5
  Collection device (filter)
6
  Measurement assembly
7
Notes: Samples are counted using the same type of equipment used for wipe counting – propotional counter, scintillator with photomultiplier tube or large area silicon detectors. Will not be discussed further

## Slide 27
Particulate Airborne Monitoring Techniques
Challenges include:
  Low number of particles per cubic meter of air corresponding to DACs for some radionuclides
  Interference due to presence of natural airborne radionuclides such as radon decay products
  Finding a representative sample
Location and loss of particles from point of release to detection
    Loss in sampling lines
  Collection efficiency of filter media used

## Slide 28
Low Particulate Issues (1)
Alpha emitting radionuclides such as plutonium have a very high committed dose per unit intake by inhalation, DAC is typically 0.2 Bq/m3
Beta emitting radionuclides in comparison have a much lower committed dose per intake, DAC is typically above 100 Bq/m3
Levels of 20 to 40 Bq/m3 from radon progeny are common.
Concentration of radon progeny varies with air changes (depends on ventilation).

## Slide 29
Low Particulate Issues (2)
Solutions include:
Use of radon / thoron discrimination techniques
More frequent filter changes (at the beginning of each shift, specifically in difficult areas due to high radon and/or dusty)
Use very high air flow rates to collect more radioactivity
  Continuous air monitor flow rates are limited
  Higher flow rates result in higher levels of radon daughters
Longer count times to achieve lower LLDs
Notes: Because Pu being an alpha emitter, high levels of dust will cause self absorption.

## Slide 30
Sources of Natural Activity
Stable Daughter
210Tl
214Bi
218Rn
222Rn
238U
234Th
234Pa
234U
230Th
226Ra
218Po
214Pb
218At
214Po
210Pb
210Bi
210Po
206Pb
Natural uranium as ore in soil
Unstable parent
Gas
208Tl
212Bi
220Rn
232Th
228Ac
228Th
224Ra
216Po
212Pb
212Po
208Pb
Gas
228Ra
Stable Daughter
Thorium as ore in soil
Unstable parent
ALPHA EMISSION
BETA EMISSION
Half life: 214Pb -27 min, 214Bi-20 min, 214Po-180 sec
Notes: Half life of radon decay products (limited by Pb-214) is about 30 minutes, half life of thoron daughters (limited by Pb-212) is 10.6 hours

## Slide 31
Natural Activity Compensation
The natural radioactivity compensation is based on:
An electronic discrimination device to discriminate between the pulses generated by the detector due to radon decay products and those from a radionuclide of an artificial origin
A data processing software, which will allow compensation for the natural activity.
In case of delayed measurement of radioactivity, the activity collected on the filter is measured after the decay of radon decay products.

## Slide 32
Natural Activity Decay
Two measurements of the collection filter are done as follows:
The first measurement, after decay of the short-lived 222Rn decay products. This measurement is typically performed at least 5 hours after the sampling is completed
The second measurement after decay of the short lived 220Rn daughters. This measurement is typically performed 5 days after the sampling is completed.
Notes: A rapid estimation can also be made by performing two short measurements within a 90 minute interval corresponding to three half lives of Rn-222 decay products – this can give an indication of the likelihood that radon daughters are present

## Slide 33
Sampling Flowrate Characteristic
  Devices used to sample are usually open face filter
    Cyclones can be used to reject particles above a respirable size
  Generally the sampling flow rate is:
  chosen between 30 l/min to 100 l/min
    High flow rates will result in burst filters
    Low flow rates will reduce particulate collection efficiency
    Typically 10 liters per minute per cm2 of filter media
  measured by mean of mass flow meter or volume flow meter down stream the collection device
Notes: Human breathing rate is appr. 20 l/min
How do differences in temperature and/or pressure affect the condition – why should sampling flow rate be adjusted?

## Slide 34
Factors influencing the
Volume of Air Sampled
The air temperature
The air pressure
The two main factors influencing the sampled air volume are:

## Slide 35
Concentrations of airborne radioactive materials in a room can vary widely in space and time.
Airflow pattern studies should be conducted to identify the locations for air samplers.
General issues were discussed in Lesson 1. Also:
  Air samples should be collected in the downstream of sources.
  To check the effectiveness of containment, the air monitor should be placed near to release point.
  To measure the intake of workers, the air monitor should be placed in breathing zone.
Location of Air Monitors

## Slide 36
Location of Air Monitors
RELATIVE MERITS OF SAMPLER LOCATIONS |  | 
Sampler Location |  | 
 | Close to release point (directly downwind) | Remote to release point (in ventilation exhaust)
Dilution | Low | High
Air monitor alarm setting | Can be higher | Must be set low
Plume Concentration | High | Low
Alarm Response Time | Short (1 min) | Long (Several minutes)
Probability of plume ‘hitting’ sampler | Lower | Higher
Detection level | Low (good) | High (poor)

## Slide 37
Smoke Test to study Airflow Pattern
No ventilation
With ventilation
Migration after 120 seconds
Notes: Images show the impact of ventilation systems on the plume of activity with the ventilation system switch on and off.
Videos available to illustrated impact of ventilation and potential migration of activity.

## Slide 38
Sampling Line Efficiency (1)
If the collection device/monitor cannot be placed in the area to be monitored a sampling line may be used to transport airborne material, but:
It is impossible to transport aerosols in a sampling line without loss of particles, and it is necessary to
Know the different particle loss mechanisms and the methods of reducing the losses;
Optimize the dimensions of the sampling circuits;
Quantify the losses in order to estimate
the sampling efficiency of the device;
Notes: If the collection device or the monitor can’t be placed in the area to be monitored a sampling line may be used to transport airborne material
This sampling line should not trap or temporarily retain the airborne material sampled

## Slide 39
Sampling Line Efficiency (2)
The main parameters influencing the loss of aerosols in the sampling line are:
size of the sampled particles
velocity of the flow in the pipes
dimensions, shape and the construction of the pipes
electrical charge on the pipes and aerosols
temperature difference between the pipe’s wall and the carrier gas
humidity
Keep sampling lines as short as possible, < 2m horizontal
  If not, it is possible to model the performance of the sampling line using commercially available codes and programmes.
A rule of thumb: 50% of 10 µm particulates is lost in 1.5 m of horizontal pipe work

## Slide 40
Minimisation of Deposition
Minimize the number of bends or elbows in the system
Avoid sharp bends (90 degrees)
Optimize the pipe diameter in relation to the flow – use 20mm with 37 lpm
Use smooth pipes of constant bore
Use ball valves which do not trap particles

## Slide 41
Air Circuit
The nature of the material of the sampling line should be chosen carefully paying particular attention to chemical corrosion or memory effects. SS tubes are preferable to PVC since it avoid deposition of particulates.
The shape of the sampling line should be designed in order to reduce as much as possible the loss of particles.
The monitor should be designed to minimize the loss of particles on the walls.

## Slide 42
Air Circuit
The aerosols should enter the measurement cell from all possible directions.
Aerosol entry from one direction only can results in non uniform collection of larger size particles (>1 μ m) on the filter paper.
Non uniform collection on the filter can alter the detection efficiency.
The impaction of radioactive aerosols on the wall of the measurement cell can increase the background, and consequently increase the detection limit of the monitor.

## Slide 43
Choice of the Collection Filter
The collection filter shall be chosen taking into account of the following parameters:
Collection efficiency should be >50% for polydispersed particles with a 5 μm (ICRP 1994)
Burial of particles and its effects on counting efficiency and spectral resolution
Pressure drop:
Adapted to the expected sampling flow rate and pump characteristics
Dust and radon levels in the facility
Variation versus the mass collected adapted to the pump characteristics
Alpha spectrometry resolution for continuous monitoring using the method of discrimination

## Slide 44
Detection Limits
The detection limit is influenced by the uncertainty of the measurement which is a factor of:
  The volume of air sampled
  The count time
  The background
  The detection efficiency
Reduce uncertainties by increasing the volume of air sampled or by increasing count time, both of which will impact the filter change frequency.

## Slide 45
Calculating Airborne Radioactivity
(long-lived)
For grab sampling:
CS = RN / (V x e x SA x CE x CF)
  CS = activity concentration at end of sample run time
  RN = net counting rate
  V = sample volume
  e = detector efficiency
  SA = self-absorption factor
  CE = collection efficiency
  CF = conversion from disintegrations per unit time to activity

## Slide 46
Parameters influencing the Detection Efficiency
The self-absorption of the radiation in the particles collected on the filter
Dust build up on the air filter
  In very dusty workplaces, sampling filters may have to be changed twice a day, or sampling flowrate may have to be reduced.
  Moving filter monitors may be used where there are dust issues as the moving filter reduces the dust loading
The location of the particles deposited on the filter and therefore the detection geometry.

## Slide 47
Parameters influencing the Detection Efficiency
The deposition pattern on filter A shows a concentration of material near the center of the filter. Hence, if the detector efficiency is determined with a uniform source of radioactivity, the activity on the filter will be overestimated.
The deposition pattern on filter B shows a concentration of material near the outer edge of the filter. Hence, if the detector efficiency is determined with a uniform source of radioactivity, the activity on the filter will be underestimated.
Filter A Filter B
Notes: Can correct for these detection efficiency challenges by taking a set of samples, bulking them and sending them for radiochemical analysis. The actual concentration can be compared to the result and corrections made to the detection efficiency to account for the actual differences

## Slide 48
Alarm Levels on Monitors
Set ALARM points as low as reasonably practical to identify changes in conditions without causing excessive false radon alarms
Review past air monitor performance
Use higher alarms when respiratory protection worn
Can use shorter and longer term alarm levels (DAC-h and DAC)

## Slide 49
CALIBRATION, TESTING AND INSPECTION

## Slide 50
Type of testing of WPM Equipment
Workplace airborne monitoring equipment is required to be type tested to demonstrate its adequacy for the workplace monitoring.
Type testing of air monitoring equipment involves extensive testing to determine the sampling and measurement performances (e.g., response, linearity, energy dependence) and the effect of environmental factors such as electrical and mechanical disturbances.

## Slide 51
Type of testing of WPM Equipment
Type tests are usually undertaken by manufacturers and acknowledged independent laboratories according to the performance criteria and procedures stipulated in the national and international standards.
They are normally performed on a prototype or on an equipment taken at random from a production batch and intended to be typical of the equipment type.

## Slide 52
Calibration
Air monitoring instruments should have a valid calibration of the detection system and the flow meter before use
Calibration should be performed with traceable radioactive source (e.g. planchets or filters) having the same characteristics as the source used during type testing sample (size, radionuclide and method of construction) and in accordance with manufacturer’s instructions
For a counting system, the detector’s sensitivity for a solid source should be determined as recommended in the appropriate IEC standard
The frequency of calibration should be defined and results documented

## Slide 53
Calibration
Instrument | Calibration
Air samplers | Pump test, flow rate accuracy Leakage (pre-use testing only) low flow alarm low differential pressure alarm, high differential pressure alarm
Air monitors (Air sampler and built-in measuring unit) | As for air samplers but also: response test, background activity, alarm test, detection efficiency
Laboratory counting equipment | Response test, background detection efficiency energy response cross response (pre-use testing only), linearity of response (pre-use testing only)

## Slide 54
Functional test
Functional testing will typically include:
  Response check
    To maintain the good quality of the standard source and prevent the source from being damaged, it is recommended that the source is not used for functional testing of the instruments.
  Background check
  Flow rate check
  Alarm check
  Parameter check
Notes: All defective equipment should be taken out of service and be tagged for maintenance and recalibration
The functional test should preferable be documented