## Slide 1
LESSON 12:
WORKPLACE MONITORING IN
OPERATING NUCLEAR REACTORS

## Slide 2
Impact of Reactor Type
Impact of reactor type, age and plant state.
Fingerprinting.
Operating reactor monitoring and challenges.
Specific examples.
Does not address nuclear emergency monitoring.

## Slide 3
Impact of Reactor Type
Different reactor types may present unique challenges:
  Pressurized water reactors
    The most common type of reactors neutron flux during power operation, corrosion products, steam generator.
  Heavy water cooled reactors
    High concentrations of tritium from heavy water moderator – can be hundreds of DACs in the workplace and thousands of DACs from a small spill of moderator water.
  Boiling Water Reactors
    Contamination can be found in turbine plant components due to design.

## Slide 4
Impact of Reactor Age
Older reactors may have had historic fuel integrity challenges potentially resulting in:
Legacy of hot particles or discrete particles dispersed in the plant.
Alpha contamination embedded in primary systems from failed fuel (which may or may not be detected by wipes, and may be shielded by subsequent deposits.
Excessive contamination with radioactive corrosion products due to wrong choice of materials used in primary system components.
1
2
3
Notes: Example on wrong material choice: use of Stellite in primary circuit components which resulted in a much larger Co-60 fingerprint as well as Ag-110m fingerprint

## Slide 5
Impact of Plant Site
In light water reactors, prompt activation of oxygen in the coolant water systems results in the production of N-16, a short half life, very high energy gamma emitter which contributes to very high dose rates at power in the primary system.
In BWRs, there can be carry over of 16N to the steam side resulting in sky shine which can impact other areas.
N-16 has a very short half life and decays rapidly after shutdown reducing the external dose rate in the primary system.
Neutrons are present at power and not upon shut down.
May be increase in dose rates following crud burst following shutdown.

## Slide 6
Radiation Fields
Gamma radiation fields due to:
  Activation of reactor systems and components.
  Activation and deposition of contaminants in primary systems.
  Contamination in primary systems from fissions products.
Gamma radiation fields from a mix of radionuclides typically dominated by easily detectable radionuclides, such as;
  Corrosion products: Co-60, Mn-54.
  Fission products: Cs-137.

## Slide 7
Image from Gamma Camera
Image of dose rates using CZT
Notes: Possible equipment for gamma spectrometry:
  Portable sodium iodide or germanium detector
  Cadmium Zinc Telluride (CZT) increasing use
  Note that gamma field characteristics can change following crud burst, i.e. the introduction of oxygen into systems and subsequent clean up efforts

## Slide 8
Contamination Fingerprint
Contamination should be characterized
Characterization should use gamma spectrometry and radiochemical analysis of samples taken from throughout the plant systems to identify:
To identify beta and possible alpha emitters to ensure accounting for internal dose properly.
The radionuclide mix in typical contaminants (scaling factor methodology is often used in waste handling)
The contribution to internal dose from radionuclides not typically detected by contamination monitors such as Fe-55, C-14
The possible alpha emitters in the contamination.
Notes: The scaling factor methodology, see IAEA NW-T-1.18

## Slide 9
Typical Contaminants
Contamination could contain a range of fission and activation products.
Generally moderate energy beta emitters present and contribute most to internal dose – Sr-90, Cs-137, Co-60.
  Mix of beta/gamma emitters can be detected by:
beta/gamma dose rate meters, in case of high levels of contamination (µSv/h to mSv/h).
most beta contamination monitors, in case of moderate levels of contamination (tens to hundreds of Bq/cm²).
thin window GM tubes, in case of low levels of contamination (range of reference levels, Bq/cm²).

## Slide 10
Some Contaminants are not easy to detect
Some contamination due to low energy beta and low energy X or gamma e.g. H-3, C-14, Fe-55, typically not contributing significantly to dose
  Specific WPM is not usually necessary, once their contribution to dose has been identified.
Alpha contamination could contribute significantly to internal dose;
  Typically is limited to the primary circuits, steam generators, but also removed primary system components in storage, and waste areas.
  Determining beta/gamma to alpha ratios can be a useful method to identify radiological significance.
    At 3000:1, approximately 50% of inhalation dose will be from alpha and 50% from beta/gamma.
  The need for monitoring of alpha contamination and airborne activity is determined by the presence of alpha activity and the ratio of beta gamma to alpha activity.

## Slide 11
Specific Radionuclides
Radionuclides which might be additionally released, especially for failed fuel elements:
  Noble gases
      Krypton, Xenon.
      These short half life gases can result in widespread contamination.
  Iodine
      Wide range of short half life radioisotopes.
Tritium
  Common in CANDU reactors.
  Not usually significant quantities in other reactor types.

## Slide 12
Typical WPM in an Operating Reactor
Routine Monitoring
-Contamination, radiation dose rate for classification of areas.
-Verification of permissible levels or clearance levels.
Task Monitoring
-Monitoring to ensure doses are ALARA, either in advance of or during work.
Special Monitoring
-Investigations to identify where shielding should be placed.
-Investigations following incidents.

## Slide 13
Typical WPM in an Operating Reactor
Radiation dose rates
  Gamma, beta and neutron
Radioactive contamination
  Beta, gamma, alpha in some cases
Airborne radioactivity
  Beta airborne activity is common, with a possibility of alpha.
  Radioactive iodine isotopes.
  Noble gases.
  H-3 and C-14 (in some applications).
  Full range of monitoring requirements:

## Slide 14
Installed WPM Equipment
Range of installed gamma and neutron monitors will be provided throughout the operating reactor to:
  Identify abnormal conditions, e.g.in the reactor or lowering of water level in the spent fuel pool
  Measure ambient radiation dose rates.
Results are usually relayed to the control room
Continuous monitoring for iodine, noble gases and particulates is typical to identify abnormal conditions, e.g. failed fuel, incident .
WPM equipment should be a part of the plant design.

## Slide 15
Area WPM Equipment
Installed WPM equipment, with alarm settings, for continuous area monitoring of
  Gamma dose rate
  Beta particulate airborne activity
  In specific reactors and applications
    Airborne tritium monitors (heavy water reactors)
    Airborne alpha monitors
Portable area monitors are used for specific tasks, e.g. outage re-fuelling and maintenance work.
Notes: Area WPM for alphas is usually needed for failed fuel, the current situation and also in the plant history.

## Slide 16
WPM Equipment
Area gamma monitor
Continuous air monitor
Courtesy: Nucleonix

## Slide 17
Portable Monitors – Gamma Dose Rate
Wide range of gamma dose rates:
Dose rate monitoring for transport
Low dose rates – down to µSv/h
Dose rates during routine operations and outages, and to reduce doses ALARA.
Typically 5 µSv/h - 50 mSv/h
Very high dose rates during power operation in some non-accessible parts of the plant.
Up to several Sv/h
Need to have range capability to measure these.

## Slide 18
Other Examples of Dose Rate Monitoring
Uniform gamma fields from wide beams e.g. around reactor system components or in a waste storage facility.
Narrow beams, e.g. from shielding weaknesses.
Other sources such as waste drums and pipework (could be visible, overhead or underground).
Hot spots of contamination trapped in systems, components or on surfaces, which can be point sources.
Dose rates underwater in spent fuel pools, which may require underwater dose rate monitoring equipment.
Neutron dose rates through leakage paths at power (streaming).

## Slide 19
Contact Gamma Dose Rates
Gamma dose rates measured on closed reactor systems to identify:
Gamma dose rates in open reactor systems or inside spent fuel pools (for diving operations).
Sources of radiation to avoid exposure.
Temporary shielding requirements prior to maintenance work .
To check dose rates with shielding.
Sources to be removed, e.g. hot spots.

## Slide 20
Other Types of Contact
Dose Rate Measurement
Contact beta/gamma dose rate from open systems during outages.
Unshielded debris may have high contact dose rates (e.g. after cavity drain-down).
High beta and gamma dose rates from fuel fragments or contamination following a recent fuel failure.
Contact beta/gamma dose rates from underwater sources (debris/fuel fragments) in spent fuel pools.
Contact beta/gamma dose rates from sealed sources used for testing and calibration.
Notes: Rememeber that contact dose rates should be avoided more often than measured
Hot particles can be 10 mSv/h to 2 Sv/h beta

## Slide 21
Portable Dose Rate Measuring Equipment
Teletector-GM
Gamma identifier-Scintillator
Ionization chamber
Neutron
Gamma spectrometry kit

## Slide 22
Specific Radiation Monitoring Challenges
Remote gamma dose rate monitoring
Measuring beta dose rates
Protecting equipment from contamination
Underwater gamma dose rate monitoring
Remote Monitoring Examples
Notes: Provide some examples to clearly explain the challenges, e.g. remote gamma dose rate measurement.

## Slide 23
Contamination WPM
Beta/gamma radiation dose rate monitors for high levels of contamination.
Plastic scintillators, proportional counters and standard GM probes for moderate levels of contamination.
Proportional or scintillation counters for low levels of contamination, typically for clearance.
Equipment to detect contamination of moderate and high energy beta/gamma:
Notes: GM probes: often used due to price and durability in field, as well as for easy use.

## Slide 24
Equipment for Contamination
Alpha Scintillation counter
GM detector
Pancake detector
BF3 counter
Thermoscientific Smartlon

## Slide 25
Monitoring Challenges
Background gamma radiation in and around reactor systems inside containment can prevent the direct monitoring of beta/gamma contamination, e.g. during re-fuelling and maintenance outages.
Direct contamination monitoring can be conducted where dose rates are lower, e.g. outside of containment and for release of materials from contaminated areas.
Hot particles should be detected as soon as possible
Beta emitting particles can be easily shielded, therefore not readily detected.
1
2
3
4

## Slide 26
Indirect Contamination Monitoring Techniques
Variety of techniques are used to measure non-fixed radioactive contamination.
  Dose rate scanning for hot particles (e.g. fuel fragments or activated stellite).
  Large area wipes monitored by a beta/gamma contamination meter or dose rate meter to measure contamination levels or confirm absence of non-fixed contamination.
  Small area wipes monitored by beta/gamma contamination meter or dose rate meter to quantify contamination levels.
  Wipes may also be monitored for alpha contamination.
  Small area wipes may be measured using counting equipment if the activity is low enough to avoid contaminating the equipment.

## Slide 27
Indirect Contamination Monitoring Techniques
Large and small area wipes used on:
  Components and surfaces, especially floors:
    Before work to check the levels of non-fixed contamination.
    During the work to check the level and spread of contamination.
    After work to confirm the absence of non-fixed contamination.
  Prior to shipment, the outside surface of packages to confirm the absence of non-fixed contamination and then to quantify for transport documentation.
  Tools and equipment to confirm the absence of non-fixed contamination prior to leaving the controlled area.
Direct monitoring can be conducted wherever background radiation levels permit.

## Slide 28
Indirect Contamination Monitoring Techniques
Sticky mats/rollers can be used to capture and measure particulates.
Small items gamma monitors can be used to identify contamination on e.g. tools and equipment.
Mop wipes used for cleaning and decontamination of large areas can be measured for contamination.

## Slide 29
Examples of Indirect Contamination Monitoring
Examples of taking wipe samples
Sticky mat
Monitor to measure wipes

## Slide 30
Contamination Monitoring
Alpha contamination probes and counting equipment may be required.
  As alpha contamination measurement is not often conducted in operating nuclear plants, it can be challenging to develop and maintain expertise.
Gamma spectrometry capability for radionuclide identification of detected contaminants.
  To identify source of contamination.
  To determine dosimetry implications.
Alpha spectrometry required for identification of alpha contaminants.
  Usually not an onsite capability.
Gamma Spectrometer
Notes: To identify sources of contamination and to determine dosimetry implications, gamma spectrometry capability is needed for radionuclide identification of gamma emitters.

## Slide 31
Airborne Contamination Monitoring
Beta/gamma monitoring in air.
  Medium to high volume installed air samplers used for routine operation
  Small volume air samplers used during specific tasks and/or in specific areas (relatively high DACs easily detected).
  Air samples can be screened using portable contamination monitor for immediate indication of contamination.
  May perform particle sizing in special circumstances.
Alpha monitoring requires larger volumes of air and the use of a scaler to measure fractions of DAC values.
Radon daughter interference was discussed in lesson 6
    May use discriminating counters.

## Slide 32
Exampled of Particulate Air Sample Monitoring
Portable particulate monitor
Counting equipment examples
Air Sample papers – clean and used

## Slide 33
Airborne Contamination
Recent fuel failures may give rise to the release of noble gases and iodine radioisotopes upon shutdown (and which require additional workplace monitoring including:
  additional monitoring of widespread contamination from noble gases.
  iodine monitoring.
  alpha monitoring (note this could also remain embedded in systems in future).
Notes:   Note regarding alpha monitoring: although the release may not immediately impact beta/gamma to alpha ratios due to large beta/gamma component.
  Additionally: fuel failure will result in greater contamination, such as fuel fragments, discrete radioactive particles and potentially higher dose rates, so enhanced monitoring is required

## Slide 34
Airborne Tritium Monitoring
Tritium can easily spread and disperse from irradiated heavy water spills and wetted materials.
  Portable H-3 detectors and passive samplers are used during work with liquids to identify tritium in the workplace.
  Installed H-3 monitors are used to measure ambient H-3 levels and detect abnormal conditions.
Installed H-3 monitor
Hand held H-3 monitors

## Slide 35
Summary
Impact of reactor age, type, plant condition.
Fingerprint (hard to detects, alpha, specific radionuclides).
Installed gamma and airborne monitors.
Range of portable monitoring equipment for radiation dose rates, contamination, air activity .
  challenges include background radiation, hot particle, shielded beta particles, remote monitoring.
Gamma spectrometry widely used.
Beta gamma airborne monitoring.
Noble gases, iodine, alpha and H-3 when required.