## Slide 1
LESSON 5:
SURFACE CONTAMINATION MEASUREMENT

## Slide 2
Outline
Calibration and Testing
Practical Measurement
Specific emissions and interpretation
Instruments
Introduction
Technique
Direct Monitoring
Indirect Monitoring

## Slide 3
INTRODUCTION

## Slide 4
Definitions and Types
Contamination:
Presence of radioactive substances in a volume or on a surface where they are unwanted and undesired (air, water, internally in the body, etc.)
Transferable, non-fixed or loose contamination: surface contamination that can be easily removed
Non-transferable or fixed contamination: surface contamination that cannot be easily removed

## Slide 5
What causes surface contamination?
Spill/leak of radioactive material
Deposition of airborne radioactive substances on surfaces
Contact of a contaminated object (e.g. hand, shoes, clothes) with the surface

## Slide 6
Need for Surface Contamination Monitoring
Occupational exposures may result from loose contamination:
Transfer of surface contamination to the atmosphere leading to internal exposure through inhalation or ingestion
External exposure to whole body, eyes or extremities if contamination is significant
Measurement of loose contamination is very important
Fixed surface contamination may cause:
Potential for future loose contamination
External exposure if contamination is significant.

## Slide 7
Purpose
Verification that levels of contamination are well within the permissible levels.
  0.4 Bq/cm2 for alpha emitters and 4 Bq/cm2 for beta/gamma emitters (IAEA guidelines*).
Classification of controlled and supervised areas
Provision of assistance with methods to determine and prevent the spread of contamination.
To detect breach in containments and departures of good working practices.
To provide relevant inputs to decide on individual monitoring or dose assessment
*Regulations for the. Safe Transport of. Radioactive Material, IAEA,TS-R-1

## Slide 8
TECHNIQUE

## Slide 9
Methods of Contamination Monitoring
Direct method
detecting contamination directly using instruments
used to detect both transferable (fixed) and non-transferable (loose).
Indirect method
Using dry or wet wipes
Used to detect transferable (loose) contamination and by calculation the total contamination.
Detecting contamination on wipes using instruments

## Slide 10
Direct Method
Measurements normally performed on a ratemeter basis
  Surveyor scans the instrument over area of interest
  Identifies contaminated areas and averages the indication
  Use of audio output is useful
Scaler readings increasingly being used to detect lower levels and make measurements more auditable and repeatable
  Place instrument and average or integrate over a short time, e.g. 10 seconds, at various locations.

## Slide 11
Direct Method
Points to consider:
Nature and location of the contamination
The average area to be covered
The detector to surface distance during measurements
The speed of monitoring.

## Slide 12
Direct Method-Nature and Location
Presence of water/grease/
dirt/paint/polish on the surface (consider self absorption)
Radio nuclides (Type and energy of radiations)
Homogeneity of potential contamination
Other radiation sources in the vicinity.
There may be a need to remove sources or shield an area before the contamination monitoring is started.
Nature and Location of Contamination

## Slide 13
Direct Method-Area of Measurement
Area of measurement:
  The sensitive area of the detector defines the minimum averaging area.
  For surfaces such as floor, wall etc, contamination should be averaged over an area of 1000 cm2 or less
    Regulation may dictate the averaging area.

## Slide 14
Direct Method-Detector to Surface Distance
Detector to surface distance
Alpha and low energy beta emitters will only be detected if the probe to surface distance is not greater than a few millimetres due to short range of these particles.
Never allow the probe to touch the surface as this might contaminate the probe. (e-g : Use a spacer)

## Slide 15
Direct Method- Speed of Monitoring
The speed of monitoring
The speed at which a probe can be moved over a surface to ensure reliable assessment of contamination levels, needs to be carefully assessed prior to any survey.
Generally one detector dimension per second is recommended from practical experience.
Monitoring speed depends on the required detection level.

## Slide 16
Example of Calculation of Surface Contamination-Direct Method
Example: Counting efficiency= 5%, Meter window area = 50 cm2, Counts=75 cpm, Background counts= 20 cpm

## Slide 17
Indirect Method
Wipe method is used for estimation of non-fixed or transferable (loose) contamination
Either dry or wet wipe method is employed
  Dry wipes identify the potential airborne hazard and PPE
  Collection efficiency could be improved by wetting the swipe materials in water or alcohol
    Never count a wet wipe for alpha
Wiped samples may be analysed using a contamination monitor in the workplace or may be analysed in a laboratory.

## Slide 18
Indirect Method-Wipe Area
Typical wipes used the area in the range of 100 cm2 (for quantification) to larger (for identification)
For localized contamination use small area wipes to prevent contamination spread.
Sub-divide large areas into smaller areas and wipe each one.
For objects wipe each one or identifiable parts.
In practice may use a combination of both methods.
The wipe area may be determined by regulation
Otherwise the qualified expert should define the area

## Slide 19
Wipe Materials
Many types of swipe materials are used including paper, styrofoam, cotton swabs, cloth patches, and glass fibre filters.
Large area wipes usually use cloth.
Adhesive tape can be used to collect particulates for monitoring/analysis.
Solvents include methanol or distilled water (the swipe must be dry before counting).

## Slide 20
Steps in Wipe Procedure
Prepare labeled vials or envelopes as appropriate
  Small boxes with separations, plastic bags or envelopes can be used
  Prevent cross contamination between wipes
Wearing gloves, use filter paper/cotton swab to take several samples from different areas.
  Rough surfaces should generally not to be wipe tested
Wipe should be rubbed with uniform pressure over the area of interest.
  Use tongs if high levels of activity or dose rate
Place samples in separate vial or envelope
Note the location of the sample as appropriate.

## Slide 21
Wipe Counting
Analysis method depends on contamination level
  For high activity, perform direct monitoring using an instrument for an adequate time and record the result
  For lower activities and low energy beta or gamma emitters, count in a laboratory counter or shielded counter in a low dose rate area
    Handle the wipes with tweezers
    Wipes are typically counted for one minute, depending on the sensitivity required
    Ensure background and a blank wipe are counted
    If wipe requires further analysis, do not use liquid scintillation counting
    Retain wipe for radiochemical analysis if required.

## Slide 22
Recording of Surface Contamination Results
It is not only good practice but also a legal requirement that contamination monitoring results are recorded. This contains:
Manufacturer, model and serial number of the contamination meter used
Background measurement
  A basic description of each location monitored
Monitoring readings in Bq/cm2 or cps / cpm (depending the local rules and regulator request).

## Slide 23
Recording of Contamination Monitoring –Area Wise
1,2,-Lab bench top, 3-Centrifuge, 4-sink, 5-workfloor,6-Refrigerator shelves,
7-Floor Refrigerator,8-Floor Sink,9-Floor bench,10-Inside door handle

## Slide 24
Factors that influence Wipe Efficiency
The wipe efficiency or pick up factor is the percentage of contamination picked up by a single swipe of the surface. It is affected by
  type of wipe used.
  pressure applied by the person when making the wipe.
  contamination distribution.
  porosity, chemical composition, texture and cleanliness of the surface.
Traditionally a pick up factor value of 10% is used.
Self absorption by the wipe (usually: 25 % for alpha and 5% for beta).

## Slide 25
Wipe Testing Advantages
Prolonged measurement in laboratory can provide a lower detection limit than direct method.
It is possible to make a spectrometry to identify the radionuclides.
Can be used in high background areas or where physical geometry and accessibility is a constraint.
For some low energy beta emitting radionuclides (Tritium for example), this is the only way you can for routine measurement of surface contamination.

## Slide 26
Wipe Testing - Disadvantages
Does not measure fixed contamination
Error associated with wipes is high . Using pick up factor can induce over or under estimation of total contamination.
Wiping might miss the area of contamination entirely or might not measure specific contaminated spots
The qualified expert should always interpret the results and identify if further monitoring is required

## Slide 27
INSTRUMENTS

## Slide 28
Choosing the Equipment
Selection of monitor according to the type and the energy of radiation
A full list of nuclides which could be encountered must be made and the decay schemes studied
The calibration of the monitor will also require knowing the radiation type and energy to determine efficiency
In practice, it is common to require a range of contamination monitoring equipment

## Slide 29
Conversion Factor Approach
The approach used here is as follows:
As the equipment usually does not display the radiation protection quantity to be measured a conversion from cps to Bq/cm2 is necessary
57.8 cps
Bq/cm2

## Slide 30
Conversion from cps to Bq/cm2
Conversion from cps to Bq/cm2 – i.e: Detection efficiency, area are to be taken into account:
Type of radiation (alpha / beta / X ray and gamma)
Energy of radiation
Window size
Background

## Slide 31
Conversion Factor
Therefore, determination of a conversion factor for a contamination monitor;
Needs to be made with a with reference sources of different radiation types and qualities compatible with those that will be measured with the equipment
Conversion factor(Nuclide) =
cps
Bq/cm2
Notes: Typical nuclides used for determination of conversion factors:
E (alpha) or Emax (beta)
14C 156 keV β
36Cl 709 keV β
90Sr+90Y 2274 keV β
241Am 5.48 MeV α, γ
239Pu 5.16 MeV α, γ
238U 4.20 MeV α, γ

## Slide 32
Determining Conversion Factors in Practice
How does it work in practice?
Reference source
Spacer
Cont.
meter
Monitoring area

## Slide 33
Determining Conversion Factors in Practice
Measure the response to the source in cps
Reference source will ideally be a large source with units of Bq/cm2 so express conversion factor directly.
If it is in other units, then estimate conversion to Bq/cm2
  Bq (divide by monitoring area)
  s-1/cm2 (multiply by two assuming is an ideal source)
  s-1 (multiply by two and divide by monitoring area)
Express conversion factor in cps/Bq/cm2

## Slide 34
Example
Calculate the response of a pancake probe where the probe response to a contamination source is 100 cps to a 1kBq source (probe area is 15 cm2)
  Source strength = 1000 Bq
  Dividing by monitoring area = 1000Bq/15cm2 or 66 Bq/cm2
  Conversion factor = 100 cps per 66 Bq/cm2
  or 1.5 cps/Bq/cm2

## Slide 35
Calculation based on loose Surface Contamination
Surface contamination = (N - Nb) / (E x 60x A x F) (Bq/cm2), where:
N = is the total count rate in counts per minute (cpm) measured directly or on the wipe.
Nb = is the count rate of the blank (in cpm)
E = instrument efficiency for specific isotope (e.g. for 26% efficiency, E = 0.26)
60 = sec/min
A = area wiped in cm2 (not to exceed 100 cm2) or area of detector (direct measurement)
F = is the pick-up factor for the wipe. (e.g. a value of F = 0.1; i.e., 10%).

## Slide 36
Frequency of checking the Conversion Factors
How often shall I check my equipment?
Standards, Tecdoc or national standards specify the calibration intervals
  Usually every 1-2 years, the test is performed using various reference sources of the radionuclide source of interest (e.g. planchets or planar sources) with a traceable surface emission rates of different activities at the same distance as contamination measurement. For counting equipment, use the same size source and counting geometry as the smear
  The detector’s efficiency should be determined as recommended in the appropriate ISO standard.
To ensure reliability and to verify stability of the equipment, a check source should be used prior to every day of measurement (function testing)

## Slide 37
CALIBRATION AND TESTING

## Slide 38
Example
The calibration time interval should be defined in the quality management system documentation
The results of calibration should be normally noted on the data sheet of the equipment or documented in another way
A daily check source need not be a traceable standard – typically just looking for a response.
To maintain the good quality of the reference source and prevent the source from being damaged, it is recommended that the traceable source is not used for functional testing of the instruments.
Probes should not be routinely exchanged between rate meters and the equipment should always be tested as a full assembly
Notes: In practice the calibration is performed on the measuring equipment separately, using traceable radioactive sources (e.g. planchets or planar sources)
The sensitivity of the detector for a solid source should be determined during type testing as recommended in the appropriate ISO standard.
This method requires the use of a reference source of a right size with a traceable surface emission rate to the radionuclide of interest
Some countries permits only laboratories authorized for calibration to use the word calibration.
The reference sources are sold as an accessory.

## Slide 39
Function Testing – Surface Contamination Equipment
Check the physical integrity of the detector, especially the window, cable and probe
For proportional counters, check gas pressure or re-fill the chamber
Check the calibration validity
Check the battery
Check the expected background reading
Check the alarm, if appropriate
Place the detector on a low activity alpha or beta check source with the source usually in the center of the window. Check whether the reading is within the expected range – for a scintillation probe, check against light to detect leaks
Notes: Check sources can be purchased as accessories.
The functional test may be written in the form of a check-list.

## Slide 40
PRACTICAL MEASUREMENT

## Slide 41
Surface Contamination Measurement(1)
Before starting a measurement:
1. Ensure to select the correct equipment (radiation type, energy and surface activity range)
2. Perform a functional test
3. Use the audio signal. Set the alarm signals at a known surface contamination level. Be careful in noisy workplaces.

## Slide 42
Surface Contamination Measurement(2)
Follow the procedure for the specific type for the measurement (more to follow later)
When an unusually high or low measurement result is indicated, act immediately on the information
Allow an appropriate time for the equipment reading to stabilize.
Before starting a measurement:

## Slide 43
Surface Contamination Measurement(3)
Performing the measurement (cont´d)
Keep the thin window away from sharp objects to avoid irreparable damage to the detector.
Turn the equipment off when not in use
If the equipment has a separate probe, do not hold the probe by the cable.

## Slide 44
SPECIFIC CONSIDERATIONS FOR ALPHA EMITTERS

## Slide 45
General Aspects: Alpha Surface
Contamination Measurements
Alpha emitters have higher inhalation dose coefficients (radio-toxicity) compared to beta and photon emitters
This leads to a lower maximum acceptable surface contamination level, usually 10 times lower than beta contamination (and not more than 0.4 Bq/cm²)
Any alpha detected is significant.

## Slide 46
General Aspects: Alpha Surface
Contamination Measurements
Measurements are difficult due to the very short range of alpha radiation in air and by absorption on the surface itself
  This makes the alphas very susceptible to attenuation in layers of grease or dust on the surface and also demands that any probe is held not more than a few mm from the surface
In practice, direct alpha monitoring is only possible on smooth, clean, non-porous, relatively flat surfaces
  This means many objects cannot in effect be monitored
Monitoring speed for alpha is typically slow
Generally, it is advisable to use equipment with the largest practical window size
  Increasing window area can increase speed of monitoring

## Slide 47
Equipment for Alpha Measurements (1)
Two types of detectors are commonly employed:
Zinc sulphide scintillation detector, either as a pure alpha detector or deposited on a plastic scintillator plate as a dual alpha and beta detector
Thin windowed refillable proportional counter
Large area at least 100 cm2 for clearance.

## Slide 48
Equipment for Alpha Measurements (2)
Scintillation detector for alpha
Scintillation detector for alpha
Notes: Explain the photmutlpier tube

## Slide 49
Equipment for Alpha Measurements (3)
Butane gas proportional chamber for alphas.
Notes: The problem is that the butane equipment needs high quality butane or argon-methane gas to work. Which may not be widely available.

## Slide 50
Equipment for Alpha Measurements (4)
Pancake halogen quenched GM tube. Window = 12 cm2
Efficiency = 15 % for 239Pu

## Slide 51
Procedure for Alpha Surface
Contamination Measurements (1)
Scanning in rate mode,
The rate of movement for low level measurements should normally be a few centimeters per second with a 100 cm2 probe
The surface scanning should be made using the audio output. When the contaminated area has been found, then the instrument should be held still and the reading should be used to estimate the surface contamination.
Static integrated mode: put the probe over the surface for a fixed time, and record the result.
Two methods can be used

## Slide 52
Procedure for Alpha Surface
Contamination Measurements (2)
Be careful to maintain a constant distance of a few mm between the detector window and the surface to be measured
To maintain the constant distance, it may be helpful to stick a disposable plastic or rubber spacer on the detector frame (not on the window). The spacer may be placed directly on the surface and replaced when necessary if contaminated
Notes: A few mm in practice means between 2 and 5 mm!

## Slide 53
Procedure for Alpha Surface
Contamination Measurements (3)
The equipment used for alpha surface contamination measurements shows a very low gamma background rate, usually around 0.1 cps for a 100 cm2 probe
If you touch the contaminated surface, you may contaminate the window or probe, and low levels of window contamination will produce a high background rate

## Slide 54
Procedure for Alpha Surface
Contamination Measurements (4)
A disadvantage of large area detectors, other than the weight and size, is if the activity is in the form of discrete hot spot which would be averaged over the area of the detector and the spot could be high.
The large window will also not allow the exact location of a small spot of contamination.
Floor monitors can use many detectors, therefore increasing sensitivity.
If necessary, a detector with a smaller window can be used to detect the spot with higher sensitivity.

## Slide 55
Procedure for Alpha Surface
Contamination Measurements (5)
Interpretation of the measurement result is dependent on the type of surface
Smooth surfaces, such as stainless steel, will have a high efficiency whilst surfaces such as unpainted wood, will have a poor efficiency or may hide contamination.
A layer only 2 mg/cm2 thick will fully absorb the alpha particles and prevent monitoring (for water, oil and grease, this corresponds to a thickness of 20 micrometre)
It is important to note the condition of surfaces monitored if there may be any concern over the efficiency of detection and this should be considered during interpretation.
Notes: Interpretation is dominated by the need to make some guess at the efficiency for real surface contamination. For good surfaces, such as stainless steel and hard plastic faced boards in really good condition, then if the surfaces look completely clean then the source efficiency can be quite high and using the measured calibration factor will normally give a reasonable answer. For poor surfaces, such as unpainted wood, carpets and obviously greasy surfaces, direct monitoring can dramatically underestimate the available activity.
ISO Publication 7503-1 [20] recommends assuming that only half the number of particles which set off towards the accessible surface of the source actually escape for alpha and lower energy beta particles. However, what must be kept in mind is that a layer only 2 mg.cm-2 thick over an alpha contaminated surface prevents its monitoring. For water, oil and grease, this corresponds to a thickness of 20 micrometre or one fiftieth of a millimetre. For uranium particles, this would correspond to a radius of closer to 1 micrometre, to absorb completely alpha particles generated at the centre.
Hence, when alpha monitoring, all that a surveyor can say is that the activity is greater than or equal to the value derived from calibration.
Uncertainties are dominated by the difficulty of estimating the efficiency of the real activity. The next most important contribution is often the difference in energy between the calibration source (5.4 MeV) and the contaminant (4 to 8 MeV), particularly when monitoring for U-238 at 4.2 MeV and the third major uncertainty is the detector to surface distance.

## Slide 56
SPECIFIC CONSIDERATIONS FOR BETA EMITTERS

## Slide 57
Plot of Efficiency vs Energy
Gamma energy (keV)

## Slide 58
Beta Surface Contamination Measurements (1)
Most beta surface contamination measuring equipment detects betas with Emax above around 150 keV
Low energy beta radiation needs to be monitored at a few mm from the surface to detect the contamination
Below this energy, the range of the beta becomes small, not enough to cross a few mm air gap and penetrate the window of a practical detector
Counting equipment can be used to detect lower energy betas, e.g. Fe-55, Ni-63, H-3 (liquid scintillation or proportional counters – see lesson 7)
Windowless gas flow proportional instruments are also available on the market, although these are difficult to use in practice.
Notes: The section covers beta emitting nuclides with energies in excess of 63Ni, Emax = 66 keV. This nuclide represents the lower limit where direct monitoring with a windowed detector is possible although many practitioners consider 14C (0.16 MeV) to be the lowest energy directly monitorable nuclide.

## Slide 59
Beta Surface Contamination Measurements (2)
For higher energy beta radiation (e.g. Sr-90, P-32) the probe can be held up to 1cm from the surface being monitored with little impact on count rate
The beta surface contamination measuring equipment will show a significant background count from artificial or natural gamma radiation.
This is around 5 cps per 100 cm2 or 1 cps per 20 cm2 of probe window.

## Slide 60
Beta Surface Contamination Measurements (3)
The authorized level of beta contamination for non-radiologically controlled area is normally higher than for alpha contamination, typically 4 Bq/cm2.
  This level can usually only be detected in low gamma background levels
    Higher contamination levels can be detected
  Monitoring for these levels should be conducted in low background or wipes should be taken and counted in a low background area.

## Slide 61
Equipment for Beta Measurements (1)
Many types of instrument can be used for direct beta contamination measurements:
Thin glass or metal walled GM detectors (Emax > 0.5 MeV)
Gas refillable proportional counters (as for alpha)
Thin window xenon filled sealed proportional counters
Thin windowed scintillation detectors (as for alpha)
Thin end window GM detectors

## Slide 62
Equipment for Beta Measurements (2)
Pancake detector
Contamination monitor

## Slide 63
Interpretation of the Measurement Results
1.
Low energy beta measurement experience similar interpretation issues as alpha.
2.
There is only a limited range of nuclides which will be available during calibration and type testing. It is often necessary to estimate the response for the actual nuclides encountered.
Use a conservative conversion factor if a mix of radionuclides
Notes: There is only a very limited range of nuclides which will be available during calibration and type testing. Sometimes there are good analogues, in the sense of a nuclide for which there is type test data which has a very similar energy to that involved in the real monitoring. For example, 14C and 35S are similar, as are 32P and 90Sr + 90Y. When there are no suitable analogues, then it will be necessary to estimate responses to the nuclides encountered in practice. Sometimes manufacturers can provide data for a wider range of nuclides than is included in the type test. Otherwise, the simplest solution is to graph the response to the type test nuclides against energy, making allowance for the decay probability. An estimate of the response to the practical nuclides can then be made.

## Slide 64
SPECIFIC CONSIDERATIONS FOR PHOTTON EMITTERS

## Slide 65
X Ray and Gamma Measurements(1)
Many nuclides produce low energy X rays, typically by a process of electron capture.
125I emits two X-rays at 27 keV and 32 keV and a gamma at 35 keV
X ray emitting nuclides can also be found in contamination around reactors.
The most prominent, low energy photon emitting nuclide in nuclear plants is 55Fe, which emits X ray at 5.19 keV.
Notes: Electron capture is a process in which a proton-rich nuclide absorbs an inner atomic electron, thereby changing a nuclear proton to a neutron and simultaneously causing the emission of an electron.
Electron capture is the primary decay mode for isotopes with a relative superabundance of protons in the nucleus, but with insufficient energy difference between the isotope and its prospective daughter for the nuclide to decay by emitting a positron. Electron capture is an alternate decay mode for radioactive isotopes with insufficient energy to decay by positron emission. It is sometimes called inverse beta decay.

## Slide 66
X Ray and Gamma Measurements(2)
Most gamma emitters also emit beta radiation and, generally, it is the beta radiation which is easiest to detect on surface and is used in contamination measurements
  Typical examples are 137Cs and 60Co
Pure Gamma/X ray emitters will not be detected by beta probes, but can also be assessed using scaling factor based on the fingerprints if there is a mix of radionuclides

## Slide 67
X Ray and Gamma Measurements(3)
There are two conflicting issues:
In practice this means that the same probe cannot be used for measuring both low and high energy gamma/X ray emissions.
For low energies, the detector window has to be relatively thin and have a low atomic number to ensure good transmission
For high energies, the problem is to ensure that the relatively penetrating radiation is attenuated sufficiently by the detector. This demands a high mass per unit area, high atomic number detector
Notes: If have both, it is better to choose the low energy option – will still see the contamination

## Slide 68
Equipment for X Ray and Gamma Measurements (1)
There are two popular forms of detector:
One is the xenon filled, titanium windowed, proportional counter, more useful for lower energies
The other is the aluminum or beryllium windowed sodium iodide scintillation detector which is useful for a wider range of energies
Xenon filled proportional counter
NaI detector with window
Notes: The gas filled detector is usually cheaper, however the large window facilitates the escape of the counting gas, which means that the detector has to be periodically returned to the manufacturer for re-filling.

## Slide 69
Equipment for X Ray and Gamma Measurements (2)
The sodium iodide scintillation detectors are usually less sensitive than the xenon filled proportional counters
Since the probes are designed to detect gamma contamination, they are very sensitive to background radiation and may have a significant background count rate
Most instruments employed for low energy photon measurements will respond to beta emitters. If in doubt, shield the betas with a filter
Photomultiplier tubes are susceptible to magnetic fields, they are often supplied with mu-metal shields to overcome this issue.
Notes: One aspect of caution which applies to all scintillation detectors is their use in magnetic fields. Photomultiplier tubes are susceptible to magnetic fields, which reduce the gain of the detector, effectively increasing the energy threshold. Some types are provided with a mu-metal shield, which acts as a magnetic short circuit around the detector, and considerably reduces the influence of the magnetic field. These will often perform correctly even in obvious magnetic fields. Other types which do not have this shielding may well be significantly affected by both deliberate magnetic fields (such as some types of tool) and accidental magnetic fields (such as structural steel work or rail wagons). In case of doubt, a low energy source is taped to the detector and checks made to insure that the count rate does not change dramatically when moved over the object of interest.

## Slide 70
Procedure for X Ray and Gamma Surface Contamination Measurements
A distance of a few cm from the surface is adequate for scanning for higher energies
Need to be closer to the surface for low energies
The audio output should be listened to locate the contamination. Then the operator should move the probe to maximize the count rate.
In high background areas, perform wipe test and measure the sample in a low background area.
Notes: Even low energy (5.9 keV) X-radiation has a much longer range than an alpha or low energy beta radiation. Hence the need to hold the detector very close to a surface is reduced. In a similar way to high energy beta radiation, drawing back a few millimetres from a surface effectively increases the averaging area, reducing the count rate from point source contamination but having little effect on distributed contamination. Monitoring is thus generally similar to dealing with high energy beta radiation. The probe should be scanned over the surface in question while the operator listens to the audio output. If an increase is noted or if a trigger point is exceeded then the operator should move the probe to maximize the count rate, and either record the value and position or clean up the contamination immediately.

## Slide 71
1
The main issue is calculation of the response to a particular nuclide due to complicated decay schemes.
2
There is only a limited range of nuclides which will be available during calibration and type testing. It is often necessary to estimate the response for the actual nuclides encountered.
3
Interpreting gamma contamination in the presence of a gamma background is a significant issue
The rationale for the estimated response should be recorded
Interpretation of the Measurement Results

## Slide 72
Counting Equipment
Alpha and Beta
  Large area silicon diode connected to a scalar timer
  Gas flow proportional counter
  Plastic scintillator with ZnS layer
Beta only
  Thin end window pancake GM
  Plastic scintillator
Alpha only
  ZnS(Ag) & PIPS
X ray & Gamma
  NaI(Tl) and HPGe