## Slide 1
LESSON 9:
MEASUREMENT AND UNCERTAINTY

## Slide 2
CONTENT
Introduction
Measurements
Uncertainty and errors
Standard deviation and its calculation
Sources of errors and types of errors
Detection limits and Minimum Detectable Activity

## Slide 3
Objectives
Provide a brief overview of basic concepts on measurements and uncertainty.
This lesson focuses on the measurement methods in counting the radioactive samples.
To provide the insights into standard deviation, detection limit and MDA.

## Slide 4
Counting Statistics
Radioactive decay is a random process. Consequently, any measurement based on observing the radiation emitted in nuclear decay is subject to statistical fluctuation
Counting statistics is also used to make a prediction about the expected uncertainty of these measurements.
The term counting statistics includes the framework of statistical analysis required to understand and interpret the results of a nuclear counting experiment.
These inherent fluctuations represent an unavoidable source of uncertainty in all nuclear measurements.

## Slide 5
Measurement
  What is a measurement?
  A measurement tells us about a property of something. It might tell us how heavy an object is, or how hot, or how long it is. A measurement gives a number to that property.
  What is not a measurement?
There are some processes that might seem to be measurements, but are not.
For example, comparing two pieces of string to see which is longer is not really a measurement.
Notes: Radioactive decay is a random process. Suppose that in a particular experiment, N counts are recorded in a time T. Then the counting rate, for that trial, is R = N/T. If the average counting rate is R (the bar means that the average is taken over many trials), then any particular trial will result in a measured R which will probably be close to, but not exactly equal to R . A very similar situation occurs if you flip a coin many times: the number of heads will be close to, but seldom exactly equal to, the number of tails. In such cases, the counting statistics have a simple rule, called the √N rule: If the number of counts recorded in a time T is √N, then the uncertainty in the number of counts is N . That is, if you repeat the experiment many times, you will get a gaussian distribution of N’s, centered on N , the average value of N, , with a standard deviation of √N .

## Slide 6
Uncertainty of Measurement
It tells something about its quality.
Uncertainty of measurement is the doubt that exists about the result of any measurement.
Expressing uncertainty of measurement
Two numbers are really needed in order to quantify an uncertainty. One is the width of the margin, or interval. The other is a confidence level, and states how sure we are that the ‘true value’ is within that margin.

## Slide 7
Error versus Uncertainty
Error is the difference between the ‘measured value’ and the ‘true value’.
Uncertainty is a quantification of the doubt about the measurement result.
Whenever possible we try to correct for any known errors: for example, by applying corrections from calibration certificates. But any error whose value we do not know is a source of uncertainty.

## Slide 8
Why is the Uncertainty of Measurement Important?
You may be interested in uncertainty of measurement simply because you wish to make good quality measurements and to understand the results.
You may be making the measurements as part of a:
  calibration - where the uncertainty of measurement must be reported on the certificate.
  test - where the uncertainty of measurement is needed to determine a pass or fail.
  tolerance - where you need to know the uncertainty before you can decide whether the tolerance is met.
Notes: When you measure the count rate near a radioactive source, the Geiger counter always measures the total rate due to the source and the background. How do you remove the background from the measured rate to get the rate due to the source only, and how does this affect your uncertainty? This question is particularly important if you suspect that you are near a weak source of radioactivity, a source which increases the radioactivity at your location only slightly above background. How do you decide whether the source is actually present?
Example: A Geiger counter is placed near a suspected source of radioactivity and it records 58 counts in 30 sec. The source is removed and the background count is found to be 48 counts in 30 sec. Can we be sure that the source is truly radioactive?

## Slide 9
Spread…standard deviation (σ)
The standard deviation is a measure of the dispersion of randomly occurring events around a mean.
The usual way to quantify spread is standard deviation. The standard deviation of a set of numbers tells us about how different the individual readings typically are from the average of the set.
As a ‘rule of thumb’, roughly two thirds of all readings will fall between plus and minus (±) one standard deviation of the average. Roughly 95% of all readings will fall within two standard deviations. This ‘rule’ applies widely although it is by no means universal.

## Slide 10
Calculating an estimated standard deviation (σ)
Consider the set of readings 1,2,3,4,5,6,7 and the average is 4.
Next, find the difference between each reading and the average and it comes -3,-2.-1,0,1,2,3
Next square each of these, i.e. 6, 4, 1, 0, 1, 4, 6.
Next, find the total and divide by n-1, (in this case n is 7, so n-1 is 6),
i.e. 6+4+1+0+1+4+6 = 3.66
6
The estimated standard deviation (σ) is found by taking the square root of the total, σ=√3.66 = 1.91.

## Slide 11
Sources of Errors and Uncertainty
GENERAL SOURCES OF ERROR
Assuming the counting system is calibrated correctly, there are five general sources of error associated with counting a sample:
Self-absorption
Backscatter
Resolving time
Geometry
Random disintegration of radioactive atoms (statistical variations)

## Slide 12
Types of Errors
Systematic Errors
  uncertainties in the bias of the data, such as an unknown constant offset, instrument mis-calibration
  implies that all measurements are shifted the same (but unknown) amount from the truth.
  measurements with a low level of systematic error, or bias, have a high accuracy.
Random Errors
  arise from inherent instrument limitation (e.g. electronic noise) and/or the inherent nature of the phenomena (e.g. biological variability, counting statistics)
  each measurement fluctuates independently of previous measurements, i.e. no constant offset or bias – measurements with a low level of random error have a high precision.

## Slide 13
Probability Distribution Function
These functions, characterize random errors with a distribution Statistical Models:
Binomial Distribution Random independent processes with two possible outcomes.
Poisson Distribution Simplification of binomial distribution with certain constraints.
Gaussian or Normal Distribution Further simplification if average number of successes is large (e.g. >20).

## Slide 14
Simple Examples
The following are measurements of counts per minute from a 22Na source. What is the count rate and its uncertainty?
  2204, 2210, 2222, 2105, 2301
  x = 2208.4 hence σ= √2208.4/5 = 21
  Therefore count Rate = (2208 ± 21)counts/min

## Slide 15
The Effect of Background (1)
!
When the measurement result includes a contribution from the radiation background, the background value must be subtracted from the measured value.
!
Results of repeated background measurements are also random and follow a normal distribution
Notes: The normal distribution of the background measurement results is also an assumption.

## Slide 16
The Effect of Background (2)
!
The difference between sample and background values also follows a normal distribution with true mean equal to the difference of true sample and background means, but with variability greater than the variability of sample or background counts alone
!
If the counting times of sample and background are equal, then the best estimate of the standard deviation of the difference is given by the square root of the sum of the two measurements, except when they consist of only a few counts
Notes: N is the sample count, B is the background count, assuming that the sample and background measurement times are equal.

## Slide 17
The Effect of Background (3)
Background count (600 s) = 380 counts
Sample count (600 s) = 650 counts
Net counts = 650 – 380 = 270 counts
  Standard deviation of the difference, σ = √(650 + 380) = 32.1 counts
  95 % conf. level (1.96 σ ) = 62.9 counts
  The net count may be expressed as 270 ± 63, at 95 % confidence level
  To get the count rate, divide by 600 s: 0.45 ± 0.11 cps

## Slide 18
LC, LD and MDA (1)
LC = critical level
LD = limit of detection
MDA = minimum detectable concentration
The critical level (LC) is the level, in counts, at which there is a predetermined statistical probability of incorrectly identifying a measurement system background value as “greater than background”
Any response above the LC is considered to be greater than background

## Slide 19
LC, LD and MDA (2)
The limit of detection (LD) is an a priori estimate of the detection capability of a measurement system, and is also reported in units of counts.
The minimum detectable concentration (MDA) is the LD in counts multiplied by an appropriate conversion factor to give units consistent with a required measurement, such as Bq/cm2 or Bq/m3.
MDA is the lowest activity value that can be achieved when a sample is measured with a detection system.

## Slide 20
LC, LD and MDA (3)
Gamma-ray spectrometry: MDA depends on the background (statistical nature), counting time, detector and sample properties, measurement geometry, nuclear 2 and sample properties, measurement geometry, nuclear decay data of the considered radionuclide.
The MDA depends on the
  Background of counting system
  Counting efficiency of a counting system
  Counting time

## Slide 21
LC, LD and MDA (4)
The critical level and the detection limit can be calculated by using the following formulae:
where
LC = critical level (counts)
LD = detection limit (counts)
k = corresponds to (1-α) and (1-β) probability levels of the standardized normal distribution (assuming α and β are equal)
B = number of background counts that are expected to occur while performing an actual measurement.
Notes: Based on CURRIE’s 1968 paper “Limits for qualitative detection and quantitative determination, Application to radiochemistry” - one of the most influential and cited papers in the area of radioactivity
measurements.
Times for counting BG and sample + BG are equal.

## Slide 22
LC, LD and MDA (4)
If values of 0.05 for both α and β are considered acceptable, then k = 1.645 (from look-up tables) and the equations can be written as:
Where C is the conversion factor

## Slide 23
MDA for Air Sampling Systems*
Rs = sample count rate in s-1
Ts = sample counting time in s
Tb = background counting time in s
E = fractional filter efficiency in (%)
F = air flow rate through sampler in m3 s-1
K = counting efficiency in s-1 Bq-1
*NUREG 1400
Notes: The above formula is a practical example, there are many different formulas, e.g. for gamma sepctrometry, etc.

## Slide 24
Conclusions
Statistics is a part of measuring
Counting values have uncertainties
  Detection level and MDA are important parameters helps in reporting the results confidently.