## Slide 1
LESSON 2:
BASIC CONCEPTS OF WORKPLACE MONITORING – PART 2

## Slide 2
Lesson 2: Basic Concepts - II
Recording the Monitoring Results
Check List to Pickup Proper Monitoring Equipment
The Measurement Process
Quality Assurance
Factors which will influence results
General Considerations for all Measurements
Survey instruments
  Selection
  Calibration and Testing

## Slide 3
CHOICE OF INSTRUMENTS

## Slide 4
Selection of Instruments
The qualified expert.
The person who is responsible for maintenance, repair and testing of the equipment.
An administrator, to ensure that an appropriate contract is put in place with the supplier.
A representative from the monitoring staff, with a wide practical experience of radiation protection measurements.
There are 4 specialists who should be involved in
instrument selection:

## Slide 5
Choice of Instruments
Factors which affect the choice of radiation monitoring instruments for any particular application include:
  The types of radiation to be measured.
  Dose, dose equivalent rate or contamination measurements.
  Expected energy range of radiations present.
  Best available detection efficiency for radionuclide(s) of concern, based on fingerprint
  Required indication (units).
  Sensitivity and range of measurements required.
1

## Slide 6
Choice of Instruments
  Whether an instrument can perform more than one function or if more than one instrument is required
  Indication goes high or off scale if over ranged.
  Possible interfering radiations present.
  The potential for the presence of very high dose rates or contamination levels.
  The speed with which an instrument responds.
  Logarithmic/linear analogue scales or digital displays and ease of use.
  Illuminated display and/or audible output.
2

## Slide 7
Choice of Instruments
  Influence of ambient temperatures, humidity, radio-frequency radiation, magnetic fields, etc.
  Safety in explosive/flammable locations.
  Ease of decontamination
  Battery availability and life expectancy
  Size, weight and portability
  Ruggedness, reliability and serviceability
  Human factors, handle design, switches, etc.
  Training requirements for users
3

## Slide 8
Choice of Instruments
Access requirements, and ease of installation or maintenance for installed equipment.
Installed or hand held equipment.
Ability to network and remote readout.
Availability of spare parts.
  Initial and ongoing maintenance cost.
Access to a country supplier representative.
Good technical support services.
Warranties and delivery times.
4

## Slide 9
TYPE TEST

## Slide 10
Type Test
A workplace monitoring instrument is required to be type tested to demonstrate its suitability.
Type tests are intended to determine the characteristics of a particular type or model of an instrument under a series of irradiation and storage conditions.
It is generally carried out by manufacturer of instruments.

## Slide 11
Type Test
Involves extensive testing to determine the radiological performance (e.g. response, linearity, energy dependence, angular dependence etc.) and the environmental, electrical and mechanical performance.
The knowledge of the instrument characteristics is the basis for the choice of a suitable instrument to be used in a specific radiation condition.

## Slide 12
CALIBRATION AND TESTING

## Slide 13
Calibration
Calibration is a process of comparing an instrument indication with the true value of the quantity of interest.
Workplace monitoring instruments used for external exposure should be calibrated in terms of the operational quantities H*(10) and H'(0.07,Ω).
Sources and dose rates employed should be traceable to a National Standards Laboratory.
For contamination monitors, a calibrated response should be provided to the user to convert the reading to surface activity.

## Slide 14
Calibration
The quantities and units used in calibration should, wherever possible, conform to the current recommendations of the ICRU. IAEA Safety Report Series No. 16 gives clear guidance on the calibration procedures.
Calibration confirms the instrument is operating as expected by the types test and should be conducted prior to use, at a frequency typically required by regulation, e.g. annually, and after repair.
A sticker should be attached to the instrument to indicate its calibration validity and a calibration certificate should be retained

## Slide 15
IEC Standards for Workplace
Monitoring Instruments
IEC 60532:2010 – applicable to dose rate meters
IEC 60846-1:2009 – specify the design requirement and performance characteristics of dose rate meters
IEC 60846-2:2007 – applies to portable radiation monitors
IEC 61005:2003 - Specifies requirements for the performance characteristics of neutron ambient dose equivalent (rate) meters.

## Slide 16
Testing
The function test verifies that the instrument is still performing as expected
  Test is conducted at regular, short intervals by the user, for example before each use
    The frequency should be defined
  Includes a set of simple on plant tests which looks for major problems
    Includes physical inspection of the instrument, zero, battery, background and source checking
Instruments which fail testing should be identified and removed from service

## Slide 17
GENERAL CONSIDERATIONS FOR ALL MEASUREMENTS

## Slide 18
General considerations for all measurements
Instrument should be appropriate for the type, energy and intensity of radiation and for the quantity of interest.
Check there is a valid calibration and test certificate.
Perform the appropriate function checks on required ranges using appropriate sources
Instrument should be appropriate for the type, energy and intensity of radiation and for the quantity of interest.

## Slide 19
General considerations for all measurements
  Orient the instrument correctly
  Choose correct range – usually low
    When entering known high dose rates, operator doses may be reduced by selecting a dose rate range that will reduce the overall time to make the measurement
  Be aware of time constants.
  Allow correct time for instrument to respond do dose rate before moving to next measurement

## Slide 20
General considerations for all measurements
Protect or allow user modification.
If an audible output is provided, use it to give an instant indication, particularly of hotspots and narrow beams.
Consider whether alarms could be useful. If so, set to the appropriate levels.

## Slide 21
General considerations for all measurements
Believe the instrument reading. Do not assume it is an instrument problem.
If the indication is not what is expected, consider leaving the area. A lower indication than expected or a zero may be caused by instrument failure.
If the indication is significantly higher than expected, leave the area. This applies particularly if an instrument goes into overload.

## Slide 22
General considerations for all measurements
For dose rate measurements, go no closer than 3 detector dimensions to the source (due to inverse square law requirement).
Be aware of narrow beams from shielded sources.
Always seek to minimise personal dose/contamination.
  A detector will average through its sensitive volume and the indication from a large area detector may underestimate the dose rate in the beam.

## Slide 23
General considerations for all measurements
Consider whether the source may be pulsed.
  Pulse counting equipment, such as a Geiger-Mueller (GM) based unit, generally under-reads when the count rate from the detector exceeds 30% of the pulse repetition frequency.
Be aware of any environmental aspects - rapid changes of temperature, high humidity, high electromagnetic and electrostatic fields.
Consider dynamic ranges.
If not using instrument for a long time, remove batteries.

## Slide 24
General considerations for all measurements
Be aware of any personal dangers in the area to be monitored.
  Be sure that the monitoring equipment is suitable when these are taken into account.
  For example, telescopic or long reach instruments can be dangerous if used in areas where there are exposed electrical conductors.
  For example, telescopic or long reach instruments can be dangerous if used in areas where there are exposed electrical conductors.
  For example, telescopic or long reach instruments can be dangerous if used in areas where there are exposed electrical conductors.

## Slide 25
General considerations for all measurements
Be aware of any potential problems with the area to be monitored which could cause damage to the equipment such as the presence of spikes which could penetrate detector windows, or contamination of contamination monitors.
Make sure you have a means to record results with you.
  Do not record zeros or background, record the reading.
  Record any relevant conditions (e.g. power level, dusty conditions) where this might affect results.

## Slide 26
General considerations for all measurements
For installed equipment, the appropriate location will have to be chosen with care to reflect occupancy and the sources of dose rate or activity. Cabling and services will have to be provided.
For portable equipment, monitoring points may well have been identified. Be sure you know where they all are.

## Slide 27
Factors which will influence the final results
Instrument reading
Background correction
Fingerprint and its application
Calibration factor, either obtained directly during the process of calibration or derived from that data combined with other information on nuclide, energy, etc.
2
1
3
4

## Slide 28
Factors which will influence the final results
Source efficiency - rate of the emission of particles from the surface being monitored to the rate of generation of that particular particle (including geometry and self absorption)
For wipe counting, the removal factor and the area wiped.
Conversion to the quantity of interest, if necessary.
Other influence quantities (temperature, pressure).
6
5
7
8

## Slide 29
RECORDING THE MONITORING RESULTS

## Slide 30
Recording the monitoring results
The records should contain following details:
Where?
    The location of the monitoring points, described clearly and unambiguously.
When?
    When the monitoring took place.

## Slide 31
The records should contain following details:
Who?
    Who performed the measurements.
With What?
    Which type of monitoring instrument was used, including serial number, the date of calibration of the equipment and confirmation that the equipment was in an adequate condition for use.
Recording the monitoring results

## Slide 32
The records should contain following details:
How?
    How that piece of equipment was employed; Any work procedure followed;
How much?
    The actual measured value of the dose rate or contamination level, including the measurement quantity (do not record “background” or “zero”).
Recording the monitoring results

## Slide 33
CHECKLIST TO CHOOSE A PROPER MONITORING INSTRUMENT

## Slide 34
Checklist to choose a proper monitoring instrument
Does it have a valid test certificate
Does it look to be in working order
Is there, for example, obvious
damage to cable or connectors
Is the battery good enough to last
out the work period?

## Slide 35
Checklist to choose a proper monitoring instrument
Is the background count rate believable
Does the count rate change when the cable
between probe and rate meter is flexed
Use the appropriate check source for functional test
What to do if the Instrument appears faulty
A user should know whom to approach to rectify the
instrument, where it should be kept and how it should be labeled
!
!

## Slide 36
Checklist to choose a proper monitoring instrument
Where
Where each monitoring point or area is and how to identify it
Where
The frequency at which the monitoring should be carried out. This could be task based or time based.

## Slide 37
THE MEASUREMENT PROCESS

## Slide 38
The measurement process
This should cover:
The reference direction
How to check battery and select range
Use of any audio output (clicker or alarm)
For contamination monitoring, the correct distance between probe and surface and how to maintain it safely

## Slide 39
The measurement process
How to take the measurement, considering the time constant of the instrument and how long to average the indication for
  This should also cover the interpretation of logarithmic scales.
  How to use collection media and preserve it.
Where and how to record the data.
How to respond to any unusual values.

## Slide 40
The measurement process
How to identify environmental conditions which might cause instrument problems ?
  High electromagnetic, electrostatic and magnetic fields can cause high count rates or the instrument may fail.
  Very bright lights can cause spurious count rates.
  Cold instruments into warm, damp environments can cause condensation, giving spurious high dose rates with ionisation chamber instruments.

## Slide 41
The measurement process
Maintenance by the User
  This could include zero setting, battery changing, cleaning and checking.
Modifications
  Do not perform unauthorized modifications on instruments.
  If modifications are required, talk to the manufacturer.
  Re-calibrate any instrument following modification.

## Slide 42
The measurement process
Precautions
  Do not swing probes by the cable.
  Avoid getting instruments damp.
  Take care of instruments.
  Do not inhibit or adjust alarms unless authorized.
Transport to avoid damage.
Avoid damaging or contaminating.
1
2
3
4

## Slide 43
QUALITY ASSURANCE

## Slide 44
Quality Assurance
Quality assurance, QA, is a process designed to give confidence in the quality of measurements and to promote feedback leading to an increase in the quality in the widest sense.
The nature and extent of the QA programme should be consistent with the number of workers potentially exposed and the level of that exposure.
All persons involved in the monitoring programme are responsible, to some degree, for implementing quality control (QC) procedures.
Overall responsibility should be given to an identified individual.
Management should motivate staff to detect, report and correct problems.
Record and control software versions.