## Slide 1
LESSON 8:
GAMMA AND ALPHA SPECTROMETRY FOR WORKPLACE MONITORING

## Slide 2
Content
Objectives
Need for Spectrometry method of analysis
Introduction to alpha and gamma spectrometry
Alpha Spectrometry system
Gamma spectrometry system
Theory
  Detectors and their applications in workplace monitoring
Theory
  Detectors and their applications to workplace monitoring

## Slide 3
Need for Spectrometry Method
  Qualitative and quantitative determination of isotopes present in the workplace is sometimes necessary to decide on suitable protective measures.
  Gross alpha and gamma activity does not provide energy information about isotopes.
Radionuclide distribution can be identified.
Natural activity may be eliminated by means of energy discrimination
NORM and man made isotopes may be distinguished through pulse height discrimination

## Slide 4
Examples of Application of
Spectrometry
Samples from workplace containing mixture of radionuclides can be analysed using spectrometry.(e.g. filter papers and air/gas samples).
Assessment of contamination from surface spectrometric measurement.

## Slide 5
Introduction to Gamma and Alpha Spectrometry
Radionuclides emit alpha and gamma radiation with characteristic energy.
Energy selective detectors (e.g. NaI(Tl), HPGe, PIPS,SiO2) yield electronic pulses with amplitudes proportional to the emitted energy.
Measurement system (spectrometer) arranges the pulses by amplitude and counts them.
Plot of count rate vs energy is called as spectrum.
Spectrum evaluation gives information on the radionuclide via energy and activity via counts.

## Slide 6
Introduction to Gamma and Alpha Spectrometry
Characteristics of alpha and gamma spectra:
Alpha and gamma spectra can provide fingerprints of radionuclides.
Depending on the resolution of detector the isotopes with closed peak energies can be resolved.
Area under the peak is proportional to activity of the given sample.

## Slide 7
GAMMA SPECTROMETRY

## Slide 8
Alpha Spectrometry-Theory
Alpha energy is converted to electronic signal (pulse) via semiconductor detector.
  Passivated Implanted Planar Silicon (PIPS) Detector
Intensity of pulse height produced is proportional to the alpha energy.
  Mono-energetic alpha particles
  Energy Calibration vs. Pulse Height
Number of pulses are proportional to the
amount of isotope present.
  Efficiency Calibration

## Slide 9
Alpha Spectrometry
Alpha spectroscopy is a widely used technique for the identification and quantification of alpha-emitting radionuclides.
  Naturally occurring alpha emitters
  Transuranic elements, special nuclear materials
It is characterized by high efficiency, low background and low detection limits.
It can be applied for the assay of a variety of samples.

## Slide 10
Typical Alpha Spectrometry
Block Diagram
Alpha spectrometer with Multi Channel Analyser

## Slide 11
Passivated Implanted Planar Silicon (PIPS) detector
active area,
resolution,
depletion depth.
Detectors are optimized for Alpha Spectroscopy applications which require high resolution, high sensitivity and low background.
There are three main parameters that define a silicon surface-barrier detector:

## Slide 12
Passivated Implanted Planar Silicon (PIPS) detector
  Operating Principle
Alpha particles are stopped in the depletion region, forming electron-hole pairs.
The number of electron-hole pairs ultimately formed is thus directly proportional to the energy of the particle.
The electric field in the depletion region sweeps the electrons to one terminal and the holes to the other.
The resultant charge pulse is integrated in a charge sensitive preamplifier to produce a voltage pulse.

## Slide 13
Depletion Layer in Surface Barrier
Courtesy: Canberra

## Slide 14
Typical Alpha Spectrum of Natural Uranium
Alpha spectrum recorded using surface barrier detectors

## Slide 15
Laboratory Alpha Spectrometry System
Most widely used are silicon semiconductor detectors (e.g. PIPS, SSB)
Sample holder in spectrometer

## Slide 16
Application of Alpha Spectroscopy in CAM
CAM for alpha detection
Typical PIPS detector
Courtesy:Canberra
Notes: Typical comercially available continuous air monitor for alpha and beta emitters.
Here the detector is a planar silicon semicondutor device.

## Slide 17
Example: Alpha Spectrum of Air Filter
Alpha spectrum of a filter paper sample
Notes: An example alpha spectrum from the air monitor. The spectrum allows the alpha particles emitted by the naturally occuring decay products of radon-222 to be distinguished from, for example, Pu-239 which would appear in the spectrum in the region of interest marked between 3 MeV and 5.7 MeV.
For example,
Po-218 , alpha energy = 6.11 MeV
Bi-214, alpha energy = 3.27 MeV
Po-214, alpha energy = 7.83 MeV

## Slide 18
Collection Media
The collection media should be carefully chosen.
Glass fiber filters being used in alpha spectrometers should have good “front surface” collection characteristics that will prevent burial of collected particles in the filter bed.
Glass fibre filter papers

## Slide 19
Portable continuous Air Monitors
Portable planar silicon detectors for alpha spectrometry

## Slide 20
GAMMA SPECTROMETRY

## Slide 21
Gamma Ray Spectrometers- Theory
Gamma spectrometry is based on three fundamental interactions of radiation with matter.
  Photo electric effect
  Compton scattering
  Pair production
The probability of gamma ray interaction is depends on atomic number, density of materials and energy of gamma rays.

## Slide 22
Gamma Ray Spectrometers- Theory
Conversion of ionizing radiation into visible light is called Scintillation.
Depending on the application many types of detectors are employed.
For high efficient detection of gamma one has to rely on NaI(Tl)
For high spectroscopic resolution, HPGe is preferred

## Slide 23
Block Diagram of Gamma Ray Spectrometer
Block diagram of a basic gamma spectrometry system
Notes: Analogue to digital converter converts the pulse amplitudes to digital format
Multi-channel Analyzer (MCA) has a set of „channels”, i.e. storage places for pulses with amplitudes of subsequent ranges

## Slide 24
Gamma Spectrometers Set Up
The spectrometer shall be first adjusted to your measurement task:
Set up the energy – channel number function by reference sources with known energies ( energy calibration).
Set up the activity – counts function by reference sources with known energy and activity ( efficiency calibration).
Reference sources (IEC recommended) should have the same geometry as your sample.
Notes: For semiconductor detectors (e.g. HPGe, SiO2), the channel scale can easily be converted to an energy scale, as the response is linear.
Geometry means the general characteristics of the sample, not only the source – detector distance! (E.g. size, shape, material, mass.)

## Slide 25
Types of Gamma Ray Spectrometers
NaI(Tl)
High efficiency but poor resolution ( 7-8 %@662 keV)
HPGe
High resolution (0.2 -1.2 %@662 keV)
CdZnTe
high efficient and very compact
Operable at room temperature
Need liquid nitrogen cooling or electrical cooling.
Moderate energy resolution (2 keV @ 662 keV)
LaBr3
Ce Low resolution (2.8-4 %@662 keV) but high efficiency
Detectors are chosen depending on the application

## Slide 26
The spectrum evaluation consists of:
Identification of peaks
Determination of peak energies
Identification of radionuclides
Calculation of peak total areas, subtraction of background
Determination of activities and uncertainties
Use of a spectrometry software
Allows the identification and quantification of gamma emitting isotopes in a variety of samples
Gamma Ray Spectrometry as Analytical Tool

## Slide 27
Components of Gamma Spectrum
Theoretical spectrum of gamma ray

## Slide 28
Semiconductor Detectors (1)
Notes: In semiconductor detectors, an electric field is applied to the detector volume. An electron in the semiconductor is fixed in its valence band in the crystal until a gamma ray interaction provides the electron enough energy to move to the conduction band. Electrons in the conduction band can respond to the electric field in the detector, and therefore move to the positive contact that is creating the electrical field. The gap created by the moving electron is called a "hole," and is filled by an adjacent electron. This shuffling of holes effectively moves a positive charge to the negative contact. The arrival of the electron at the positive contact and the hole at the negative contact produces the electrical signal that is sent to the preamplifier, the MCA (multi channel analyzer), and on through the system for analysis. The movement of electrons and holes in a solid-state detector is very similar to the movement of ions within the sensitive volume of gas-filled detectors such as ionization chambers.
Common semiconductor-based detectors include germanium (e.g. HPGe), cadmium telluride, and cadmium zinc telluride.
Cryogenic(< -150 °C) temperatures are vital to the operation of germanium detectors.

## Slide 29
Semiconductor Detectors (2)
HPGe detector types and efficiency curves
Courtesy: Canberra
Notes: Older type of Ge detectors (GeLi) needed continuous cooling, the recently used hyper-pure Ge detectors (HPGe) shall be cooled only when the high voltage is on the detector.
HPGe detectors differ not only in size and geometry (e.g. planar or coaxial), but also in the arrangement of p-n regions

## Slide 30
Spectrometry Process

## Slide 31
Energy Resolution
Notes: NaI(Tl) means that the crystal of NaI is doped with Tl.
The explanation to the better energy resolution in a semiconductor detector is that the energy needed to create a ion-hole pair in the semiconductor is much less than the energy to produce a light photon of light in the scintillation detector. The statistical fluctuation and hence the width of the full energy peak will be much less.
This spectrum also shows that to identify and quantifty a sample containing many gamma emitting radionuclides, a Ge detector is necessary.
In compensation, the NaI detection efficiency is higher as the NaI crystal can be considerably larger than that of the Ge crystal.

## Slide 32
APPLICATION OF GAMMA RAY
SPECTROMETRY

## Slide 33
Analysis of lab samples using HPGe spectrometry systems
HPGe system for radionuclide analysis
Notes: Examples of commercial laboratory gamma spectrometry systems.
Note the shielding and the container filled with liquid N2.

## Slide 34
Application of portable HPGe in contamination monitoring
An example of one portable commercial system with an electric cooling system.
Used mainly for qualitative analysis, i.e. identification of contamination and of radionuclides
Used to check the surface contamination in workplace.
Notes: An example of one portable commercial system with an electric cooling system.

## Slide 35
Application in identification unknown isotopes inside the container
Identification of unknown radioactive materials using HPGe

## Slide 36
Different configurations of portable gamma spectrometric systems
Used mainly for qualitative analysis, i.e. identification of contamination and of radionuclides.
Suitable for a few radionuclides present with distinct energies, due to the poor energy resolution.
Measures surface contamination in workplace.
Hand held portable spectrometer with NaI(Tl)
CdZnTe probes
CdZnTe probes
Notes: A typical commercially available portable NaI gamma spectrometry system.

## Slide 37
Comparison of different systems for in-situ spectrometric monitoring
HPGe
CZT 500
CZT 60
Gamma spectrum recorded using different configurations of spectrometer

## Slide 38
Gamma Spectrometry for Air Monitoring
Typical gamma spectrum of filter cartridge
Cartridge samples
Notes: Filters with a mixture of radionuclides of known activities traceable to national standards are commercially available.
These filters are used to set the energy and efficiency functions of the gamma spectrometry systems for air filter measurements.
In the above Ge spectrum, Ba-133 is used as a substitute („mock”) for I-131, as it has a much longer half-life (10.7 years), while half-life for I-131 is
only 8.02 days.