## Slide 1
LESSON 7:
TRITIUM AND C-14 MONITORING

## Slide 2
TRITIUM MONITORING

## Slide 3
Tritium Monitoring
Background and significance
        Techniques
      Real Time Measurements
  Estimation of low energy beta activity

## Slide 4
Background and Significance
Tritium is a low energy beta emitter with Emax of 18 keV and needs special techniques to measure
Tritium significantly produced in PHWR reactors.
  Nearly 2/3 of internal dose is due to tritium intakes
Fusion reactors use tritium as fuel
To a lesser extent:
  H-3 is a byproduct of nuclear ternary fission (0.01%) in light water
  H-3 found due to neutron capture by Li6 and B10 in control rods
H-3 can be found as tritiated water (HTO), Gaseous tritium (HT) and organically bound tritium (OBT)
Notes: Note that this lesson does not cover organically bound tritium – this should be collected as per particulate sampling and analyzed by liquid scintillation couning

## Slide 5
Tritium Dosimetry and Radiotoxicity
Exposure to tritium oxide (HTO) is by far the most important type of tritium exposure and HTO enters the body by inhalation or skin absorption.
The total effective dose from airborne HTO exposure is about 10,000 times more than the total effective dose from an equal exposure to HT gas.
Tritium is classified as low toxicity according to IAEA classification scheme (Group 4).
Tritium should be in higher concentration to pose airborne contamination hazard.
Therefore, safe handling of tritium is recommended.

## Slide 6
Tritium Contamination
Highly contaminated metal or plastic surfaces may release some of the loosely-bound tritium immediately after exposure to the contaminating tritiated atmosphere or liquid. This is called outgassing or offgassing and can result in significant airborne concentrations.
Tritium as HT or HTO will readily adsorb onto the surface of most metals (such as stainless steel, copper, or aluminum), plastics, and rubbers.
  The tritium will remain fairly close to the surface unless the metal is heated to a high temperature.
HTO diffuses even metallic cladding resulting in contamination in unexpected places

## Slide 7
MONITORING TECHNIQUES

## Slide 8
Tritium Monitoring
In the workplace tritium can exist as a gas (HT) or as an oxide vapour in the form of HTO.
  The sampling and measurement technique depends on the chemical form of tritium.
Due to its low beta energy direct monitoring of tritium is not possible unless the air is mixed thoroughly with the detector medium.
Monitoring of tritium may be accomplished by real time measurement or by air sampling and delayed measurement of the sample in an appropriate measurement assembly.

## Slide 9
Tritium Sampling
Sampling of tritiated water vapour can be carried out by the following techniques:
Dynamic sampling
Passive Sampling
Condensation either using a cold strip or a cold finger.
Absorption in solid absorber such as silica gel, zeolite, etc.
Condensation either using a cold strip or a cold finger.
The activity of water collected from the bubblers or from the solid absorbers (through heating or condensation) is measured by liquid scintillation
1
2

## Slide 10
Bubbler for HTO sampling
Diagram Bubbler sampler operation
Notes: Simple bubbler for HTO only

## Slide 11
Delayed Measurement using Bubblers
The sampling air passes through two sets of bubblers in series filled with clean de-mineralized water and are separated by an oven.
The first bubbler traps the tritiated vapour.
The tritium gas is oxidised in the oven using a suitable catalyst to convert it into HTO.
The oven and second bubbler are not required if it is known that tritium is present only in the form of water vapour.
When sampling with bubblers during a long time, for example 1 week, it is recommended to use a bubbler equipped with a cooling device in order to avoid the evaporation of the water.
Teflon should be avoided for sampling lines

## Slide 12
HTO and HT Sampling
Airborne and gas form Tritium sampling device
Courtesy: Southern Scientific

## Slide 13
HTO and HT Sampling
Trapping yield of 99%
Reduced evaporation due to the cooled trapping system, allowing weekly collection
Easy to use, with instant opening cabinet for sample retrieval
Accurate – electronic, accredited airflow meter

## Slide 14
Dynamic Samplers
Good points:
  Rugged equipment.
  Sensitive method.
  High sampling rate.
  Sampling rate adjustable (typically from 10 up to 60 l/h).
  High collection efficiency (>95%).
  Stability of the sampling flow rate.
  Sampled volume is accurately known.
  Possible to use for a long time sampling.

## Slide 15
Dynamic Samplers
Weaknesses:
  Require external power.
  Dilution of HTO sampled.
Check for:
  Temperature of the oven.
  Change the catalyst as required by the manufacturer (generally once a year).

## Slide 16
Dynamic Samplers
20 ml vial containing tritium free water or water absorbent as the HTO trap
In some cases, solid absorbers such as zeolite or silica gel are used as a tritiated water trap.
The top of the vial is fitted with a diffusion orrifice
The steady flow into the vial is maintained by the concentration gradient produced by the reduced level in the closed vessel compared with ambient air
A desorber is required to remove the water from the solid absorber
  The desorbed water is collected in a device at low temperature.

## Slide 17
Passive Samplers
Good points:
  Small samplers.
  Require no external power
  Inexpensive to purchase, maintain and service.
  Simple to use.
  No dilution of the HTO sampled

## Slide 18
Passive Samplers
Two different samplers to know HT and HTO concentration (one with a a catalyst).
Only use for short sampling time in case of total tritium sampler (between 15 and 30 minutes).
Low sampling rate.
Evaporation of the liquid.
Change the catalyst after each sampling for total passive sampler.
Weaknesses:
Check for:

## Slide 19
REAL TIME MEASUREMENTS

## Slide 20
Real Time Measurement
For real time measurement of tritium location of the detector should be done judiciously so that area background radiation is low.
For real time measurement, the sampled air passes through a measurement cell.
Measurement cell is warmed to prevent condensation
Because of the low beta energy of tritium, the measurement cell is the detector itself, for example an ionization chamber or proportional counter or a plastic scintillator cell of the flow- through type.

## Slide 21
Real Time Measurement
A prefilter is put upstream of the measurement cell to trap particulates from the air sampled.
  The prefilter should include an iodine trap if necessary - iodine can be a precursor to noble gases - ensure does not impact the result
  The pre-filter should not trap or temporarily retain tritium or decrease the flowrate or the pressure inside the measurement cell over the limits specified by the manufacturer
Interference of noble gas should be minimised while tritium measurement.
  Compensation for noble gases contribution is difficult to achieve in ionization chamber based instruments.
  In the case of proportional counters or plastic scintillators, the contribution due to noble gas can be reduced by taking advantage of difference in pulse height produced by beta particles of noble gas and tritium.

## Slide 22
EXAMPLES OF REAL TIME MONITORS

## Slide 23
Tritium air monitoring using ionization chamber
Fixed ionisation chamber instruments are most suitable for gaseous form.
Tritium beta particle ionizes the surrounding gas.
Portable instrument could supplement the fixed ones
Chamber volumes ranges from tenth to few tens of litres depending on required sensitivity – flow rates a few lpm
Gamma and radon compensation is crucial.

## Slide 24
Portable and Installed Tritium Air Monitore
Portable tritium monitor
With dual ion chamber
Installed tritium monitor
With dual ion chamber

## Slide 25
Ionisation Chambers
Good points:
  Small detector due to the distance covered by the beta particle of tritium in air.
  Good sensitivity for tritium.
  Low maintenance cost.
Weaknesses:
  Generally non specific for tritium measurement.
  Sensitive to temperature.
  Sensitive to pressure: the response is quite proportional to the pressure.

## Slide 26
CARBON -14 MONITORING

## Slide 27
Sources of 14C in Nuclear environment
Carbon-14 and tritium are weak beta emitters and do not present an external radiation hazard.
The major 14C producing neutron activation reactions in nuclear power reactors are:
  (a) The 14N(n,p)14C reaction
(b) The 17O(n,α)14C reaction
(c) The 13C(n,γ)14C reaction
(d) The 15N(n,d)14C reaction
(e) The 16O(n,3He)14C reaction

## Slide 28
Chemical Form of Carbon -14
C-14 can be in the forms of CO2, CO or hydrocarbons.
Carbon mono-oxides and hydrocarbons can be converted to CO2, by a catalyst at high temperature.
Techniques for sampling 14CO2 in air can be either active or passive.
The real time monitors used for tritium sampling may be suitable for C-14 also.
This is required as it is easy to measure C-14 activity in CO2 form.
Notes: Less commonly, real time measurement can be conducted for C-14

## Slide 29
The carbon in CO2 form may be trapped by bubbling through a solution of sodium hydroxide or in a solid absorber like zeolite.
  The air to be sampled passes through bubblers filled with sodium hydroxide solution in the range 0.8–4M (or tubes filled with solid absorber), an oven and a second set of bubblers (or tubes filled with solid absorber).
  The first bubbler (or tube) traps carbon in CO2 form.
  The organic molecules are broken and the carbon oxidized to CO2 form in the oven using a catalyst made of Platinum/Alumina pellets.
  CO2 is trapped in the second bubbler or solid absorber connected at the outlet of the oven.
Sampling

## Slide 30
Release of C-14 for Measurement
When the CO2 is trapped in a solid absorber it has to be desorbed by heating at the end of the sampling time.
  A temperature of 400 degrees Celsius is required for zeolite
CO2 desorbed is trapped by bubbling through a solution of sodium hydroxide.
C-14 activity counted in LSS using Toluene based scintillator.

## Slide 31
Calibration and Testing
The following should be calibrated:
  The flow meter associated with the sampling assembly should be calibrated against a national standard.
  The oxidation efficiency should be determined by sampling standard tritrium gas or organic carbon. Application of low concentration hydrogen or carbon and mass spectrometry can also be used for determination of the oxidation efficiency.
  The measurement assembly with a standard source of tritiated water, C-14 or, for monitors, with a radioactive source
Functional checking should include the detection efficiency of the measurement assembly and the sampling flow rate or volume

## Slide 32
ANALYSIS METHODS

## Slide 33
Liquid Scintillation Counting
For Liquid Scintillation Counting the sample to be counted is mixed with a mixture of certain organic chemicals, in a specified proportion, called a cocktail.
Generally the cocktail consists of a primary solvent (Dioxane for aqueous samples and Toluene for organic samples), a secondary solvent (Naphthalene only for aqueous samples), a primary solute (PPO) and a secondary solute (POPOP).
A vial containing the sample and the cocktail is placed in the measurement cell of the liquid scintillation counting system.
Notes: LSC technique detects Alpha, Beta radiation, Weak gamma, X rays, Auger electrons, Gamma
Is the best technique available for low energy emitters such as Ni-63, H-3, Pu-241, Fe-55

## Slide 34
Liquid Scintillation Counting
The scintillations produced in the cocktail by beta particles from tritium are detected by suitably positioned photomultiplier tubes.
The light output resulting from the scintillations can get reduced in the cocktail due to various mechanisms, commonly known as quenching. This will result in a reduction in the efficiency of the detector.

## Slide 35
Beta Spectrum and Counting

## Slide 36
Liquid Scintillation Operations
Light flashes detected
Number determined
Intensity determined
Number of flashes per minutes is recorded as counts per minute
Converted to nuclear desintegration rate
Data analysis

## Slide 37
Liquid Scintillation Counting
Notes: LSC holds a large number of vials, Operates automatically, Energy ranges to be set on the counter/May be built in efficiency, but should verify response using sources, Blank wipe and cocktail should also be counted
  For wipe counting, the manufacturer of the liquid scintillation cocktail may recommend a material for the wipes

## Slide 38
Liquid Scintillation Counting
Interference in liquid scintillation counting;
Background (noise and PMT crosstalk)
Static electricity
Wall effect (organic solvent penetrate the wall)
Radionuclide mixture (spectral overlapping due to presence of two radionuclides in sample)
Luminescence
  Chemiluminescence
  Bioluminescence and photoluminescence.
Notes: Luminescense is the increase in count rate caused either by exposure to ultraviolet light or the presence of chemicals in the sample. Photoluminesence will decay with removal from light and time. Chemiluminesence will not decay

## Slide 39
Practical Aspects
Solvent has to be compatible with the liquid scintillation cocktail.
Spills of cocktail on outside of vial should be avoided to prevent quenching.
Process should be clean to prevent luminescence or quenching.
Make sure energy range is correct for radionuclides to be detected.