## Slide 1
LESSON 3:
X AND GAMMA RADIATION MONITORING

## Slide 2
X and Gamma Radiation Monitoring
X and Gamma Radiation Monitoring using Portable and installed instruments
  1. Monitoring Technique
  2. Instruments
  3. Calibration and Testing
  4. Practical measurements

## Slide 3
X AND GAMMA RADIATION MONITORING INSTRUMENTS

## Slide 4
X and Gamma Radiation Monitoring Instruments
Purpose:
To measure ambient dose equivalent rate, H*(10) at the points of interest.
To assess the intensity of radiation fields at workplace for regulatory compliance.
To apply the measurement values for controlling external exposures.
1
2
3
Notes: The last two purposes are not in the TECDOC specifically for external radiation, but are consistent with previous slides.

## Slide 5
1. Monitoring Technique
The monitoring process generally involves placing the reference point of a suitable monitor.
The reference point should be as close as possible to the intended measurement point.
It measures ambient dose equivalent rate, at the measurement point.
In theory, there is no need to point the instrument in a particular direction but in practice no instrument is totally isotropic in its response.
In addition, the operator will produce considerable shielding for radiation behind the instrument. Hence, it is important to identify a direction for each monitoring point.
For this reason, area gamma monitors should be located at higher elevations.
Notes: Direction is frequently intuitive, but sometimes it is not as obvious. If it is not obvious then select a direction which gives the maximum indication and note this for future monitoring.
Low energies could occur because of scatter.
  The presence of low energies can be identified by noting whether the indication of an instrument drops significantly when a 1 mm thick steel or copper sheet is placed over the window of a wide energy range instrument.

## Slide 6
1. Monitoring Technique
Is easier as sensitivity of instrument increases.
Searching:
In areas where conditions are poorly understood
Shielding weaknesses or source searches
Highest dose rates on outside of package
Helped by using audio output.
Portable instruments are used.

## Slide 7
Installed instruments continuously monitor.
1. Monitoring Technique
Alarms and remote indicators possible.
Emphasis is on selection of best position for detectors
Consider scenarios of potential dose rate
Ensure instrument is not shielded and is located to measure point of interest
Facility design basis may dictate locations

## Slide 8
INSTRUMENTS

## Slide 9
2. Instruments
The instrument required is determined mainly by the anticipated range of dose rates and whether relatively low energy radiation is present.
It is important to choose equipment which has a desirable response at the lowest dose rate of interest.
Another main point of consideration is whether the radiation field has a significant low energy component (below 60 keV) and the range of an instrument.

## Slide 10
2. Instruments
The equipment that can
be employed are:
GM counters
Ionisation chambers
Proportional counters
Plastic scintillation detectors
Semiconductor detectors

## Slide 11
Schematic diagram of gas filled detector
Sketch of functioning of a typical GM tube
Notes: Description of what is happening in the detector:
- Photon enters and generates an energetic electron in the walls or gas
- Electron loses energy by collision, generating ions and electrons
- In air filled types, electrons stick to oxygen molecules
- In others, the electrons remain free
- Ions and electrons drift in electric field
- The charge collected on collector is amplified and displayed

## Slide 12
GM Detectors in Workplace Monitoring
End Window | Here radiation enters the sensitive volume of the detector through a very thin mica window (1.5 - 3 mg/cm2). Protection of thin window is taken care by a mesh. End windows GM tubes can detect alpha, beta, and gamma.
Side Wall | This detector has a sliding sleeve. Beta particles (300keV and above) and gamma rays can be detected with the window open. By closing the window the beta contribution is removed.
Pancake | A pancake G-M is similar to the end window in that a very thin mica covering is used. Its design offers a greater detection area than the end window probe.
Notes: GM tubes are classified as gas-filled detectors. All such detectors consist of a volume of gas between two electrodes. Such devices indicate the number of interactions occurring in the detector and are thus called counters.
End window detectors can still see betas so may need to add a 3mm plastic cover to reduce this when measuring gamma
Telescopic detectors usually use steel walled GM detectors calibrated so they should be used side on to the radiation field
Pancake also sees beta and needs to have plastic added to only respond to gamma

## Slide 13
GM Counters - Merits
Higher sensitivity compared to ionisation chamber
Small size and volume is possible due to higher sensitivity.
Can be operated in ‘pulse’ mode or ‘current’ mode, depending on electronics.
Energy compensated instruments suitable for measuring ambient dose equivalent (H*10).
Relatively inexpensive.
Notes: GM counters connected to a count rate meter (scaler) are the most widely used instrument due to their low cost, simplicity, reliability, and high sensitivity.
The GM counter registers “counts” produced by the incident radiation which in turn produces an ionization in the sensitive volume.
The GM survey meter can be made into a fairly accurate indicator of dose rate over a wide range of photon energies. This is done through the use of filters around the detector.

## Slide 14
GM Counters - Limitations
Dead time is about 300 μsec
  Dead time correction may be needed.
Saturates at high dose rates.
No energy discrimination is possible.
Cannot be used to accurately measure the pulsed source of radiation from accelerators.
Shows energy dependency, hence filters used.
Slope of the plateau must be designed reasonably flat for reproducible results.
Notes: The GM can saturate at high dose rates. In this case, ionization chamber survey instruments should be used.

## Slide 15
Energy Dependency of GM Tubes
Bare detectors have an energy dependent response.
Energy compensation for X, gamma dose below 80 keV.
10 100 1000 keV
response
energy uncompensated
energy compensated
Notes: Also watch for beta radiation – must be shielded out

## Slide 16
GM Survey Meter
Polar response of a typical cylindrical GM for X-Ray and Gamma Ambient Dose Equivalent Rates
Good polar response

## Slide 17
Typical Examples GM Based Gas Detectors
Miniature GM tubes
End window GM tubes
Pancake GM tubes
Side wall GM tubes
Courtesy: Canberra

## Slide 18
Proportional Counters as Survey Meters
Dose rate meters using proportional counters are uncommon.
They are more sensitive than ionisation chambers
and suitable for measurements in low intensity
radiation fields.
Although smaller than equivalent sensitivity
ionization chambers they require highly stable power
making them not suitable for workplace monitoring.
Notes: Use: low energy X and gamma contamination measurement and beta contamination measurement
These are generally 10 cm x 10 cm or larger xenon filled proportional counters with titanium windows about 5 mg.cm-2 thick.
Good points
Can be obtained in large sizes.
Also sensitive to low energy X, gamma emitters such as 125I and 55Fe.
Can be re-windowed, although the cost is a large fraction of a new detector.
Tougher than the thin end window GM detectors although still vulnerable to sharp points such as tweezers, nails, wire and swarf.
Weaknesses
Expensive compared to GM detectors.
Requires a better stabilised and higher voltage supply than GM detectors, typically 1.6 to 2 kV.
On rare occasions can deteriorate over a period of days or weeks as a consequence of an undetectable leak.
Can give energy information
Good uniformity of response
Easy to repair
Higher cost, more complex, more prone to damage
Some types require gas refilling
Relatively long lived

## Slide 19
Typical Example of Proportional Counters
Proportional counters in various configurations.

## Slide 20
Ion Chamber - Characteristics
Detector operates in current mode with air as fill gas.
No gas amplification is required for operation
Designed to measure x-ray and gamma rays
Ideal for exposure rate measurements; can measure very high radiation levels with virtually no dead time
Flat energy response above 100 keV
By increasing the pressure of fill gas the sensitivity can be enhanced

## Slide 21
Ion Chamber - Strengths
Excellent X, gamma ambient dose equivalent and polar response
More reliable in case of high dose rate where GM tubes cannot be used due to overload effect.
Suitable for accident dose rate and therefore installed for the purpose of area monitoring.
Can be used in pulsed radiation fields
Can be used to measure accurately measure beta radiation if window slide provided.
Gold standard for exposure measurements.
1
2
3
4
5
6

## Slide 22
Ion Chamber - Weaknesses
Detector operates in current mode with air as fill gas.
No gas amplification is required for operation
Designed to measure x-ray and gamma rays
Ideal for exposure rate measurements; can measure very high radiation levels with virtually no dead time
Flat energy response above 100 keV
By increasing the pressure of fill gas the sensitivity can be enhanced
1
2
3
4
5
6

## Slide 23
Ionisation Chambers in Workplace Monitoring
Portable Ionisation Chamber Survey Meter
Measuring: 1 μSv/h to 1Sv/h
Measure volume: 350 cm3
Area monitor
Range:1 μSv/h to 1Sv/h
Detector: Ionisation chamber probe
Courtesy: Mirion
Courtesy: Canberra

## Slide 24
GAMMA AND X RAY DOSE RATE MEASUREMENT
(SOLID STATE DETECTOR)
Notes: These are not addressed in the TECDOC

## Slide 25
Solid State Detector
Scintillators
Semiconductors
Two types

## Slide 26
Scintillation-Phenomenon
Scintillation is the process through which ionizing radiation is converted into visible light photons.
  This type of luminescence called radio-luminescence.
Scintillation can provide the energy information and hence can be useful in nuclear spectroscopy.
  The size of the light pulse is determined by the amount of energy deposited in the event.
It takes place in the time scale of nano seconds.
Plastic scintillator, NaI(Tl), CsI(Tl) or LaBr:Ce3+ used.

## Slide 27
Schematic diagram of scintillator
Interaction of gamma with NaI(Tl)
Notes: The total current from the photomultiplier is a function of the energy deposited in the scintillator crystal which can be used to estimate the dose rate.
Converts light to an electric pulse.

## Slide 28
Scintillator - Characteristics
High atomic number and hence high efficiency for gamma
Sensitivity is 103 to 104 times higher than gas detectors depending on the material.
Higher stopping power enhance the probability of energy deposition.
Intensity of light is proportional to the energy.
Can be operated at room temperature.

## Slide 29
Scintillator as Survey Meters
NaI(Tl) based scintillators are used both for as area monitors and portable as survey instruments.
Sleek plastic scintillator based instruments with a wide range of 0.01μSv/h to 10 Sv/h are also available.
Palm or credit card sized radiation dosimeters based on Si diode with a wide range of dose(1 μSv to 10 Sv) and dose rate (1mSv/h to 10Sv/h) have been introduced recently.
Some configurations also provide spectral information of radionuclide present.

## Slide 30
Examples of Scintillators as Dose Rate Meters
Area Gamma Monitor:
Range: 0.1 μSv/h to 100 mSv/h
Courtesy: Canberra
Micro R survey meter:
Range: 0.01-100 mSv/hr
Courtesy: Nucleonix
Plastic scintillator: Survey meter
Range:1 nSv/h to 100 μSv/h.

## Slide 31
Semiconductor - Characteristics
Types
– Si diode
– Cadmium Zinc telluride (CdZnTe)
Dimension- small in size
High Energy resolution (HPGe) and therefore suitable for radiation spectroscopy.
HPGe is available with higher efficiency (nearly 100 %) with along with high resolution.

## Slide 32
Semiconductor - Examples
Courtesy: Fuji electric
X-Gamma Silicon Survey Meter
Range: 0.01μSv/h to 99.9 mSv/h

## Slide 33
CALIBRATION AND TESTING

## Slide 34
Calibration
Calibration is defined as the quantitative determination, under a controlled set of standard conditions, of the indication given by a radiation measuring instrument as a function of the value of the quantity the instrument is intended to measure.
Will typically include:
Response to high dose rates
Linearity
  Energy dependence
Directional dependence
Notes: HIGH DOSE RATE Expose the instrument to a dose rate no more than 50 % above the preset alarm threshold.
LOW DOSE RATE Confirm that the low level alarm may be activated

## Slide 35
Calibration
Calibration should use radiation qualities defined in (ISO 4037-3 – 662 keV of 137 Cs, 241Am-60keV).
Calibration is conducted at dose rates representing between 1/3 and 2/3 of each measurement scale.
Calibration factor between ±1.2 is an acceptable calibration factor.
Generally, the linearity should be determined using either 137Cs or 60Co gamma radiation over the range of dose rates for which the instrument will be used.

## Slide 36
Calibration
Any uncalibrated ranges should be identified on the instrument.
Calibration should be made at least every year
Calibration of installed monitors is performed by the manufacturer. After that, it is tested in-situ using a calibrated instrument or a source, including:
  Function check
  Background indication
  Alarm check (including high & low dose rate response)
  Response to high dose rates

## Slide 37
Functional Testing
For an installed instrument, will typically include:
Functional check
Background indication
check indicator lights are functioning
visual check of physical condition
  check alarm using the check function, if available.
check display operates correctly
1
2

## Slide 38
Functional Testing
For a portable instrument, will typically include:
Physical integrity of the detector, cable, probe
  Calibration validity
  Battery condition
  Background reading
Check source reading within the expected range
    Alarm check

## Slide 39
PRACTICAL DOSE RATE MONITORING

## Slide 40
Practical Dose Rate Monitoring
1.High activity sealed source; keep distance.
2.Searching for contamination
3.Gamma
radiography with a high activity sealed source
4.Installed gamma monitoring
Notes: First picture: High activity sealed source; keep distance
Second picture: searching for contamination
Third picture: gamma radiography with a high activity sealed source
Fourth picture: installed gamma monitoring

## Slide 41
Practical Dose Rate Measurement(1)
Choose the right equipment for the purpose.
Perform a functional test.
Check the calibration validity.
Set alarm signals for the task (dose rates and integrated doses).
Make a survey of the radiation field and point of interest.
1
2
3
4
5
Notes: Ensure you have chosen the right equipment for your environment and the purpose (radiation type, energy, dose rate ranges).
Perform a functional test.
Check the calibration validity on the detector.
Set alarm signals at known dose rates and integrated doses. Be careful in noisy workplaces. Headphones can be bought as an extra equipment. Think also at rotating beacon.

## Slide 42
Practical Dose Rate Measurement(2)
Direct the detector to the highest dose rate
Be aware of collimated beams
At unexpected high or low dose rate, act immediately
Use appropriate time constant for detector reading to stabilize
GM or proportional counters should not be used in pulsed beams
1
4
2
3
5
Notes: Make a survey – radiation fields present and points of interest.
Point the detector in the direction of the highest dose rate. Normally make the measurement at 1 meter above ground height.
Be aware if collimated beams are present; you may be in the beam and the detector out of the beam.
When an unusually high or low measurement result is indicated, act immediately on the information. Low measurement result may indicate a faulty instrument!
Allow an appropriate time for the equipment reading to stabilize.

## Slide 43
Practical Dose Rate Measurement(3)
Use telescopic detectors (or remote detectors for installed applications) for high dose rates source
Use audio signal and/or rotating beacon
Minimize the time spent in the radiation field
Cover the probe with sleeve to avoid contaminating the equipment
Turn the equipment off when not in use.
Don’t hold the probe by the cable.
Notes: Be aware that pulsed beams should not be measured with GM or proportional counters.
For high activity sources, use a telescopic detector, preferably not GM based.
Use the audio signal.
For measuring dose rates above 1 mSv/h, minimise the time spent in the radiation field.
Use also rotating beacon if possible
Avoid touching the detector on possibly contaminated surfaces. For photon detection the detector may be protected by thin plastic foil.
Keep the thin window away from sharp objects to avoid irreparable damage to the detector.
Turn the detector off when not in use.
If the equipment has a separate probe, do not hold the probe by the cable.

## Slide 44
SPECIALIZED EQUIPMENT -EXAMPLES
Notes: Check the physical integrity of the detector, cable, probe and the calibration validity.
Check the battery.
Check the expected background reading.
Place a low activity gamma emitting check source on the reference point of the detector (usually marked with a spot or small hole on the detector case). Check whether the reading is within the expected range.
Some equipment have built-in check sources.
The IAEA makes available on request all procedures related to workplace monitoring
Explain that in certain workplaces with lots of staff working with this kind of equipment it is valuable to document the functional tests in a simple way.

## Slide 45
Energy Compensated Pancake G-M
Courtesy: Thermo-Fisher
Features:
1.8 to 2.2 mg/cm2 Ultra Thin Mica Window
28 or 45 mm effective diameter
High alpha, beta and gamma efficiency
Cost effective with unmatched quality/cost ratio
Measurement range: less than 2 mSv/h
Photon energy range:17 keV to 1.3 MeV

## Slide 46
Gamma Camera
Gamma camera employed for detecting contamination in workplace monitoring

## Slide 47
Gamma Camera Employed at Fukushima

## Slide 48
Gamma Camera Employed at Fukushima Plant
Visualization results in the reactor building of the Fukushima Dai-ichi nuclear power plant. P1-A and P2-A are images of the same penetration holes observed from different locations with the gamma camera built with CdZnTe.
P1A
P2A
Ref: Progress in Nuclear Science and Technology, Volume 4 (2014) pp. 14-17

## Slide 49
Backpack for Searching of Sources
Alarm at +20% BG
Energy range: 50 keV to 3 MeV
BF3 neutron
detectors
Plastic scintillator
for gamma

## Slide 50
Novel Radiation Detectors
Courtesy: Thermo fisher
Key Features
• Pocket-sized gamma neutron pager
• Very high neutron and gamma sensitivity
• Immediate classification of gamma source
(NORM/non-NORM)
• Energy compensated gamma dose rate
• Dual gamma/neutron display
• No false neutron alarms for even intense
gamma sources
• Ideal for law enforcement officers
and first responders
• Gamma efficiency: 900 cps per μSv/h (Am-241)
• Neutron efficiency: 4.3 cps/20,000 n/s Cf252

## Slide 51
Miniature Radiation Detectors
G-M watch
NaI scintillator
Mobile phone with app
Miniaturisation results in novel and tiny radiation detectors
Notes: The iPhone can only be used for high dose rates.
Observe that the quality of these kind of instruments are very shaky!

## Slide 52
Summary: Photon Dose Rate Measurement
Always execute measurements with great care
Never take the performance of the instrumentation for granted
Regularly check the equipment
Ensure periodic calibration
Protect yourself from unnecessary exposures
Notes: Always execute measurements with great care and never take the performance of the instrumentation for granted
Continually check operation of the measuring equipment with test equipment or test runs and analyse the results critically for consistency and reasonableness
Always bear in mind your own personal safety