## Slide 1
TRAINING COURSE
ON
WORKPLACE MONITORING

## Slide 2
INTRODUCTION
OBJECTIVES OF THE TRAINING PROGRAMME

## Slide 3
OBJECTIVES
To strengthen the radiation protection programme in Member states by assisting in establishing workplace monitoring programme by providing insights on;
    Background and significance of monitoring.
    Techniques applied for monitoring,
    Instruments employed for monitoring,
    Demonstrating the practical use of instruments and interpretation of results,
    Calibration and Quality considerations.

## Slide 4
STRUCTURE OF THE TRAINING
The training course on workplace monitoring contains three modules:
Module 1 pertains to Radiation Dose Rates (X, gamma, beta and neutron) and Surface Contamination Monitoring.
There are 6 lessons in this module.
Module 2 pertains to airborne contamination
monitoring and noble gas, iodine, tritium monitoring and specific workplace applications.
There are 6 lessons in this module.
Module 3 pertains to practical exercises and demonstrations.
There are 5 practical lessons in this module.

## Slide 5
Module 1: Fundamentals of Radiation Monitoring
CONTENTS
Lesson 1: Fundamentals of Workplace Monitoring-I
Lesson 2: Fundamentals of Workplace Monitoring-II
Lesson 3: Measurement Uncertainties
Lesson 4: Gamma Monitoring
Lesson 5: Neutron Monitoring
Lesson 6: Surface Contamination Monitoring

## Slide 6
Module 2: Airborne Contamination and Specific Applications
CONTENTS
Lesson 7: Airborne Contamination Monitoring
Lesson 8: Tritium and C-14 monitoring
Lesson 9: Alpha and Gamma Spectrometry
Lesson 10: Iodine Monitoring
Lesson 11: Noble gas Monitoring
Lesson 12: Monitoring in Nuclear Power Plants

## Slide 7
MODULE 3: Practical Lessons
CONTENTS
Practical 1: Workplace Monitoring
  Design of a Monitoring Program
  Preparation/Review of Documentation
Practical 2: Dose Rate Monitoring
  The selection and use of radiation survey instruments for gamma and neutron measurements

## Slide 8
MODULE 3: Practical Lessons
CONTENTS
Practical 3: Surface Contamination Monitoring
  The selection and use of alpha and beta contamination monitors and counting systems
  Calculation of results from measurements
  Counting statistics and the impact on results

## Slide 9
CONTENTS
Practical 4: Workplace monitoring
  Performance of surveys using dose rate meters, surface contamination monitors for direct measurements and wipe testing, including liquid scintillation counting
  Monitoring follow up to contamination and radiological incidents
MODULE 3: Practical Lessons

## Slide 10
CONTENTS
Practical 5: Airborne Monitoring and Spectrometry
  Airborne contamination monitoring using portable air samplers and fixed air monitors.
  Alpha spectrometry discussion/demonstration
  Gamma spectrometry using NaI scintillation detectors and hyperpure germanium (HPGe) – discussion/demonstration
MODULE 3: Practical Lessons

## Slide 11
LESSON 1:
BASIC CONCEPTS OF WORKPLACE MONITORING– PART 1

## Slide 12
LESSON 1: BASIC CONCEPTS- I
Introduction
Purpose of Workplace Monitoring
Designing the Monitoring Programme
  Operational Quantities for External Radiation
  Fingerprints
  Types of Monitoring
  Frequency and Location of Monitoring

## Slide 13
INTRODUCTION

## Slide 14
INTRODUCTION
Ionising radiations have the potential to harm human body if exposed to excessive levels
Occupational exposures are associated with various activities of the nuclear fuel cycle
A framework of radiation protection is established
  To limit exposure
  To optimise protection within these limits.

## Slide 15
Reliable measurements of radiation exposures from radioactive materials are essential
  to demonstrate compliance with regulatory limits
  to allow optimisation of working conditions
  to identify any failures or breakdown in control
The general term “monitoring” describes measurements related to identification, quantification and control of exposure due to radiation and radioactive material.
  Monitoring is more comprehensive than simple measurement.
  It comprises both assessment and interpretation.
INTRODUCTION
Notes: The understanding of the fact that ‘Monitoring’ is more than just ‘Measurement’ is important. One must understand what the value received in a measuremant stands for and what to to with it. For example a measurement of airborne activity may say it is reading 100 Bq/m3 but in fact the masurement has measured radon and long lived airborne activity. The interpretation of the measurement provides the monitoring result.

## Slide 16
INTRODUCTION
Workplace monitoring (WPM)
  comprises measurements made in the working environment
Individual monitoring (IM)
  monitoring by equipment worn by individual workers or measurement of quantities of radioactive material invivo or invitro and the interpretation of such measurements.
Notes: Equipment includes dosimeters

## Slide 17
INTRODUCTION
From experience it has been seen that:
  Sometimes only WPM is necessary and/or possible
  Sometimes only individual monitoring is necessary
  Normally both are necessary and possible
As an Agency guideline, if potential for committed effective dose could exceed 1 mSv/year then IM should be considered.
WPM could be made below 1 mSv/year for confirmatory reasons in supervised area.
Notes: The effective dose constraints/guidance values vary nationally and according to company guidelines.

## Slide 18
PURPOSE OF WORKPLACE MONITORING

## Slide 19
PURPOSE OF WORKPLACE MONITORING
Monitoring may be part of the prior radiological assessment
Determination of whether radiological exposures in the workplace are under adequate control or whether operational changes have improved or worsened the situation.
  These measurements facilitate the control of exposure of workers by early detection of deteriorating or abnormal conditions, thereby allowing appropriate remedial actions.

## Slide 20
PURPOSE OF WORKPLACE MONITORING
Selection of appropriate personal protective equipment
Assistance in preventing the spread of contamination
Provision of information that can be used to allow workers to understand how, when and where they are exposed and to motivate them to reduce their exposure.

## Slide 21
PURPOSE OF WORKPLACE MONITORING
Demonstration of good working practices and engineering controls (e.g. the adequacy of supervision, training, etc.)
Detection of failures of containment or departures from good operating procedures
Classification of controlled and supervised areas.
Provision of a prompt and reliable alert to allow workers to evacuate to a place of safety

## Slide 22
PURPOSE OF WORKPLACE MONITORING
Contribution to the dose assessment or estimation for individual or groups of workers and demonstration of compliance with dose constraints
Provision of information for
  The planning of programmes of individual monitoring
  Defining operational procedures and facilities
  Evaluation of doses in the event of accidental exposures

## Slide 23
DESIGNING THE MONITORING WORKPLACE

## Slide 24
DESIGNING THE MONITORING PROGRAM
The Basic Safety Standard states that the type and frequency of workplace monitoring shall be based on:
  Expected fluctuations of:
    Dose rate, Surface contamination,
    Activity concentrations in air
  The likelihood and magnitude of exposures in anticipated occupational occurrences and accident conditions

## Slide 25
REQUIREMENTS OF MONITORING PROGRAM
According to BSS(2014), the monitoring programme should;
  Specify the quantities to be measured.
  Identify the most appropriate measurement methods.
  Outline the procedures to be adopted.
  Recommend the frequency of monitoring and monitoring points.
  Stipulate the reference levels and actions to be taken if these levels are exceeded.

## Slide 26
OPERATIONAL QUANTITIES

## Slide 27
NEED FOR OPERATIONAL QUANTITIES
Operational quantities are needed for monitoring external exposures because:
The protection quantities generally are not measurable,
exposure limits are given in terms of protection quantities.
for area monitoring a point quantity is needed
Quantities are defined by ICRU-86,88,92

## Slide 28
EXTERNAL RADIATION
For the purpose of area monitoring ambient dose equivalent H*(d) for gamma radiation and the directional dose equivalent H’(d,Ω) for beta gamma radiation are defined. H’(3,Ω) for dose equivalent to the eye.
The units of effective and equivalent dose are based on these measurement values
It is good practice to have all monitors at the workplace using the same quantities

## Slide 29
EXTERNAL RADIATION
The ambient dose equivalent at a point in a radiation field is the dose equivalent that would be produced by the corresponding expanded and aligned field in the ICRU sphere (30 cm diameter tissue) at a depth ‘d’, on the radius opposing the direction of the aligned field. The recommended value of ‘d’ for penetrating radiation is 10mm.
Notes: It is impractical to align the radiation field in the sense of moving all the sources present so that they lie along the same line. The equivalent effect is produced by designing instruments so that their response depends as little as possible on the direction of irradiation. This is generally termed an isotropic response. This is important because we do not always know the direction of the radiation that workers are being exposed to.

## Slide 30
EXTERNAL RADIATION
Primary standards for ambient and directional dose equivalent, H*(d) and H´(d), do not exist.
Reference fields for calibration of instruments are usually realised in terms of
  radiation fluence rate, Ǿ, (for neutrons)
  air kerma rate, Ka , (for photons) ,
  absorbed dose to tissue (electrons),
  and the application of fluence-(or air kerma or tissue absorbed dose) to-dose equivalent conversion coefficients.

## Slide 31
EXTERNAL RADIATION – H’ (0.07)
Directional Dose Equivalent, H(0.07), expressed in Sv, is the dose equivalent at a depth of 0.07 mm in a similar sphere, and is the quantity used to control dose equivalent to skin.
H(0.07) is most important for weakly penetrating radiation, generally  radiation and low energy x and  radiation, although penetrating radiations also contribute if present
The unit of H’(0.07) is the Sievert or sieverts/hour.
Notes: The fields are not aligned in this case, therefore direction is important – you need the find the radiation and point the meter towards it
The operational units are directly measured – Seiverts/hour – and this can be directly related back to the dose limit

## Slide 32
SURFACE CONTAMINATION
Surface contamination is the ratio between the activity of the radionuclides present on a surface to the area of the surface.
Surface contamination is given in terms of “activity per unit area” (expressed in Bq/cm2)
For contamination monitoring, the instrument typically reads in counts per minute and has to be converted to the operational unit based on the efficiency of the instrument, either the area of the instrument (for direct monitoring) or the area wiped (for loose contamination monitoring), the kind of material/surface and contamination.

## Slide 33
AIRBORNE CONTAMINATION
Airborne activity is the ratio between the activity of the radionuclides in the air and the volume of the air
Airborne activity is given in terms of “activity per unit volume” (expressed in Bq/m3)
Airborne activity is typically measured in units of activity and then has to be converted to the operational unit using the efficiency of the detector and the volume of air sampled
Notes: Not only is the operational unit not directly measured, the significance of the operational unit also has to be compared with a derived measurement to relate it back to dose

## Slide 34
DERIVED QUANTITIES

## Slide 35
EXTERNAL RADIATION
The operational quantities - gamma or beta gamma radiation dose rate – can be directly related to occupational exposure through time
An acceptable occupational radiation dose rate can also be derived for comparison with the measured dose rate
  e.g. 2.5 µSv/h for 2000h gives rise to 5 mSv annual exposure

## Slide 36
SURFACE CONAMINATION
The operational quantity of surface contamination cannot be directly related to occupational exposure.
A relationship between surface contamination and occupational exposure can be derived making assumptions about the radionuclide, route and duration of exposure.

## Slide 37
SURFACE CONAMINATION
An acceptable derived surface contamination level can be derived for comparison with the measured level.
More typically today, generic derived surface contamination levels are used
  DWL for beta, gamma 0.4Bq/cm2 and 0.04 Bq/cm2 for alpha surface contamination in transport regulations
  More recently, radionuclide specific contamination levels have also been derived.

## Slide 38
AIRBORNE CONAMINATION
The operational quantity of airborne contamination cannot be directly related to occupational exposure.
  Can derive a relationship
  Limit of intake is the intake of exposure to a single radionuclide (in Bq) of a specific physical and chemical form which would result in the annual committed effective dose limit.

## Slide 39
An acceptable airborne contamination level can then be derived for a specific radionuclide in a specific chemical form for comparison with the measured level.
  For example, Derived Air Concentrations (DAC)
AIRBORNE CONAMINATION

## Slide 40
AIRBORNE CONAMINATION
Airborne contamination can then be expressed as DAC fractions which correspond to fractions of Limits on Intake if breathed for a year.
  e.g. 0.3 DAC would result in 0.3 of an annual dose limit if breathed for a year
An estimate of intake can also be made in DAC-h by multiplying the fraction of DAC by the time of exposure.
  2000 DAC-h results in annual dose limit

## Slide 41
FINGERPRINTS
Most facilities handle a range of radionuclides.
In order to assess surface contamination or airborne activity, the nuclides present and their relative activity concentrations are required to be known.
This is called a fingerprint or nuclide vector
The fingerprint is used to help identify the workplace monitoring requirements and interpretation of results.
Notes: Though the design report may indicate the possible activity concentration of probable radionuclide, the precise value may be determined only if the facility is made operational. This finger print is useful to strengthen the protection measures.
This is particularly required for surface contamination and airborne contamination. For external radiation, need to know energy and emission as is a directly measured operational quantity

## Slide 42
FINGERPRINTS
Fingerprints may be determined from the following sources of information:
Safety report based on Modelling codes (irradiation of nuclear fuel in a neutron flux).
Operational history of the facility.
Radiochemical analysis of representative samples of the radioactive material.
Prior radiological safety assessmen.t
A combination of the above.
Notes: Without exception, when determining any fingerprint it is necessary to review the operational history of the radioactive materials, noting possible contamination mechanisms, the nature of radioactivity that might therefore be present and the physical distribution of this activity.
If it can be shown that sources of contamination are fairly limited and tend to be well-defined, it may be possible to develop a fingerprint from reference to historical records such as the radioactive species handled in connection with the materials to be characterised.
This is the case where only one material has been handled in a facility. For facilities which have seen wider use, historical records may provide a strong indication of the radionuclides which are present in the fingerprints. However, their relative activities may continue to be unknown without further investigation. It should not be assumed that the most recent history of the radioactive material is the only factor to be considered.
Where contamination is strongly linked to irradiated fuel or non-fuel materials, modelling codes such as FISPACT , FISPIN, or ORIGEN (for more information, see the TECDOC) may be used to predict the relative activities of radionuclides produced as a result of the irradiation of nuclear fuel in a neutron flux. However, in order for modelling codes to be useful, parameters which could critically affect the quality of their output must be known with confidence.
Will discuss more about fingerprints in the specific module on workplace monitoring on day 5 of the course

## Slide 43
FREQUENCY AND LOCATION OF MONITORING

## Slide 44
TYPES OF MONITORING
The type of monitoring (related to the purpose) is required to be known to determine the frequency and location.
Monitoring is classified into:
  Routine
  Task Related
  Special Monitoring
  Confirmatory Monitoring
Notes: These definitions are given in ICRP75 and the ORP

## Slide 45
TYPES OF MONITORING
Routine monitoring is:
  Intended to verify the suitability of continuing operations
  Intended to demonstrate that the working conditions remain satisfactory and meet regulatory requirements

## Slide 46
ROUTINE-EXTERNAL RADIATION
Typically a network of monitoring points will be established as a routine monitoring regime. Focus locations on:
  High occupancy
  Boundaries of controlled areas
  Expected operational activities
  High dose rates
  Where dose rates are liable to change
  Which would indicate early failure of equipment or process condition changes

## Slide 47
ROUTINE-EXTERNAL RADIATION
Frequency depends on occupancy and likelihood of change
  Where no substantial change expected and low occupancy, frequency can be low, e.g. monthly.
  Where no substantial change and occupancy high, or some change and low occupancy, then typical frequency would be weekly or daily.
  When sudden increases in dose rate may occur, or a worker may be significantly exposed then dose rates may require continuous monitoring.

## Slide 48
ROUTINE - SURFACE CONTAMINATION
Removable contamination is more typically monitored as it can affect occupational exposure or loss of control of contamination. Focus locations on:
  High traffic or high occupancy areas, boundaries of controlled areas and work areas
    Larger wipes more frequently for control
    Spot check with smaller wipes to confirm absence of contamination
    Direct monitoring of touch points or at lower frequency
  Areas less likely to be contaminated
    Use small wipes and direct monitoring

## Slide 49
ROUTINE - SURFACE CONTAMINATION
Frequency of loose contamination monitoring is typically high.
  High risk areas monitored daily
  Lower risk areas monitored weekly or monthly
  Very low risk areas monitored every 3 or 6 months or annually – typically this frequency is associated with direct monitoring.
Mix high frequency of loose contamination with lower frequency of direct in same area.

## Slide 50
ROUTINE - AIRBORNE CONTAMINATION
All potential airborne sources should be identified whether from routine or accidental conditions
Migration routes from potential sources should be identified
Identify potential exposure of personnel in terms of annual intake and decide whether continuous monitoring is required or sampling
  If greater than annual limit do continuous sampling
  If < 0.02 annual limit, then sampling not necessary
Notes: This will be discussed more in section on air sampling

## Slide 51
ROUTINE - AIRBORNE CONTAMINATION
Airborne monitoring locations chosen:
  To provide alert with least delay from onset of release – monitor as close as practical to a potential release, but also use secondary method (e.g. stack or further distance from source).
  To monitor migration pathways.
  Considering ergonomics, practicality, air flow patterns.
  Away from vents, inlets, doors and heat sources.
  Where necessary remotely.
  For installations, use fixed location samplers.

## Slide 52
EXAMPLES OF ROUTINE MONITORING
Frequency of airborne monitoring is typically high, depending on the potential airborne exposure and occupancy of the area. Typical frequencies are:
  Continuously
  Once per shift
  Daily
  Weekly
Notes: Added back in this last point as it is an important one I think
Discuss how often it is necessary to make the different measurements.
A hint can be:
Dose rate measurements (weekly)
Surface contamination measurements (weekly)
Airborne contamination measurements (continuous)
Hand-foot monitoring at exit (always)
Outside the laboratory (once per month)

## Slide 53
ROUTINE – WORKPLACE MONITORING
HFM
Balance
CAM
Basin
Basin
Basin
basin
1
2
3
4
5
6
7
8
9
10
Workbench
Workbench
Workbench
basin
Fume hood
Fume hood
Fume hood
Fume hood
Storage
Workbench
CAM
basin
CAM
Scheme of routine monitoring inside a lab
Notes: x# identifies the contamination measurement points
Additional points may be direct on taps, door handles, knobs, etc.
May take daily large area wipe measurements on exit from lab and at fume hoods.
May also monitor outside the laboratory
CAM: Continuous air monitor
HFM: Hand and foot monitor
Print out this overhead and add in dose rate monitoring points - at occupational areas – at fume hoods, at storage areas, at sink traps, at entrance to lab, on walls outside of lab

## Slide 54
TYPES OF MONITORING
Task related monitoring
  Conducted during a specific operation/task, provides data to support the immediate decisions on the management of the operation.
  Recommended to confirm the predictive estimates of radiological conditions and optimize protection.
  Where working conditions change and operations of limited duration (shorter than routine monitoring).
  Provides data to support immediate decisions on the operation (establish protective measures).
  Complements individual monitoring by providing useful indicators to predict doses.

## Slide 55
TASK RELATED MONITORING
The need for task related WPM shall be agreed with the operator and the radiation protection staff
The objectives and methods to be agreed on a case-by-case basis
Investigation levels need to be established, these can differ from routine monitoring

## Slide 56
TASK RELATED MONITORING - EXTERNAL
Locations defined for the operation.
Use installed monitoring equipment to give warning – could be routed to a control room.
Outside controlled area to confirm dose rates during tasks are acceptable (e.g. outside of waste build up areas). For beta radiation, could be related to areas where hands may come into contact either before touching or to verify the dose rates.

## Slide 57
TASK RELATED MONITORING - CONTAMINATION
Verify contamination levels as work proceeds and to assist in contamination control.
Floors adjacent to work areas could be monitored during and after work.
Tools may also be monitored or gloves.
Outside of containment before removal from the area.
For airborne monitoring, may use portable air samplers for the duration of the work activity

## Slide 58
EXAMPLES
Filter change operations
Non-routine maintenance operations
      Clean-up operations
Breakdown activities

## Slide 59
TASK RELATED MONITORING - EXAMPLE
Decontamination operation carried out in a specially erected tent to ensure prevention of spread of contamination. Airborne contamination monitoring is task related.
Notes: Example shows a temporary tented area used for drilling operations. Image shows a mobile filtration system, with a portable air monitor sampling into the tented area. A second monitor is located outside the area to detect the potential spread of contamination. If a continuous monitor is not used, then the air sampler could be changed each time the operation is conducted or on a shiftly basis. But continuous monitoring will give direct results, without need for decay counting.

## Slide 60
TYPES OF MONITORING
Special monitoring:
Investigative in nature
Covers a situation where there is inadequate information to demonstrate adequate control
The WPM programme should define when special monitoring is to be carried out
The WPM programme should define how and how quickly the results are to be reported, and relevant investigation levels.

## Slide 61
SPECIAL MONITORING
Workplace monitoring is done:
Prior to the initial use of the radiation installation, or as soon as possible after radiation sources are brought into the area.
Whenever changes in procedures, equipment, shielding, or sources have occurred.
Periodically to detect the effect of long-term changes in equipment, environment and work habits.
When an accident or incident is suspected or after it has occurred.

## Slide 62
SPECIAL MONITORING - EXAMPLES
Commissioning
Decommisioning
  Identify gamma hot spots, clearance surveys
Modifications
  Identifying where shielding should go or amount
Abnormal conditions
  Identify hot spots to prioritize removals
  Identify man access routes or dose uptake minimization
  Identify operational challenges, for example blockages
  Extent of spillage
  Finding lost sources

## Slide 63
SPECIAL MONITORING - EXAMPLES
Example: Hotspot identification using gamma imaging camera.

## Slide 64
TYPES OF MONITORING - CONFIRMATORY
Confirmatory monitoring is performed where there is a need to check assumptions made about exposure conditions.
Used to verify the robustness of shielding and other protection measures.
Used, for example, to confirm the effectiveness of protective measures.

## Slide 65
SUMMARY
Workplace monitoring is necessary to identify the potential exposures. It involves measurement of
  External exposure measurements
  Surface contamination measurements
  Air activity monitoring
Workplace monitoring helps to control the exposure by taking suitable protective measures.