## Slide 1
LESSON 11:
NOBLE GAS MONITORING

## Slide 2
Noble Gas Monitoring
Techniques
Equipment
  Examples of Noble gas Monitors
Calibration and Testing
Factors Influencing Volumetric Activity

## Slide 3
Noble Gas Monitoring
Some of the important radioactive noble gases, which are produced in the nuclear reactors, are;
Activation product 41Ar,
Fission product 85Kr, 85mKr, 87Kr, 88Kr and 133Xe, 133mXe, 135Xe and 138Xe.
A release of noble gases into the working atmosphere indicates a breach of containment including possible clad rupture.
Noble gases are inert and insoluble deliver;
external dose due to immersion
lung dose
Detecting the Fission Product Noble Gases (FPNG) limits and reduce the risk of exposures.

## Slide 4
Techniques
Noble gases cannot be trapped by filters.
By special techniques such as cryogenic techniques they can be collected on suitable medium.
They are monitored either in real time or by taking a grab sample which is subsequently counted by an appropriate detector in the laboratory.
In real time monitoring the sampled air passes through a measurement cell.
The measurement cell may be the detector itself, for example an ionization chamber, or a proportional counter of flow through type, or a sampling cell equipped with a plastic scintillator.

## Slide 5
Techniques
Gamma radiation background in the vicinity is compensated by a sealed identical detector as used for the flow cell.
The reading may be either direct or after software treatment to compensate for the activity of the radon and to improve the detection limit of the instrument.
The readings have to be provided in derived units (e.g. Bq/m3).

## Slide 6
EQUIPMENT

## Slide 7
Equipment
A data processing software where appropriate.
Alarm device
A closed detector dedicated to compensation of background gamma radiation.
The associated electronics.
An air sampling circuit.
A measurement cell.
A typical Noble gas monitor consists of:

## Slide 8
Points to consider for Equipment
In order to avoid error in the gas measurement the sampling line should include upstream of the detector:
  A humidity trap (e.g. desiccant)
  A pre-filter in order to trap particulates
    The pre-filter should not decrease the flow rate or the pressure inside the measurement cell over the limits specified by the manufacturer.
  A charcoal filter
    if iodine is suspected to be present in the area. Iodine isotopes are decay precursors of noble gas (Xe) so trapping the iodine should not result in a false indication of the noble gas measurement.

## Slide 9
Points to consider for Equipment
The measurement cell should be equipped with a pressure measurement device in order to verify that the pressure is in the normal range of operation.
The instrument should be chosen depending on its detection limit and the response time required for each particular case.

## Slide 10
Location of Sampling Head
The most important criterion in selecting the location of sampling head is the representativeness of the sample.
The most appropriate location for the sampling head may be selected by the determination of transfer coefficient using a gaseous tracer like helium.

## Slide 11
Sampling Line
If the detector is not directly installed in the area of interest, the gas may be sampled and drawn from the controlled area to the measurement cell.
In this case, the time delay between the sampling point and the measurement cell should be taken into account due to the short half lives of the isotopes.
Noble gases are not subject to trapping on the wall of the sample line.

## Slide 12
EXAMPLES OF MONITORS

## Slide 13
Measurement Cell Equipped with Silicon Diode or Scintillation Detector
Good points:
Small size monitor.
Good detection limit: a chamber of 100 ml has a detection limit similar to an ionisation chamber of 10 Litre.
Usable with high gamma ray background due to the small size and possibility of efficient shielding without significant excessive weight.
Low influence on sensitivity of the variation of temperature.

## Slide 14
Insensitive to humidity.
Low cost.
Low maintenance cost.
Points to consider for Equipment
Measurement Cell Equipped with Silicon Diode or Scintillation Detector
    Weakness:
Sensitive to pressure: the response is proportional to the pressure inside the measurement chamber.
Interference of gamma. So the detector should be shielded.

## Slide 15
Measurement Cell Equipped with Silicon Diode or Scintillation Detector
Check for:
Detection efficiency.
Calibration of the pressure measurement device.
Clogging of the inlet filter.

## Slide 16
Ionization Chamber
Good point:
  Low maintenance cost.
  Continuous monitoring of
  Noble gases in workplace
Weaknesses:
  Detector larger than measurement cell equipped with silicon diode.
Courtesy:Doza

## Slide 17
Ionization Chamber
Poor detection limit with reasonable size.
Usable where gamma radiation background is low.
Due to the size of the detector it is quite difficult to install a shielding.
Sensitive to temperature.

## Slide 18
Sensitive to pressure: the response is proportional to the square of the pressure.
Very sensitive to humidity: it is necessary to include a humidity trap at the inlet of the monitor.
Sensitive to electromagnetic field: the large chambers should be well screened electrically.
Ionization Chamber

## Slide 19
Ionization Chamber
Check for:
Detection efficiency.
Calibration of the pressure measurement device.
Clogging of the inlet filter.
Saturation of the water trap.

## Slide 20
Flow through Proportional Counter
Good points:
Small detector.
Due to the small size detector
can be well shielded.
Low maintenance cost.
Proportional counter for WPM.
Courtesy: Thermo scientific
Weaknesses:
Poor detection limit.
Sensitive to temperature.
Sensitive to pressure: the response is proportional to the square of the pressure.
Sensitive to humidity.

## Slide 21
Flow through Proportional Counter
Check for:
Detection efficiency.
Calibration of the pressure device.
Clogging of the inlet filter.
Saturation of the water trap.

## Slide 22
Examples of Modern Noble Gas Monitors
Detector: HPGe detector
Detector: Flow through ionisation chamber

## Slide 23
EXAMPLES OF MONITORS

## Slide 24
Type Test Report and Certificate
As per IEC 60761-1, the manufacturer shall provide, with each equipment, a certificate giving the following information:
radioactive noble gas or gases for which the assembly is designed;
detector type and general characteristics;
response as a function of the volumetric activity per unit under reference conditions;
response to the check source;
response to the other radioactive gases;
response to interfering gases of concern.

## Slide 25
Calibration and Functional Testing
Each air monitoring instrument is required to have a valid calibration certificate before use
Functional testing should be defined by the manufacturer and usually includes testing the detection efficiency with a radioactive source at regular intervals.
The detection efficiency measured with a checking source is only an equipment functional check.
Unless otherwise indicated by the manufacturer, this value should not be used as a calibration factor.
The pressure measurement device should be verified at regular time intervals by comparing with a calibrated pressure measuring instrument

## Slide 26
FACTORS INFLUENCING VOLUMETRIC ACTIVITY

## Slide 27
Factors Influencing Activity Concentration
The energy of the radionuclide detected.
The pressure of the gas inside the measurement cell.
The presence of natural radioactive gas such as radon.
Determination of the volumetric activity of the noble gas may be influenced by three parameters:

## Slide 28
Energy of Radionuclide Measured
The detection efficiency is affected by the radiation type and energy of the radionuclide measured.
The response of the instrument to different energies should be measured during type testing and provided by the manufacturer.

## Slide 29
Pressure of the Gas Measured
The volumetric activity is normally indicated for temperature and pressure of the gas at the inlet of the assembly.
  The manufacturer should identify if any correction is required and the method to address this.
The manufacturer should indicate the influence of other radioactive gases such as radon on the indication of the monitor.