## Slide 1: Cardiac CT - radiation doses, dose management and practical issues
L 11
Notes: Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Answer True or False
Patient dose from a cardiac CT is equivalent to 20 chest conventional radiographies.
In cardiac CT the radiation dose to the different organs is very similar to the catheterization procedures.
For cardiac CT, patient doses are typically measured in Gy•cm2.

## Slide 3: Educational Objectives
To understand the principles and the technology of CT for cardiology examinations
To understand the dosimetric quantities for patients in CT and the factors influencing these doses

## Slide 4: Number of CT Procedures in US
IMV Benchmark Report on CT, 2006
Notes: <number>
<footer>
<header>
<date/time>

## Slide 5: Distribution of CT procedures 50.1 million in 2003
IMV Benchmark Report on CT, 2004
HCAP: ~70% of all CT procedures
Notes: <number>
<footer>
<header>
<date/time>

## Slide 6: Introduction
Computed Tomography (CT) was introduced into clinical practice in 1972 and revolutionized X ray imaging by providing high quality images which reproduced transverse cross sections of the body.
Tissues are therefore not superimposed on the image as they are in conventional projections
The technique offered in particular improved low contrast resolution for better visualization of soft tissue, but with relatively high absorbed radiation dose

## Slide 7: Computed Tomography
CT uses a rotating X ray tube, with the beam in the form of a thin slice (about 1 - 10 mm)
The “image” is a simple array of X ray intensity, and many hundreds of these are used to make the CT image, which is a “slice” through the patient

## Slide 8: The CT Scanner

## Slide 9: Helical Scan Principle
Scanning Geometry
Continuous Data Acquisition and Table Feed
X ray beam
Direction of patient
movement

## Slide 10: Multislice CT collimation
0,5mm
1mm
2,5mm
5mm
Multislice CT: several slices can be collected simultaneously

## Slide 11: Pitch factor
Inter-slice distance is defined as the couch increment minus nominal slice thickness. In helical CT the pitch factor is the ratio of the couch increment per rotation to the nominal slice thickness at the axis of rotation. In clinical practice the inter-slice distance generally lies in the range between 0 and 10mm, and the pitch factor between 1 and 2.
The inter-slice distance can be negative for overlapping scans which in helical CT means a pitch < 1.
(EUR 16262: European Guidelines on Quality Criteria for CT)

## Slide 12: Pitch redefined for MDCT
I - Table feed (mm/rotation)
W - Beam width (mm)
I
T
W
Beam Pitch =
I
W
Detector Pitch =
I
T
T - Single DAS channel width (mm)
N - Number of active DAS channels
Beam Pitch =
Detector Pitch
N
=
I
N*T
= Pitch†
† IEC Part 2-44, 2003
Notes: <number>
<footer>
<header>
<date/time>

## Slide 13: Typical Effective Dose Values for CT
Head CT 1 - 2 mSv
Chest CT 5 - 7 mSv
Abd & Pelvis CT 8 - 11 mSv
Average U.S. background radiation per year
3.6 mSv
Typical chest X ray ~ 0.1 - 0.2 mSv
Ca-Scoring 1.5 - 5.0 mSv
Cardiac CTA 10 - 25 mSv
Notes: <number>
<footer>
<header>
<date/time>

## Slide 14: Why doses are high in MDCT?
Shorter scan times and thinner slices requires higher tube current to maintain good image quality
For cardiac CT, excessive tissue overlap (low pitch) is often required to reduce motion artifacts
Translates to higher patient dose!
Notes: <number>
<footer>
<header>
<date/time>

## Slide 15: Factors influencing MDCT radiation dose and image quality
      Tube current (mA)
      X ray on time
      Pitch
      mAs or effective mAs
      X ray beam energy (kVp and filtration)
      Slice thickness
      Geometric and detector efficiency
      Beam filters
      Reconstruction algorithms, …
      Patient size
Notes: <number>
<footer>
<header>
<date/time>

## Slide 16: Dose (at the detector) vs. Noise
20 nGy per frame
150 nGy per frame
240 nGy per frame
Notes: <number>
<footer>
<header>
<date/time>

## Slide 17: Effect of X ray beam energy
120 kVp
135 kVp
Notes: <number>
<footer>
<header>
<date/time>

## Slide 18: Effect of Pitch on Dose and Image Quality
P = 0.83
CTDI = 37 mGy
P = 1.48
CTDI = 20.6 mGy
45% lower
P = 0.64
CTDI = 47.8 mGy
30% higher
Notes: <number>
<footer>
<header>
<date/time>

## Slide 19: Radiation Measurements

## Slide 20: CT Dosimetry
CTDI100
MSAD
CTDIw
CTDIvol
DLP
Effective
Dose
Notes: <number>
<footer>
<header>
<date/time>

## Slide 21: Experimental Setup
Radcal 1015C Electrometer, with CT ion chamber in body phantom
Notes: Polymethyl methacrylite phantom
CT chamber placed 1 cm underneath the 32 cm diameter CT phantom

## Slide 22: Typical dose distribution in CT
100
Body 32 cm
Head 16 cm
100
100
100
100
50
100
100
100
100
Dose uniform on surface and decreases towards center
Notes: <number>
<footer>
<header>
<date/time>

## Slide 23: CTDIW (weighted average of center and periphery doses)
CTDIvol (Pitch factor is considered)
CTDIw=(2/3) CTDIedge+(1/3) CTDIcenter
CTDIvol=(1/pitch)CTDIw
Notes: <number>
<footer>
<header>
<date/time>

## Slide 24: Dose Length Product (DLP)
Indicates radiation dose of entire CT exam
Includes number of scans and scan width
DLP = CTDIvol (mGy) • scan length (cm)
Displayed on monitors
Notes: <number>
<footer>
<header>
<date/time>

## Slide 25: Effective doses for Cardiac Imaging
Procedures | Modality | Effective Dose (mSv)
Ca Scoring | EBCT | 1.0 - 1.3
 | MDCT | 1.5 - 6.2*
CTA | EBCT | 1.5 - 2.0*
 | MDCT | 6.7* - 25.0
Cardiac SPECT w Tc-99m or Tl-201 |  | 6.0 - 15.0
CA (diagnostic only w fluoroscopy) |  | 2.1* - 6.0
Chest x-ray |  | 0.1
*Hunold P, et al., Radiology, 2003
Notes: <number>
<footer>
<header>
<date/time>

## Slide 26: CT Dose Modulation

## Slide 27: Recent Advances: CT Dose Reductions
X ray attenuation lower in AP and higher in lateral projection
Higher attenuation
high mA
Low attenuation
low mA
However, CT doses are uniform on the surface and decrease radially towards center
Various dose reduction options are being considered
Notes: <number>
<footer>
<header>
<date/time>

## Slide 28: Dose Reduction Options
Dose reduction based on patient anatomy
Lower mA in AP, higher mA in lateral directions
1 GE, 2 Toshiba and 3 Siemens MDCT
200 mA
150 mA
130 mA
150 mA
180 mA
210 mA
200 mA
170 mA
180 mA
Methods
Patient attenuation measured during scout scan (AP & Lat) and alter mA for each gantry rotation (Smart mA1, Real AEC2) or “on-the-fly” (Care dose3)
Dose reduction of 20-40% is possible
Notes: <number>
<footer>
<header>
<date/time>
Other than changing technique factors, manufacturers are also trying to reduce radiation dose by adopting several dose reduction options

## Slide 29: ECG-controlled tube current modulation
Tube current monitoring ECG signal is lowered in systole region and ramped up during diastole region
Mean radiation reduction of up to 45% has been reported*
*Jakobs, et.al. Euro Radiol, 2002
Notes: <number>
<footer>
<header>
<date/time>

## Slide 30: Impact of Dose Modulation: Chest CT
Radiation dose: Lateral: 16% increase, AP: 25% reduction
*Mahesh, Kamel & Fishman, Evaluation of ‘CareDose’ on Siemens Sensation 64 MDCT scanner
Notes: <number>
<footer>
<header>
<date/time>

## Slide 31
Notes: <number>
<footer>
<header>
<date/time>

## Slide 32: Dual Source CT* (DSCT)
Two X ray tubes positioned at right angle
Two detector arrays opposite to X ray tubes
Temporal resolutions less than 100 ms is possible by combining data from one-fourth of data acquired by two detectors
*Siemens ‘Definition’ at Johns Hopkins, 2006
Tube A
Tube B
Notes: <number>
<footer>
<header>
<date/time>

## Slide 33: Conclusions
Radiation dose estimates for CT exams are best expressed as CTDIvol (mGy), DLP (mGy.cm) and Effective Dose E (mSv).
With increasing number of CT scans, there is concern about radiation burden to general public.
Notes: <number>
<footer>
<header>
<date/time>

## Slide 34: Conclusions
Physicians referring or performing cardiac CT exams should understand radiation doses associated with various protocols and should be able to justify the appropriateness of CT exam
“Genie is out of the box, it is now left to the user how to tame it”
Notes: <number>
<footer>
<header>
<date/time>

## Slide 35: Answer True or False
Typical patient dose values for cardiac CT angiography are in the range of 1-2.5 mSv.
Shorter scan times and thinner slices require higher tube current to maintain good image quality.
Patient dose increases if pitch factor increases (if all other parameters are maintained constant).

## Slide 36: Answer True or False
When patient radiation dose for CT is given as 500 mGy•cm, it is understood that skin dose is 500 mGy.
During CT cardiac angiography, tube current monitoring ECG signal is lowered in systole region and ramped up during diastole region.