## Slide 1: X ray production and Angiography Equipment
L 4

## Slide 2: Answer True or False
Radiation collimation is not necessary in the new flat panel detectors.
The extra filtration in some of the new X ray systems is used to improve image quality.
Radiation dose to the patients can only be measured by a specialized person standing in the catheterization laboratory during the procedure.

## Slide 3: Answer True or False
When changing the field of view from 23 cm to 18 cm, it is necessary to collimate the radiation field because it is not done automatically.
Most of the modern X ray systems for cardiology include a transmission ionization chamber to measure staff dose.
Most of the modern X ray systems have the capability to produce patient dose report at the end of the procedure.

## Slide 4: Educational objectives
X ray generation and imaging (fundamentals)
What are equipment standards for catheterization equipment (FDA, IEC), particular needs for pediatric patients equipment?

## Slide 5: X ray Generation and image formation

## Slide 6: Filmless X ray Cinefluorographic Unit
Notes: The unit consists of a patient positioning system, an X-ray source, an X-ray image
detector, and a digital video image processor, recorder and display system. X-rays are produced in the X-ray tube from highly controlled electrical power
that is applied by the X-ray generator. X-rays that penetrate the patient form an X-ray image that is detected and converted to a visible light image by the
image intensifier. The visible light image is detected by the video camera and converted to a digital video signal that is processed and displayed as a visible
light image on video monitors. Feedback circuitry from the digital video processor communicates with the X-ray generator. This enables modulation of
X-ray output to achieve appropriate subject penetration by the X-ray beam and, accordingly, proper image brightness. X-ray systems that have flat panel
detectors rather than image intensifiers do not have video cameras as the flat-panel detector produces a digital video image directly without the intermediate
visible light stage

## Slide 7
Operator / Automatic control of
tube voltage (kVp)
tube current (mA)
exposure duration
Provide
environment
source of electrons (cathode)
source of X rays (anode)
induction motor to rotate anode
heat dissipation
electrical insulation
x-ray shielding
Notes: <number>
<footer>
<header>
<date/time>
FIGURE 3. X-ray generator and x-ray tube components are illustrated. The x-ray generator provides operator control of the radiographic techniques, including tube voltage (kVp), tube current (mA), and exposure duration, and delivers power to the x-ray tube. The x-ray tube provides the environment (evacuated x-ray tube insert and high-voltage cable sockets), source of electrons (cathode), source of x-rays (anode), induction motor to rotate the anode (rotor/stator), transformer oil and expansion bellows to provide electrical and heat build-up protection, and the tube housing to support the insert and provide protection from leakage radiation.

## Slide 8
A typical X-Ray Tube
Current to cathode (mA)
no. of electrons liberated
no. of X ray photons
Voltage across
(kilovolts-peak [kVp])
energy of electrons
energy of X ray photons

## Slide 9
Typical Photon Energy Spectrum from
a Machine Operating at KV = 80
(from The Physical Principles of Medical Imagings, 2Ed, Perry Sprawls)

## Slide 10
Comparison of Photon Energy Spectra
Produced at Different kVp Values
(from The Physical Principles of Medical Imagings, 2Ed, Perry Sprawls)

## Slide 11
Filter
Notes: <number>
<footer>
<header>
<date/time>
The addition of filtration significantly alters the shape of the x-ray spectrum, as shown below. Since filtration selectively absorbs the lower energy photons, it produces a shift in the effective energy of an x-ray beam. The figure below compares an unfiltered spectrum to spectra that passed through 1-mm and 3-mm filters. It is apparent that increasing the filtration from 1 mm to 3 mm of aluminum produces a noticeable decrease in the number of x-ray photons. It should be observed, however, that most of this decrease is in photons with energies less than approximately 40 keV. These are the photons with a low probability of penetrating a typical patient and contributing to image formation. They do, however, contribute to patient exposure.
Adding filtration increases the penetration (HVL) of an x-ray beam by removing the low-energy photons.

## Slide 12
Transmission chamber and collimators
Other important elements of the X ray systems

## Slide 13: To obtain the images …
Two technologies are used:
  Image intensifier
  Flat panel detector

## Slide 14
Motorized Iris
Video Camera
Image Intensifier
DETECTOR
Photons
Cesium Iodide (CsI)
Light
Amorphous Silicon Panel
(Photodiode/Transistor Array)
Digital Data
Electrons
Read Out Electronics
Photons
Cesium Iodide (CsI)
Light
Photo-cathode
Video Signal
Electrons
Output screen
Light
CCD or PUT
Electrons
Readout Electronics
1
3,000
400
400,000
2,400
Particles #
Image
Intensifier
Flat-panel

## Slide 15: Fields of view … magnification
Imaging detectors allow different fields of view (magnification) to improve the spatial resolution.
This magnification usually increases the skin dose to the patient.
Because only a part of the detector is used during magnification, the radiation field is automatically collimated to the visualized area.

## Slide 16
X ray room dosimetric information

## Slide 17: High filtration
The introduction of additional filtration in the X ray beam (commonly copper filters) reduces the number of low energy photons and, as a consequence, saves skin dose for the patients.

## Slide 18: Reduction of Radiation Exposure with extra filtration
Additional Cu filters can reduce the skin dose by more than 70%.
Some systems offer variable extra filtration (0.2 mm - 0.9 mm) that is automatically set according to patient weight and angulation of the C-arm.
Automatic filter insertion try to keep the dose as low as possible without degrading image quality.

## Slide 19: Pulsed fluoroscopy
Pulsed fluoroscopy can be used as a method of reducing radiation dose, particularly when the pulse rate is reduced.
But … pulsed fluoroscopy does not mean that dose rate is lower in comparison with continuous fluoroscopy!!.
Dose rate depends on the dose per pulse and the number of pulses per second.

## Slide 20
Collimation

## Slide 21: Wedge filter
Wedge filter. GE Advantx X ray system

## Slide 22: Importance of wedge filters
The wedge filter has not been used to obtain this cine series. Note the important difference in contrast.
The wedge filter has been used to obtain this cine series.

## Slide 23: Reduction of Radiation Exposure with virtual collimation
Radiation-free Collimation
Manipulation of diaphragms in Last Image Hold
No fluoroscopy required

## Slide 24: Antiscatter grid
To avoid that scatter radiation from the patient reaches the detector producing a degradation in image quality
Shall be easily removable in new X ray systems (according with IEC standards)

## Slide 25
Antiscatter grid
Flat panel system
Siemens Axiom
Notes: <number>
<footer>
<header>
<date/time>

## Slide 26
Notes: <number>
<footer>
<header>
<date/time>

## Slide 27: Equipment standards for Cath Lab

## Slide 28: X ray Equipment Standards and Regulations
Standards are consensus guides from the manufacturing community, not regulatory
Several groups set standards regarding equipment, e.g., International Electrotechnical Equipment (IEC)
Apply to electrical, mechanical, and radiation safety
Apply to equipment at time of manufacture and installation
Notes: Explain the differences between the various standards setting groups, etc. What purpose does IEC serve? What about ISO? How do these impact the cardiologist in day-to-day practice?

## Slide 29: What to look for while establishing a cath. lab.
If the relevant Standards are fulfilled
If a medical physicist is available
If radiation protection tools are available
If patient dose measuring and recording system is available
If acceptance tests, commissioning and quality assurance programme have been foreseen
Notes: Explain the differences between the various standards setting groups, etc. What purpose does IEC serve? What about ISO? How do these impact the cardiologist in day-to-day practice?

## Slide 30: What to look for while establishing a cath. lab.
If the X rays system selected is appropriate for the procedures to be carried out in the catheterization laboratory
If some other relevant information described in ACC/AHA Guidelines and AAPM-70 (described in this lecture) have been taken into account
Notes: Explain the differences between the various standards setting groups, etc. What purpose does IEC serve? What about ISO? How do these impact the cardiologist in day-to-day practice?

## Slide 31: Limitation in entrance exposure rate
Federal Register: May 19, 1994.
21 CFR Part 1020.
Federal Performance Standard for Diagnostic X-Ray Systems and Their Major Components;
Final Rule.
DEPARTMENT OF HEALTH AND HUMAN SERVICES
Food and Drug Administration

## Slide 32: Limitation in entrance exposure rate
The Standard for Diagnostic X Ray Systems (May 19, 1994), limits the entrance exposure rate of fluoroscopic X ray systems during normal fluoroscopy to 10 R/min unless an optional high-level control (HLC) is activated.
If HLC is activated, the entrance exposure rate must be limited to 20 R/min.
The entrance exposure rate limits do not apply during the recording of images.

## Slide 33
Proposed Rule
December 10, 2002

## Slide 34
mGy (total)
mGy/min
(at 15 cm from the isocenter towards the x-ray source)
Fluoroscopy time

## Slide 35
Fluoroscopic equipment manufactured on or after May 19, 1995:
Shall not be operable if AKR is higher than 88 mGy/min (10 R/min).
Exceptions:
When a mode a high-level control is activated: 180 mGy/min (20 R/min). A continuous signal audible to the fluoroscopist shall indicate that the high-level control is being employed.
During the recording of images (archiving of fluoroscopic or radiographic images in analog format with a video-tape or video-disc recorder does not qualify as an exception).
Limits:
88 mGy/min
180 mGy/min

## Slide 36
IEC Standard 2000

## Slide 37: IEC standard on Interventional Radiology
Radioscopically guided invasive (and interventional) procedures
Interventional reference point
Isokerma maps shall be provided
The anti-scatter grid should be removable without the use of tools
Dosimetric indications: reference air kerma rate, cumulative reference air kerma. cumulative area kerma product, (shall be accurate to within  50 %)
Supplementary indications: cumulative time of radioscopy, cumulative number of radiographic irradiations, integrated reference air kerma

## Slide 39
Collimation: Dual-shape collimators incorporating both circular and elliptical shutters may be used to modify the field for cardiac contour collimation. Partially absorbent contoured filters are also available to control the bright spots produced by the lung tissue bordering the heart.

## Slide 40
Philips systems

## Slide 41
Example of the influence of wedge filter in the skin dose

## Slide 42
Image intensifiers. Because of the necessity of imaging large fields (e.g., for ventriculography, aortography) as well as small fields (coronary arteries), multimode (double or triple) cesium iodide image intensifiers are recommended. Formats available vary with the manufacturer but are typically 9 in/ 6 in/4.5 in (9/6/4.5), 9/6, 10/4, and 9/5.

## Slide 43: ACC/SCAI Clinical Expert Consensus Document on Catheterization Laboratory Standards JACC Vol. 37, No. 8, June 2001:2170-214
This document addresses provider and patient exposure in the cardiac catheterization laboratory, with detailed information on ionizing radiation, its risks, and reducing unnecessary radiation exposure. The document addresses the key components of a quality assurance program focused on individual and laboratory outcomes for both diagnostic and interventional cardiac catheterization laboratories: clinical proficiency, equipment maintenance and management, a quality improvement process, and radiation safety.

## Slide 44
A freely movable lead glass or acrylic shield suspended from the ceiling should be used. Its sterility may be maintained by using disposable plastic covers.
Each procedure room should have a detailed determination of exposure levels performed by a qualified radiation physicist.
There is a tendency in the busy laboratory to assign a low priority to preventive maintenance and quality assurance inspections.

## Slide 45
Coronary interventionalists must also have a thorough knowledge of specialized equipment, techniques, and devices used to perform PCI competently

## Slide 46: AAPM-70 (2001)
The generator should be capable of generating 80 to 100 kilowatts (kW) of power.
The generator design should result in “square wave” kVp pulses to achieve optimum patient dose savings.

## Slide 47: AAPM-70 (2001)
For adult studies, a 9 to 11 inch (23 to 27 cm) size is used.
Pediatric cardiac studies use smaller FoVs due to the small size of the pediatric heart.
The 4.5 inch (11 cm) FoV would be commonly employed for most pediatric imaging studies.

## Slide 48: AAPM-70 (2001). Pediatrics.
Cine frame rate capability should extend up to at least 60 fps for small children.
The generator should support an X ray tube with a minimum of three focal spots. Patients up to 3 to 4 years old can be imaged with an 0.3 mm focal spot size, and patients up to
8 to 9 years old can be imaged with cine using an 0.6 mm focal spot.

## Slide 49: Some self evaluation questions …

## Slide 50: Answer True or False
Collimation of the radiation field is always automatically made by the X ray system.
Some new systems include the feature of “virtual collimation” meaning that unnecessary radiation is removed numerically by the software.
Filtration in the X ray tube should be as low as possible.

## Slide 51: Additional information

## Slide 52: IAEA survey 2001-2003
X ray systems evaluated:
9-15 from 5 countries

## Slide 53: Conclusions from the IAEA survey
Patient dose and image quality depend largely on the settings made at the commissioning of the radiological equipment.
For different systems and different operation modes, entrance air kerma can increase by a factor of 20 (including electronic magnification) for the same patient thickness.

## Slide 54: Conclusions from the IAEA survey
Increasing phantom thickness increases dose by an additional factor of up to 12.
Differences in radiation doses from the evaluated systems show a potential for dose reduction whilst maintaining image quality.

## Slide 55: Importance of testing X ray equipment
Characterization of the X ray system, that should be part of the acceptance and status tests, should inform cardiologists about the dose rates and dose/frame for the different operation modes and for the different patient thicknesses. Image quality shall also be evaluated.
Regular constancy checks should verify if important changes may have occurred.

## Slide 56: Examples of patient dose reports

## Slide 57
Example of the data included in the study report (Siemens)

## Slide 58
Example of the data included in a dosimetric report:
Philips Integris 5000.
Coronary angiography
65% cine; 35% fluoroscopy
13 series, 728 frames
1,54 Gy.cm2/min
0,368 Gy.cm2/10 fr
1 min fluoroscopy =
39 fr = 3 s cine

## Slide 59: Examples of information contained in the DICOM header

## Slide 60
Notes: <number>
<footer>
<header>
<date/time>

## Slide 61
Notes: <number>
<footer>
<header>
<date/time>