## Slide 1: Patient Dose Management
L 5a

## Slide 2: Are these statements “True” or “False”?
Typically about 40% of radiation entering patient body penetrates through to form X ray image.
You are likely to receive more scattered radiation when performing cardiac catheterization for an obese person, compared to one done for a thin person.
During coronary angiography, patient receives more radiation dose in AP (anterior-posterior) projection, compared to LAO cranial angulation.

## Slide 3: Educational Objectives
Understand the various factors affecting radiation dose to patient
Understand operator’s role in patient dose management
How to manage patient dose using procedural and equipment factors

## Slide 4: X ray Image Formation

## Slide 5
determine energy of electrons  energy of X-ray photons
determine number of electrons  number of X ray photons
Notes: <number>
<footer>
<header>
<date/time>
FIGURE 3. X-ray generator and x-ray tube components are illustrated. The x-ray generator provides operator control of the radiographic techniques, including tube voltage (kVp), tube current (mA), and exposure duration, and delivers power to the x-ray tube. The x-ray tube provides the environment (evacuated x-ray tube insert and high-voltage cable sockets), source of electrons (cathode), source of x-rays (anode), induction motor to rotate the anode (rotor/stator), transformer oil and expansion bellows to provide electrical and heat build-up protection, and the tube housing to support the insert and provide protection from leakage radiation.

## Slide 6
X-ray tube
Photons entering the human body will
either penetrate, be absorbed,
or produce scattered radiation

## Slide 7
To create image some X rays must interact with tissues while others completely penetrate through the patient.
(1) Spatially uniform beam enters patient
(2) X rays interact in patient, rendering beam non-uniform
(3) Non-uniform beam exits patient, pattern of non-uniformity is the image
Reproduced with permission from Wagner LK and Archer BR. Minimizing Risks from Fluoroscopic Radiation, R. M. Partnership, Houston, TX 2004.

## Slide 8
Image Contrast
No object image
is generated
Object image
is generated
Object silhouette
with no internal
details
Notes: Image contrast. Differential absorption of the X-ray beam by different parts of an object renders its internal structure visible in an X-ray transmission image. X-ray photons are represented as arrows. The left-hand example shows an object that allows all X-ray photons to pass through it unattenuated. In this circumstance, no image of the object is generated. The center example shows an object, different parts of which absorb different fractions of the incident photons modulating the beam intensity and generating an X-ray image of the object. The right-hand example is completely opaque to X-ray photons and absorbs all of them. The X-ray image of this object would be a silhouette with no definition of the object’s internal structure.

## Slide 9: Detector Dose and Patient Dose
Detector Dose
  The total X ray dose, which reaches the detector
  Contributes to the image quality, and should therefore be as high as possible.
  Significantly lower than the patient dose (~ 1% of patient dose)
Patient Dose
  The total X ray dose applied to a patient
  Harmful for both the patient and the surrounding staff in scattered radiation.
  Therefore, patient dose should be as low as possible
X-ray tube
Detector dose
Patient dose

## Slide 10
Because image production requires that beam interact differentially in tissues, beam entering patient must be of much greater intensity than that exiting the patient.
Beam entering patient typically ~100x more intense than exit beam
As beam penetrates patient, X rays interact in tissue causing biological changes
Only a small percentage (typically ~1%) penetrate through to create the image.
Reproduced with permission from Wagner LK and Archer BR. Minimizing Risks from Fluoroscopic Radiation, R. M. Partnership, Houston, TX 2004.

## Slide 11
Lesson:
Entrance skin tissue receives highest dose of x rays and is at greatest risk for injury.
Beam entering patient typically ~100x more intense than exit beam in average size patient
As beam penetrates patient X rays interact in tissue causing biological changes
Only a small percentage (typically ~1%) penetrates through to create the image.
Reproduced with permission from Wagner LK and Archer BR. Minimizing Risks from Fluoroscopic Radiation, R. M. Partnership, Houston, TX 2004.

## Slide 12: Factors Affecting Radiation Dose

## Slide 13: Factors that influence Patient Absorbed Dose
Patient-related factors
Equipment-related factors
Procedure-related factors

## Slide 14: Factors that influence Patient Absorbed Dose
Patient-related factors
  Patient body weight and habitus

## Slide 15: Factors that influence Patient Absorbed Dose
Equipment-related factors
  Movement capabilities of C-arm, X ray source, image receptor
  Field-of-view size
  Collimator position
  Beam filtration
  Fluoroscopy pulse rate and acquisition frame rate
  Fluoroscopy and acquisition input dose rates
  Automatic dose-rate control including beam energy management options
  X ray photon energy spectra
  Software image filters
  Preventive maintenance and calibration
  Quality control

## Slide 16: Factors that influence Patient Absorbed Dose
Procedural-related factors
  Positioning of image receptor and X ray source relative to the patient
  Beam orientation and movement
  Collimation
  Acquisition and fluoroscopic technique factors on some units
  Fluoroscopy pulse rate
  Acquisition frame rate
  Total fluoroscopy time
  Total acquisition time

## Slide 17
Image Handling
and Display
Image Receptor
X-Ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
Generator and Feedback Schematic

## Slide 18: Factors that influence Patient Absorbed Dose
Patient-related factors
  Patient body weight and habitus

## Slide 19
Factors affecting the penetration of radiation through an object
Notes: <number>
<footer>
<header>
<date/time>
One of the characteristics of x- and gamma radiations that makes them useful for medical imaging is their penetrating ability. When they are directed into an object, some of the photons are absorbed or scattered, whereas others completely penetrate the object. The penetration can be expressed as the fraction of radiation passing through the object. Penetration is the inverse of attenuation. The amount of penetration depends on the energy of the individual photons and the atomic number, density, and thickness of the object,
The probability of photons interacting, especially with the photoelectric effect, is related to their energy. Increasing photon energy generally decreases the probability of interactions (attenuation) and, therefore, increases penetration. As a rule, high-energy photons are more penetrating than low-energy photons,

## Slide 20
Thicker tissue masses absorb more radiation, thus much more radiation must be used to penetrate a large patient. Risk to skin is greater in larger patients!
[ESD = Entrance Skin Dose]
15 cm
20 cm
25 cm
30 cm
ESD = 1 unit
ESD = 2-3 units
ESD = 4-6 units
ESD = 8-12 units
Example: 2 Gy
Example: 4-6 Gy
Example: 8-12 Gy
Example: 16-24 Gy
  Patient Weight and Habitus

## Slide 22
Thicker tissue masses absorb more radiation, thus much more radiation must be used when steep beam angles are employed. Risk to skin is greater with steeper beam angles!
  Tissue Mass and Beam Orientation
Quiz: what happens when cranial tilt is employed?

## Slide 23
100 cm
80 cm
Dose rate:
20 – 40 mGyt/min
Thick Oblique vs. Thin PA geometry
100 cm
50 cm
Dose rate:
~250 mGyt/min
40 cm

## Slide 24
Variation in exposure rate with projection
anthropomorphic phantom (average-sized) measurements
Cusma JACC 1999

## Slide 25: Unnecessary body parts in direct radiation field
Unnecessary body mass in beam
Reproduced from Wagner – Archer, Minimizing Risks from Fluoroscopic X Rays, 3rd ed, Houston, TX, R. M. Partnership, 2000
Reproduced with permission from Vañó et al, Brit J Radiol 1998, 71, 510-516

## Slide 26
Wagner and Archer. Minimizing Risks from Fluoroscopic X Rays.
At 3 wks
At 6.5 mos
Surgical flap
Following ablation procedure with arm in beam near port and separator cone removed. About 20 minutes of fluoroscopy.

## Slide 27
Big problem!
Lessons:
Output increases because arm is in beam.
Arm receives intense rate because it is so close to source.
Arm positioning – important and not easy!

## Slide 28: Examples of Injury when Female Breast is Exposed to Direct Beam
Reproduced with permission from MacKenzie, Brit J Ca 1965; 19, 1 - 8
Reproduced with permission from Vañó, Br J Radiol 1998; 71, 510 - 516.
Reproduced with permission from Granel et al, Ann Dermatol Venereol 1998; 125; 405 - 407

## Slide 29: Lesson
Keep unnecessary body parts, especially arms and female breasts, out of the direct beam.

## Slide 30: Factors that influence Patient Absorbed Dose
Equipment-related factors
  Movement capabilities of C-arm, X ray source, image receptor
  Field-of-view size
  Collimator position
  Beam filtration
  Fluoroscopy pulse rate and acquisition frame rate
  Fluoroscopy and acquisition input dose rates
  Automatic dose-rate control including beam energy management options
  X ray photon energy spectra
  Software image filters
  Preventive maintenance and calibration
  Quality control

## Slide 31
Image Handling
and Display
Image Receptor
X ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
Image receptor degrades
with time

## Slide 32
Image Handling
and Display
Image Receptor
X ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
Feedback circuitry from the image receptor communicates with the X-ray generator  modulates X-ray output to achieve appropriate subject penetration by the X-ray beam and image brightness.

## Slide 33: Field of View of Image Receptors

## Slide 34: Equipment Selection
Angiography equipment of different FOV (Field of View)
dedicated cardiac image intensifier (smaller FOV, 23-25cm) is more dose efficient than a combined cardiac / peripheral (larger FOV) image intensifier
larger image intensifier also limits beam angulation (difficult to obtain deep sagittal angulation )
9-inch
(23 cm)
12-inch

## Slide 35
Dose rate dependence on image receptor active field-of-view or magnification mode.
In general, for image intensifier, the dose rate often INCREASES as the degree of electronic magnification of the image increases.

## Slide 36
IMAGE INTENSIFIER
Active Field-of-View (FOV)
RELATIVE PATIENT ENTRANCE DOSE RATE
FOR SOME UNITS
12" (32 cm) 100
9" (22 cm) 200
6" (16 cm) 300
4.5" (11 cm) 400

## Slide 37
How input dose rate changes with different FOVs depends on machine design and must be verified by a medical physicist to properly incorporate use into procedures.
A typical rule is to use the least magnification necessary for the procedure, but this does not apply to all machines.

## Slide 38: Beam Energy, Filter & kVp

## Slide 39
Image Contrast
No object image
is generated
Object image
is generated
Object silhouette
with no internal
details
Notes: Image contrast. Differential absorption of the X-ray beam by different parts of an object renders its internal structure visible in an X-ray transmission image. X-ray photons are represented as arrows. The left-hand example shows an object that allows all X-ray photons to pass through it unattenuated. In this circumstance, no image of the object is generated. The center example shows an object, different parts of which absorb different fractions of the incident photons modulating the beam intensity and generating an X-ray image of the object. The right-hand example is completely opaque to X-ray photons and absorbs all of them. The X-ray image of this object would be a silhouette with no definition of the object’s internal structure.

## Slide 40
Effect of X ray Beam Penetration on Contrast, Body Penetration, and Dose
Notes: <number>
<footer>
<header>
<date/time>
The first step in obtaining optimum contrast in an x-ray image is to adjust the x-ray beam spectrum for the specific anatomy and clinical purpose. The penetration and the resulting contrast of a specific object or structure in the body generally depends on the photon energy spectrum.
Contrast is not the only thing that must be considered in selecting the spectrum for a specific procedure. The spectrum also affects the penetration through the body section being imaged. This has a significant effect on the radiation dose to the patient. Also, as the penetration through a body section is reduced, the amount of radiation required from the x-ray tube is increased with a resulting increase in x-ray tube heating.

## Slide 41
Beam energy:
In general, every x-ray system produces a range of energies. Higher energy X ray photons  higher tissue penetration.
Low energy X rays: high image contrast but high skin dose
Middle energy X rays: high contrast for iodine and moderate skin dose
High energy X rays: poor contrast and low skin dose

## Slide 42
Beam energy:
The goal is to shape the beam energy spectrum for the best contrast at the lowest dose. An improved spectrum with 0.2 mm Copper filtration is depicted by the dashes:
Middle energy X rays are retained for best compromise on image quality and dose
Low-contrast high energy X rays are reduced by lower kVp
Filtration reduces poorly penetrating low energy X rays

## Slide 43
Beam energy:
kVp controls the high-energy end of the spectrum and is usually adjusted by the system according to patient size and imaging needs:
  kVp (kiloVolt-peak)
Reproduced with permission from Wagner LK, Houston, TX 2004.

## Slide 44
Comparison of Photon Energy Spectra
Produced at Different kVp Values
(from The Physical Principles of Medical Imagings, 2Ed, Perry Sprawls)

## Slide 45
Beam energy:
Filtration controls the low-energy end of the spectrum. Some systems have a fixed filter that is not adjustable; others have a set of filters that are used under differing imaging schemes.
  Filtration
Reproduced with permission from Wagner LK, Houston, TX 2004.

## Slide 46
Filter
Notes: <number>
<footer>
<header>
<date/time>
The addition of filtration significantly alters the shape of the x-ray spectrum, as shown below. Since filtration selectively absorbs the lower energy photons, it produces a shift in the effective energy of an x-ray beam. The figure below compares an unfiltered spectrum to spectra that passed through 1-mm and 3-mm filters. It is apparent that increasing the filtration from 1 mm to 3 mm of aluminum produces a noticeable decrease in the number of x-ray photons. It should be observed, however, that most of this decrease is in photons with energies less than approximately 40 keV. These are the photons with a low probability of penetrating a typical patient and contributing to image formation. They do, however, contribute to patient exposure.
Adding filtration increases the penetration (HVL) of an x-ray beam by removing the low-energy photons.

## Slide 47
Filters:
(1) Advantages -- they can reduce skin dose by a factor of > 2.
(2) Disadvantages -- they reduce overall beam intensity and require heavy-duty X ray tubes to produce sufficient radiation outputs that can adequately penetrate the filters.
Beam energy spectrum before and after adding 0.2 mm of Cu filtration. Note the reduced intensity and change in energies. To regain intensity tube current must increase, requiring special x-ray tube.
  Filtration – possible disadvantage

## Slide 48
If filters reduce intensity excessively, image quality is compromised, usually in the form of increased motion blurring or excessive quantum mottle (image noise).
Lesson: To use filters optimally, systems must be designed to produce appropriate beam intensities with variable filter options that depend on patient size and the imaging task.
  Filtration –potential disadvantage

## Slide 49: Dose vs. Noise
2 µR per frame
15 µR per frame
24 µR per frame
Notes: <number>
<footer>
<header>
<date/time>

## Slide 50: Efficient Dose and Image Quality Management
0.25
2
6
10
14
Detector Dose [GY/s]
0.2 mm Cu-eq MRC
0.5 mm Cu-eq MRC
No Cu-eq
Conventional
0.5
0.75
1
-50%
Same Image quality
30cm water
Patient Dose [cGY/min]
Achieving significant patient pose savings and yet keeping image quality at the same level
Notes: This slide is a technical explanation of the previous slide.This slide compares just like the previous graph slide the patient dose on the vertical axis to the detector dose on the horizontal axis.
Vertical: The higher the patient dose, the more harmful this is to the patient and to the doctor.
Horizontal: The higher the detector dose, the better the Image Quality.
The lines provide a technical visualization for the patient dose for a patient with the appropriate detector dose for a conventional x-ray tube with no Copper Filtering (orange line), a Philips MRC x-ray tube with 0.2 mm Copper Filtering (black line), and finally 0.5 mm Copper filtering (green line).
At a certain Image Quality, there is a certain dose at the detector
Follow the first blue dotted line: At this Image Quality and a conventional tube, there is a certain patient dose.
However, follow the second blue dotted line. At the same Detector Dose (is same Image Quality) in combination with a Philips MRC x-ray-tube and 0.5 mm copper filtration, there is a significantly lower patient dose.
This patient dose is 50% lower for a Philips MRC x-ray tube than a conventional x-ray tube at the same Image Quality.
In summary: Significant Patient Dose reduction is done by pre-filtration. Achieving 50% patient dose reduction by keeping the Image Quality at the same level.

## Slide 51: Revision Qs: “True” or “False”?
The higher the kVp, the higher the energy of the X ray photons, and the more contrast is the X ray image.
When acquiring angiography with image intensifier, it is always better to use as magnified a field-of-view (FOV) as possible, because more details can be visualized.

## Slide 52: Revision Qs: “True” or “False”?
To avoid physical injury to patient, and to facilitate C-arm movement, it is advisable to keep the image receptor as far away from patient as possible.
Patient has complex triple-vessel disease for angioplasty/stenting. Doing the angioplasty for all narrowings in one procedure will increase the risk of deterministic radiation injuries.

## Slide 53: Revision Qs: “True” or “False”?
Scattered radiation has no impact on the X ray image quality.
Angiography table should be kept as near to the X ray source as possible.
Keeping the same pulse intensity, reducing fluoroscopy pulse rate from 30 to 15 pulses/sec will reduce radiation dose to patient by 50%.