## Slide 1: Patient Dose Management
L 5b

## Slide 2: Factors that influence Patient Absorbed Dose
Procedural-related factors
  Positioning of image receptor and X ray source relative to the patient
  Beam orientation and movement
  Collimation
  Acquisition and fluoroscopic technique factors on some units
  Fluoroscopy pulse rate
  Acquisition frame rate
  Total fluoroscopy/acquisition time

## Slide 3: Positioning of image receptor and X ray source relative to the patient

## Slide 4
Beam entering patient typically ~100x more intense than exit beam in average size patient
Only a small percentage (typically ~1%) penetrate through to create the image.

## Slide 5
X ray intensity decreases rapidly with distance from source; conversely, intensity increases rapidly with closer distances to source.
1 unit of intensity
4 units of intensity
16 units of intensity
64 units of intensity
  Inverse Square Law
70 cm
35 cm
17.5 cm
8.8 cm

## Slide 6
Image Handling
and Display
Image Receptor
X ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
Feedback circuitry from the image receptor communicates with the X ray generator  modulates X ray output to achieve appropriate subject penetration by the X ray beam and image brightness.
Automatic
Brightness
Control (ABC)

## Slide 7
All other conditions unchanged, moving image receptor toward patient lowers radiation output rate and lowers skin dose rate.
4 units of intensity
Image
Receptor
2 units of intensity
Image
Receptor
Image
Receptor
  Inverse Square Law (1)

## Slide 8
4 units of intensity
Image
Receptor
2 units of intensity
Image
Receptor
Image
Receptor
  Inverse Square Law (1)
Lesson: Keep the image intensifier as close to the patient as is practicable for the procedure.

## Slide 9
Distance between patient and detector

## Slide 10
All other conditions unchanged, moving patient toward or away from the X ray tube can significantly affect dose rate to the skin
Lesson: Keep the X ray tube at the practicable maximum distance from the patient.
  Inverse Square Law (2)
2 units of intensity
4 units of intensity
16 units of intensity
64 units of intensity

## Slide 11
Distance between patient and X ray source

## Slide 12
Tall vs. Short Operators - Impact on Patient Dose?

## Slide 13: Beam Orientation

## Slide 14
Positioning anatomy of interest at the isocenter permits easy reorientation of the C-arm.
This usually shortens the distance between the X ray tube and the patient, increasing the patient’s entrance port skin dose.
  ISOCENTER

## Slide 15
When isocenter technique is employed, move the image intensifier as close to the patient as practicable to limit dose rate to the entrance skin surface.
  ISOCENTER

## Slide 16
  Physical factors and challenges to radiation management
Lesson: Reorienting the beam distributes dose to other skin sites and reduces risk to single skin site.
  Beam Orientation
This is especially important in coronary angioplasty
for chronic total occlusion.

## Slide 17
Lesson: Reorienting the beam in small increments may leave area of overlap in beam projections, resulting in large accumulations for overlap area (red area). Good collimation can reduce this effect.
  Overlap Areas in Beam Re-orientation
Reproduced with permission from Wagner LK, Houston, TX 2004.

## Slide 18
  Physical factors and challenges to radiation management
Conclusion: Orientation of beam is usually determined and fixed by clinical need. When practical, reorientation of the beam to a new skin site can lessen risk to skin. Overlapping areas remaining after reorientation are still at high risk. Good collimation reduces the overlap area.
  Beam Orientation

## Slide 19: Imaging modes – Fluoroscopy, (Cine) Acquisition, Digital Subtraction Angiography

## Slide 20: Fluoroscopy vs Cine Acquisition
Influence of operation modes: from low fluoroscopy to cine, radiation / scatter dose rate could increase in a factor of 10-15

## Slide 22
Can you tell ……….
Which image is FLUOROSCOPY ? Which one is ACQUISITION?

## Slide 23
Radiation
Dose
Image
Quality
Better image quality with higher radiation dose reaching
the image receptor.
Tradeoff: higher patient dose!!

## Slide 24
ALARA
As Low As Reasonably Achievable
No known safe limit of magnitude of radiation exposure.
Patients
Professional
staff
Physicians

## Slide 25
Siemens Axiom Artis
Cine normal mode
20 cm PMMA
177 Gy/fr (entrance PMMA)
Siemens Axiom Artis, Fluoro low dose
20 cm PMMA
13 Gy/fr (entrance PMMA)

## Slide 26
Set the default fluoroscopy mode to LOW
Lowest input dose needed to
generate a USABLE image

## Slide 27
Influence of operation modes: from low fluoroscopy to cine, radiation / scatter dose rate could increase in a factor of 10-15
Duration of Fluoroscopy/Cine Acquisition
Important to keep in mind DURATION of fluoroscopy
 fluoroscopy x 10-15 sec ~ cine x 1 sec

## Slide 28: Digital Image Subtraction (DSA)
Obtained by subtracting one image from another  electronically removes information that is identical in 2 images
Subtraction process accentuates image noise  counter this effect by acquiring each of the original images at a substantially (up to 20x) higher dose per frame.
Generally, studies that use DSA employ larger aggregate doses than do studies that employ unsubtracted cinefluorography.

## Slide 29: Pulsed Fluoroscopy

## Slide 30
  Design of fluoroscopic equipment for proper radiation control
Understanding Variable Pulsed Fluoroscopy
Background: dynamic imaging captures many still images every second and displays these still-frame images in real-time succession to produce the perception of motion. How these images are captured and displayed can be manipulated to manage both dose rate to the patient and dynamic image quality. Standard imaging captures and displays 25 - 30 images per second.
  Pulsed Fluoroscopy

## Slide 31
[ video clip]
Each angiographic ‘run’ consists of multiple still images taken in quick succession.

## Slide 32
30 images in 1 second
Continuous fluoroscopy
X rays
In conventional continuous-beam fluoroscopy there is an inherent blurred appearance of motion because the exposure time of each image lasts the full 1/30th of a second at 30 frames per second.
Continuous stream of X rays produces blurred images in each frame
Images

## Slide 33
Each X ray pulse shown above has greater intensity than continuous mode, but lasts for only 1/100th of a second; no X rays are emitted between pulses; dose to patient is same as that with continuous fluoroscopy
Pulsed fluoroscopy, no dose reduction
Images
Pulsed fluoroscopy produces sharp appearance of motion because each of 30 images per second is captured in a pulse or snapshot (e.g., 1/100th of a second).
X rays
30 images in 1 second

## Slide 34
Fluoroscopic pulsing X rays are produced during a small portion of the video frame time. The narrower the pulse width, the sharper the image. ( “Faster shutter speed” in camera )
Pulsed Fluoroscopy

## Slide 35
  Physical factors and challenges to radiation management
Pulsed imaging controls:
Displaying 25–30 picture frames per second is usually adequate for the transition from frame to frame to appear smooth.
This is important for entertainment purposes, but not necessarily required for medical procedures.
Manipulation of frame rate can be used to produce enormous savings in dose accumulation.
  Pulsed Fluoroscopy

## Slide 36
Pulsed fluoroscopy, dose reduction at 15 pulses per second
Sharp appearance of motion captured at 15 images per second in pulsed mode. Dose per pulse is same, but only half as many pulses are used, thus dose is reduced by 50%. The tradeoff is a slightly choppy appearance in motion since only half as many images are shown per second
Images
X rays
15 images in 1 second

## Slide 37
Pulsed fluoroscopy at 7.5 images per second with only 25% the dose
Pulsed fluoroscopy, dose reduction at 7.5 pulses per second
Images
X rays
Average 7.5 images in 1 second

## Slide 38
Pulsed fluoroscopy, dose enhancement at 15 pulses per second
Dose per pulse is enhanced because pulse intensity and duration is increased. Overall dose is enhanced.
Images
X rays
15 images in 1 second
Reproduced with permission from Wagner LK, Houston, TX 2004.
Images
X rays
15 images in 1 second

## Slide 39
  Design of fluoroscopic equipment for proper radiation control
Lesson: Variable pulsed fluoroscopy is an important tool to manage radiation dose to patients but the actual effect on dose can be to enhance, decrease or maintain dose levels. The actual effect must be estimated by a qualified physicist so that variable pulsed fluoroscopy can be properly employed.
  Variable Pulsed Fluoroscopy

## Slide 40: Collimation

## Slide 41: Collimation

## Slide 42: A word about collimation
What does collimation do?
Collimation confines the X ray beam to an area of the user’s choice.

## Slide 43: Collimation
Why is narrowing the field-of-view beneficial?
Reduces stochastic risk to patient by reducing volume of tissue at risk
Reduces scatter radiation at image receptor to improve image contrast
Reduces scatter radiation to in-room personnel
Reduces potential overlap of fields when beam is reoriented

## Slide 44: Scattered Radiation
Patient
Operator
Lab
Personnel
X-Ray
Scattered
radiation
Two undesirable effects:
  (1) predominant source of radiation exposure to the laboratory personnel;

## Slide 45: Scattered Radiation
Two undesirable effects:
  (2) scattered radiation that continues in the forward direction and reaches the image receptor decreases the quality (contrast) of the image
Reduction of Image Contrast
by Scattered Radiation

## Slide 46
Collimation: Contrast Improvement by Reducing X ray Beam Size
Notes: <number>
<footer>
<header>
<date/time>
The amount of scattered radiation is generally proportional to the total mass of tissue contained within the primary x-ray beam. This is, in turn, determined by the thickness of the patient and the area or field size being exposed. Increasing the field size increases the total amount of scattered radiation and the value of the scatter contrast-reduction factors. Therefore, one method of reducing scattered radiation and increasing contrast is to reduce the field size with x-ray beam collimators, cones, or other beam-limiting devices, as illustrated below. This method is limited by the necessity to cover a specific anatomical region. However, in most situations, contrast can be improved by reducing the field size to the smallest practical value.

## Slide 47
Lesson: Reorienting the beam in small increments may leave area of overlap in beam projections, resulting in large accumulations for overlap area (red area). Good collimation can reduce this effect.
  Beam Orientation, Overlap and Collimation

## Slide 48: Collimation
What collimation does NOT do –
It does NOT reduce dose to the exposed portion of patient’s skin
In fact, dose at the skin entrance site increases, sometimes by a factor of 50% or so, depending on conditions.

## Slide 49: Factors that influence Patient Absorbed Dose
Equipment-related factors
  Movement capabilities of C-arm, X ray source, image receptor
  Field-of-view size
  Collimator position
  Beam filtration
  Fluoroscopy pulse rate and acquisition frame rate
  Fluoroscopy and acquisition input dose rates
  Automatic dose-rate control including beam energy management options
  X ray photon energy spectra
  Software image filters
  Preventive maintenance and calibration
  Quality control

## Slide 50
Image Handling
and Display
Image Receptor
X ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
Image receptor degrades
with time

## Slide 51
Image Handling
and Display
Image Receptor
X ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
Feedback circuitry from the image receptor communicates with the X ray generator  modulates X ray output to achieve appropriate subject penetration by the X ray beam and image brightness.

## Slide 52: Field of View of Image Receptors

## Slide 53: Equipment Selection
Angiography equipment of different FOV (Field of View)
dedicated cardiac image intensifier (smaller FOV, 23-25cm) is more dose efficient than a combined cardiac / peripheral (larger FOV) image intensifier
larger image intensifier also limits beam angulation (difficult to obtain deep sagittal angulation )
9-inch
(23 cm)
12-inch

## Slide 54
Dose rate dependence on image receptor active field-of-view or magnification mode.
In general, for image intensifier, the dose rate often INCREASES as the degree of electronic magnification of the image increases.

## Slide 55
IMAGE INTENSIFIER
Active Field-of-View (FOV)
RELATIVE PATIENT ENTRANCE DOSE RATE
FOR SOME UNITS
12" (32 cm) 100
9" (22 cm) 200
6" (16 cm) 300
4.5" (11 cm) 400

## Slide 56
How input dose rate changes with different FOVs depends on machine design and must be verified by a medical physicist to properly incorporate use into procedures.
A typical rule is to use the least magnification necessary for the procedure, but this does not apply to all machines.

## Slide 57: Beam Energy, Filter & kVp

## Slide 58
Image Contrast
No object image
is generated
Object image
is generated
Object silhouette
with no internal
details
Notes: Image contrast. Differential absorption of the X-ray beam by different parts of an object renders its internal structure visible in an X-ray transmission image. X-ray photons are represented as arrows. The left-hand example shows an object that allows all X-ray photons to pass through it unattenuated. In this circumstance, no image of the object is generated. The center example shows an object, different parts of which absorb different fractions of the incident photons modulating the beam intensity and generating an X-ray image of the object. The right-hand example is completely opaque to X-ray photons and absorbs all of them. The X-ray image of this object would be a silhouette with no definition of the object’s internal structure.

## Slide 59
Effect of X ray Beam Penetration on Contrast, Body Penetration, and Dose
Notes: <number>
<footer>
<header>
<date/time>
The first step in obtaining optimum contrast in an x-ray image is to adjust the x-ray beam spectrum for the specific anatomy and clinical purpose. The penetration and the resulting contrast of a specific object or structure in the body generally depends on the photon energy spectrum.
Contrast is not the only thing that must be considered in selecting the spectrum for a specific procedure. The spectrum also affects the penetration through the body section being imaged. This has a significant effect on the radiation dose to the patient. Also, as the penetration through a body section is reduced, the amount of radiation required from the x-ray tube is increased with a resulting increase in x-ray tube heating.

## Slide 60
Beam energy:
In general, every X ray system produces a range of energies. Higher energy X ray photons  higher tissue penetration.
Low energy X rays: high image contrast but high skin dose
Middle energy X rays: high contrast for iodine and moderate skin dose
High energy X rays: poor contrast and low skin dose

## Slide 61
Beam energy:
The goal is to shape the beam energy spectrum for the best contrast at the lowest dose. An improved spectrum with 0.2 mm copper filtration is depicted by the dashes:
Middle energy X rays are retained for best compromise on image quality and dose
Low-contrast high energy X rays are reduced by lower kVp
Filtration reduces poorly penetrating low energy X rays

## Slide 62
Beam energy:
kVp controls the high-energy end of the spectrum and is usually adjusted by the system according to patient size and imaging needs:
  kVp (kiloVolt-peak)
Reproduced with permission from Wagner LK, Houston, TX 2004.

## Slide 63
Comparison of Photon Energy Spectra
Produced at Different kVp Values
(from The Physical Principles of Medical Imagings, 2Ed, Perry Sprawls)

## Slide 64
Beam energy:
Filtration controls the low-energy end of the spectrum. Some systems have a fixed filter that is not adjustable; others have a set of filters that are used under differing imaging schemes.
  Filtration
Reproduced with permission from Wagner LK, Houston, TX 2004.

## Slide 65
Filter
Notes: <number>
<footer>
<header>
<date/time>
The addition of filtration significantly alters the shape of the x-ray spectrum, as shown below. Since filtration selectively absorbs the lower energy photons, it produces a shift in the effective energy of an x-ray beam. The figure below compares an unfiltered spectrum to spectra that passed through 1-mm and 3-mm filters. It is apparent that increasing the filtration from 1 mm to 3 mm of aluminum produces a noticeable decrease in the number of x-ray photons. It should be observed, however, that most of this decrease is in photons with energies less than approximately 40 keV. These are the photons with a low probability of penetrating a typical patient and contributing to image formation. They do, however, contribute to patient exposure.
Adding filtration increases the penetration (HVL) of an x-ray beam by removing the low-energy photons.

## Slide 66
Filters:
(1) Advantages -- they can reduce skin dose by a factor of > 2.
(2) Disadvantages -- they reduce overall beam intensity and require heavy-duty X ray tubes to produce sufficient radiation outputs that can adequately penetrate the filters.
Beam energy spectrum before and after adding 0.2 mm of Cu filtration. Note the reduced intensity and change in energies. To regain intensity tube current must increase, requiring special X ray tube.
  Filtration – possible disadvantage

## Slide 67
If filters reduce intensity excessively, image quality is compromised, usually in the form of increased motion blurring or excessive quantum mottle (image noise).
Lesson: To use filters optimally, systems must be designed to produce appropriate beam intensities with variable filter options that depend on patient size and the imaging task.
  Filtration –potential disadvantage

## Slide 68: Dose vs. Noise
2 µR per frame
15 µR per frame
24 µR per frame
Notes: <number>
<footer>
<header>
<date/time>

## Slide 69: Efficient Dose and Image Quality Management
0.25
2
6
10
14
Detector Dose [GY/s]
0.2 mm Cu-eq MRC
0.5 mm Cu-eq MRC
No Cu-eq
Conventional
0.5
0.75
1
-50%
Same image quality
30cm water
Patient Dose [cGY/min]
Achieving significant patient pose savings and yet keeping image quality at the same level
Notes: This slide is a technical explanation of the previous slide.This slide compares just like the previous graph slide the patient dose on the vertical axis to the detector dose on the horizontal axis.
Vertical: The higher the patient dose, the more harmful this is to the patient and to the doctor.
Horizontal: The higher the detector dose, the better the Image Quality.
The lines provide a technical visualization for the patient dose for a patient with the appropriate detector dose for a conventional x-ray tube with no Copper Filtering (orange line), a Philips MRC x-ray tube with 0.2 mm Copper Filtering (black line), and finally 0.5 mm Copper filtering (green line).
At a certain Image Quality, there is a certain dose at the detector
Follow the first blue dotted line: At this Image Quality and a conventional tube, there is a certain patient dose.
However, follow the second blue dotted line. At the same Detector Dose (is same Image Quality) in combination with a Philips MRC x-ray-tube and 0.5 mm copper filtration, there is a significantly lower patient dose.
This patient dose is 50% lower for a Philips MRC x-ray tube than a conventional x-ray tube at the same Image Quality.
In summary: Significant Patient Dose reduction is done by pre-filtration. Achieving 50% patient dose reduction by keeping the Image Quality at the same level.

## Slide 70: Multiple Procedures

## Slide 71: Procedure Planning
Diagnostic coronary angiography  PTCA
    Same day?
    Different day?
Multivessel PTCA
    Treat all lesions during same procedure?
    Staged PTCA?
Restenosis, Repeat Procedures

## Slide 72: “Dose Fractionation” in Interventional Cardiology
Reduce deterministic risk
  think of it as similar to risk of contrast-related nephropathy
No significant impact on stochastic risk ( cumulative effective dose)

## Slide 73
Dose
Effect
Deterministic effects
Cataract
Infertility
Erythema
Epilation
Cancer
Genetic
Prob  dose
Stochastic

## Slide 74
Patient
Operator
Lab
Personnel
X ray
Scatter
radiation
Measures taken to reduce radiation
exposure to patient will also benefit
the operator/cath lab staff

## Slide 75: Revision Qs: “True” or “False”?
The higher the kVp, the higher the energy of the X ray photons, and the more contrast is the X ray image.
When acquiring angiography with image intensifier, it is always better to use as magnified a field-of-view (FOV) as possible, because more details can be visualized.

## Slide 76: Revision Qs: “True” or “False”?
To avoid physical injury to patient, and to facilitate C-arm movement, it is advisable to keep the image receptor as far away from patient as possible.
Patient has complex triple-vessel disease for angioplasty/stenting. Doing the angioplasty for all narrowings in one procedure will increase the risk of deterministic radiation injuries.

## Slide 77: Revision Qs: “True” or “False”?
Scattered radiation has no impact on the X ray image quality.
Angiography table should be kept as near to the X ray source as possible.
Keeping the same pulse intensity, reducing fluoroscopy pulse rate from 30 to 15 pulses/sec will reduce radiation dose to patient by 50%.