## Slide 1: Optimization of Radiation Protection in Cardiology
L 9

## Slide 2: Educational Objectives
Optimization in interventional cardiology
Ways to improve the radiation protection aspect of procedure (balancing diagnostic information versus patient doses)
A single-centre’s (Udine, Italy) experience with optimization and flat panel detector

## Slide 3: Is this statement “True” or “False”?
Switching from old angiography machine with image intensifier to a new machine with flat-panel detector will definitely reduce patient radiation dose.

## Slide 4
In room dosimetric indications
Have you ever paid
attention to this?
Notes: <number>
<footer>
<header>
<date/time>

## Slide 5: Do you know how to interpret these data?

## Slide 6: Reality check…… Dose Information and Dose Reports
Understanding the dose information available in the cath. lab.?
Do you follow (and archive) patient dose reports?
Understanding the dose values in your occupational dose report?
Can you make out if the dose values as depicted in the monitor in cath. lab. are “normal” or “too high”?

## Slide 7: Reality check…… Knowing your X ray machine
What is the “cost” in radiation dose you are “paying” for the image quality you want?
Have you evaluated whether fewer images with lower image quality are acceptable?
Do you know the dose rate values and dose/image for the different operation modes?
Do you know the difference in dose for the different field of view (FOV) formats (magnification)?
Do you know how to use the newer dose-reducing features?

## Slide 8
Siemens Axiom Artis
Cine normal mode
20 cm PMMA
177 Gy/fr (entrance PMMA)
Siemens Axiom Artis, Fluoro low dose
20 cm PMMA
13 Gy/fr (entrance PMMA)

## Slide 9: Optimization means ...
To avoid acquiring more images than necessary:
  Take care of the fluoroscopy time.
  Take care of the number of series.
  Take care of the number of frames per series.
To avoid acquiring images with more quality (and more dose) than necessary:
  It could be possible to accept sometimes some noisy images in fluoroscopy and also in cine acquisitions.

## Slide 10: Optimization of Radiation Protection
Minimization of dose to patient and staff should not be the goal
Must optimize dose to patient and minimize dose to staff
Optimized patient dose rate should have sufficient dose rate to provide adequate image quality
If image quality is inadequate, then any radiation dose results in needless radiation dose!
Notes: In medical imaging the goal is to provide the image quality necessary for the physician to make a diagnosis. Consequently, a certain level of radiation dose is required and it is not possible to minimize patient radiation exposure.

## Slide 11: One of the center's experience with new angiographic machine having flat panel detector

## Slide 12
Motorized Iris
Video Camera
Image Intensifier
DETECTOR
Photons
Cesium Iodide (CsI)
Light
Amorphous Silicon Panel
(Photodiode/Transistor Array)
Digital Data
Electrons
Read Out Electronics
Photons
Cesium Iodide (CsI)
Light
Photo-cathode
Video Signal
Electrons
Output screen
Light
CCD or PUT
Electrons
Readout Electronics
1
3,000
400
400,000
2,400
Particles #
Image
Intensifier
Flat-panel

## Slide 13: Anticipated per-frame dose reduction with Digital Flat Panel technology is 30%

## Slide 14: Coronary Angiography and Angioplasty Udine, years 1990-2002
Philips OM 200 (1983)
Philips Integris 3000 (1995)
performed by 3 interventionalists except in 1998

## Slide 15: GE Innova 2000 (angiographic machine with digital flat panel technology) at Udine Center
Activity started
  04/12/2002
Jan - Oct 2003
  1421 procedures (79% of total)
    1019 diagnostic coronary angiography
    402 percutaneous coronary angioplasty

## Slide 16: Comparison of Philips H 3000 and Innova 2000 in PCI -- Characteristics of patients
diseased vessels (%)
H 3000: 588 pts, 90% of tot. treated in the year 2002
Innova: 274 pts, 67% of tot. treated between Jan-Oct 2003

## Slide 17
Comparison of Philips H 3000 and Innova 2000 in PCI -- Characteristics of procedures & lesions (1)
%

## Slide 18
Comparison of Philips H 3000 and Innova 2000 in PCI -- Characteristics of procedures & lesions (2)
%

## Slide 19
(m’)
Comparison of Philips H 3000 and Innova 2000 in PCI -- Performance & complexity index
1,37
1
1,47
0,93
0
0,2
0,4
0,6
0,8
1
1,2
1,4
1,6
1,8
2
Comp. Index
GISE Index
r (with fluoro time)
0.30 0.29 0.34 0.26
11,6
11,5
40
48
59
84
24,1
21,6
H 3000
Innova

## Slide 20: Comparison of Philips H 3000 and Innova 2000 in diagnostic procedures -- Characteristics of patients & procedures
(%)
H 3000: 1401 pts, 92% of tot. studied in the year 2002
Innova: 702 pt, 69% of tot. studied between Jan-Oct 2003

## Slide 21
m’
Gy * cm2
measured
Comparison of Philips H 3000 and Innova 2000 in diagnostic procedures
-- performance indexes and exposure parameters
calculated
4.2
24
54
15.6
20.39
10.67
31.06
4.4
28
54
15.8
27.05
18.83
45.88
35.32
0
10
20
30
40
50
60
Fluoro T
proced. T
room occ.
contrast (dl)
cine DAP
fluoro DAP
tot. DAP
1,3
H 3000
Innova

## Slide 22: Entrance surface dose: H3000 and Innova 2000
Entrance surface dose rates in Fluoro LOW for Innova is 30% less
Notes: <number>
<footer>
<header>
<date/time>

## Slide 23
?
Why the anticipated 30% per-frame dose reduction of digital flat panel technology does not translate into an effective dose reduction to patients

## Slide 24: Differences in operating conditions of the two systems
H 3000
field of view (cm)
  23/18/14
cine mode
  12,5/25 fps
fluoro mode
  low/medium/high
filter
  automatic
Innova
field of view (cm)
  20/17/15/12
cine mode
  15/30 fps
  pref 1/pref 2 (lower dose)
fluoro mode
  low/normal
filter
  manual

## Slide 25
20 cm
23 cm
400 cm2
375 cm2
Detectors area are similar
H3000
Innova200

## Slide 26
17 cm
18 cm
290 cm2
230 cm2
area +26%  DAP +26% !!!!
Using similar nominal field of view (FOV) sizes corresponds to very different area of the two X ray beams
H3000
Innova200

## Slide 27: Other possibilities……
patients may not be the same
procedures may not be the same
operators’ behavior
  filters/collimation
  use of “difficult” projections (fluoro/cine)
  focus-detector mean distances
………

## Slide 28
Collimators use in INNOVA to reduce exposure
FOV 15
dose reduction 25%
[still]

## Slide 29
FOV 20
Collimators use in INNOVA to reduce exposure
[still]

## Slide 30
proper filtering
improper filtering causes image deterioration
H 3000
Filtering prevents image saturation in
low absorption areas

## Slide 31: INNOVA
improper filtering does not cause image deterioration

## Slide 32
Skin exposure variation in exposure rate (DAP rate) with projection
anthropomorphic phantom (average-sized) measurements
Cusma JACC 1999
Projection | Fluoroscopy entrance dose rate (mGy/min) | Cine entrance dose rate (mGy/min)
AP | 31 | 388
RAO 30° | 19 | 203
LAO 40° | 20 | 216
LAO 40°, Cran 30° | 80 | 991
LAO 40°, Cran 40° | 99 | 1236
LAO 40°, Caud 20° | 29 | 341

## Slide 33
Distance between patient and detector

## Slide 34
d
2d
Because the same energy is spread over a surface 4 times larger at a doubled distance, the same object will receive only a fourth of the dose when moved away from “d” to “2d”
Source
Doubling the distance from the source divides the dose by a factor of 4
The inverse square law

## Slide 35
The inverse square law

## Slide 36
Collimation

## Slide 37
Anti-scatter grid
Increase DAP and skin dose x 2 times
Improve image quality
To be removed for pediatric patients !!

## Slide 38: Coronary Angiography optimization DIMOND Quality Criteria
Use of the wedge filter on bright peripheral areas
2-3 sequences (except for difficult anatomic details)
12.5-15 frames/s (25-30 only if heart rate exceeds 90-100 bpm or in paediatric patients)
60 images per sequence at average (12.5-15 fr/s) except if collaterals have to be imaged or in case of slow flow
Aspects of an optimized angiographic technique

## Slide 39: Innova 2000. Changes in exposure parameters over time -- diagnostic procedures Jul 2003 - Feb 2004

## Slide 40: Optimization Process

## Slide 41: Optimization requires……….
Knowledge of factors contributing to patient and staff radiation dose
  patient factors
  procedural factors
  equipment (machine) factors
Knowledge of dose reduction capabilities of our X ray system
Periodic update of our clinical and technical working protocols

## Slide 42: Optimization process involves ............
Data collection
  procedures, DAP, fluoro time
Data analysis
  reliability of data
Discussion & processes review
  collimators/filters use, FOV, projections
Implementation of changes
  more precise data collection, collimators/filters use, FOV 17 whenever possible, avoiding LAO projections
Data verification

## Slide 43: Reference levels
3rd level
“Patient risk”
2nd level
“Clinical protocol”
1st level
“Equipment performance”
Dose rate and dose/image
(BSS, CDRH, AAPM)
Level 1 + No. images + fluoroscopy time
Level 2 + DAP + Maximum Skin Dose (MSD)
Reference levels: an instrument to help operators to conduct optimized procedures with reference to patient exposure
Required by international (IAEA) and national regulations
For complex procedures reference levels should include:
more parameters
and, must take into account the protection from stochastic and deterministic risks
(Dimond)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 44
Modern X ray systems display dosimetric indications directly on the console in the control room and inside the catheterization laboratory, allowing cardiologists to know the level of radiological risk during the procedure.
Typically Dose Area Product and Cumulative Dose (*) are displayed.
(*) Cumulative Dose (CD) is the air kerma accumulated for a procedure at a specific point in space relative to the fluoroscopic gantry for a procedure (it does not include tissue backscatter). It can give an indication of the skin dose.

## Slide 45
Example of the data included in the study report (Siemens)

## Slide 46
The proposed reference levels for Coronary Angiography and PTCA were DAP 45 Gy•cm2 and 75 Gy•cm2; fluoroscopy time 7.5 min and 17 min and number of frames 1250 and 1300, respectively.

## Slide 47: Procedure optimization in the cath. lab. patients and staff share a lot……
  correct indications
  fluoro time reduction
  frame rate reduction (25 12,5/sec)
  collimation/filtering
  LAO cranial projection limitation
  distance from X ray source
  lead apron and thyroid protection
  protective glasses and suspended screen
(staff)
(patient)

## Slide 48
Optimization is especially important in more complex PTCA procedures
  chronic total occlusion
  bifurcation lesion
  degenerated saphenous
  vein graft lesion
  lesion in severely
  tortuous vessel
  ostial lesion

## Slide 49
Procedure optimization annual hand dose (cardiologist)
- 27%
+ 2%
- 49%
- 23%
Cardiologia & Fisica Sanitaria - Udine
mSv
1994-1998 - 71%

## Slide 50
Gy*cm2
Procedure optimization DAP measurements at Udine Hospital (all procedures)

## Slide 51
Patient
Operator
Lab
Personnel
X-Ray
Scatter
radiation
Measures taken to reduce radiation
exposure to patient will also benefit
the operator/cath. lab. staff