## Slide 1: Radiation Protection in Paediatric Interventional Cardiology
L 10

## Slide 2: Answer “True” or “False”?
Children are more sensitive to radiation than adults.
Exposure parameters on X ray machines are often not adjusted for paediatric patients.

## Slide 3: Educational Objective
Unique considerations in paediatric patients
Special consideration regarding equipment
How can dose be managed in paediatric patients

## Slide 4: Unique Considerations for Radiation Exposure in Children
Some unique considerations in children:
  Children are considerably more sensitive to radiation than adults
      Risk factors for cancer induction in children is between 2 and 3 times higher than for adults
  Children have longer life expectancy greater potential for manifestation of possible harmful effects of radiation

## Slide 6: Unique Considerations for Radiation Exposure in Children
Compared with a 40-year old, a neonate is several times more likely to produce a cancer over the child's lifetime, when exposed to the same radiation dose
  Radiation doses used to examine young children must generally be smaller than those employed in adults

## Slide 7: Need for Exposure Parameters Adjustment
Currently, exposure parameters are sometimes not adjusted for paediatric populations
e.g. CT examinations in children that are not optimized --
  the same exposure parameters used for a child and an adult will result in comparatively larger doses to the child
  there is no need for these larger doses to children.

## Slide 8: Some interventional procedures in pediatric cardiology
  Balloon dilatation / stenting
    vascular stenoses
      aortic coarctation
    valvular obstructive lesions
      pulmonary stenosis
      mitral stenosis
  Transcatheter closure
    atrial septal defects (ASD)
    ventricular septal defect (VSD)
    patent ductus arteriosus (PDA)
  Electrophysiology
    ablation

## Slide 9: Trends in pediatric interventional cardiology
Future of interventional cardiology in pediatrics Levi DS, Alejos JC, Moore JW. Curr Opin Cardiol. 2003 Mar;18(2):79-90.
A trend toward use of less invasive, non-surgical approaches to the treatment of congenital heart disease
Fetal catheter-based interventions are being developed for the treatment of severe congenital heart disease in utero
Trend toward catheter-mediated treatment is certain to continue, care must be taken to regulate safely the introduction of novel techniques and devices into clinical use

## Slide 10: Radiation Exposure to Children during various interventions (I)
Coil occlusion of the patent ductus arteriosus (PDA) as well as other more complex pediatric interventions has raised concern regarding radiation exposure
No correlation between fluoroscopy time and measured entrance dose
Strongest correlation cumulative dose vs. patient weight and BSA J. Donald Moore, David Shim, John Sweet, Kristopher L. Arheart and Robert H. Beekman III,Catheterization and Cardiovascular Interventions 47:449–454 (1999)
David Shim, Thomas R. Kimball, Erik C. Michelfelder, Lisa Koons, RN and Robert H. Beekman, Catheterization and Cardiovascular Interventions 51:451–454 (2000)
Procedure | No. patients | Fluoroscopy time (min) | Cine time (sec) | Total cumulative skin dose (mGy)
PDA | 8 | 10.1 ±1.8 | 11.5±1.8 | 97±25
PBV | 5 | 19.3 ±2.3 | 10.9±2.3 | 86±32
Amplatzer | 14 | 23.5 ±2.1 | 19.9±3.3 | 102±34
DC | 12 | 13.2 ±1.5 | 18.7±1.5 | 108±21
DC=diagnostic catheterization, PDA=Coil occlusion of the patent ductus arteriosus, PBV=pulmonary balloon valvuloplasty

## Slide 11: Radiation Exposure to Children during various interventions (II)
Cumulative skin dose is well correlated with patient size and not with fluoroscopy time

## Slide 12: Radiation Exposure to Children during various interventions (III)
Comparison of surface entrance doses of radiation A: Present study (Amplatzer atrial septal defect closure) B: Moore et al. [6] (patent ductus coil occlusion) C: Moore et al. [6] (pulmonary valvuloplasty) D: Wu et al. [8] (pulmonary valvuloplasty) E: Park et al. [10] (arhythmia ablation) F: Rosenthal et al. [11] (arhythmia ablation)

## Slide 13: X ray Equipment Consideration

## Slide 14: X ray equipment for pediatric cardiology
The generator should have enough power to allow short exposure times (3 milliseconds).
Fluoroscopic pulsing X rays are produced during a small portion of the video frame time. The narrower the pulse width, the sharper the image. ( “Shutter speed” in camera )

## Slide 15: X ray equipment for pediatric cardiology
The generator should be of high frequency (i.e can produce higher pulsed fluoroscopy) to improve the accuracy and reproducibility of exposures.
  E.g. children have faster heart rate. Coronary angiography in children is often acquired at 25-30 frames/sec, instead of the usual 12.5 – 15 frames/sec for adult patients.

## Slide 16: X ray equipment for pediatric cardiology
Automatic exposure control (AEC) devices should be used with caution in pediatrics
Careful manual selection of exposure factors usually results in lower doses
High kV technique should be used

## Slide 17: X ray equipment for pediatric cardiology
Image Handling
and Display
Image Receptor
X ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
  Image intensifier should have high conversion factor to reduce patient dose
  Image intensifiers should have
  high conversion factors
  for reducing patient dose

## Slide 18: Anti-scatter Grid
The anti-scatter grid in pediatrics gives limited improvement in image quality and increases patient dose given the smaller irradiated volume (and mass) the scattered radiation is less

## Slide 19
Increase DAP and skin dose typically by ≥ 2 times
  Does NOT improve image quality very much in paediatric patients (unlike in adults)
  To be removed for paediatric patients !!
Anti-scatter Grid

## Slide 20: Procedure optimization in the pediatric cath. lab. patients and staff share a lot……
correct indications
fluoro time reduction
frame rate reduction
collimation/filtering
distance from X ray source
/ image receptor
protective organ shielding
e.g gonad, thyroid
lead apron and thyroid protection
protective glasses and suspended screen
(staff)
(patient)

## Slide 21: ICRP-ISR “smart” message for pediatrics

## Slide 22: Summary
Increased radiation risks for pediatric patients
Trend of increasing number of pediatric interventional procedures
Radiation doses can be high
Very few dosimetric studies
Radiological technique must be optimized and tailored to small body sizes