## Slide 1: Why Talk about Radiation Protection in Cardiology?
L 1
Notes: Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Answer True or False
Serious radiation injuries are only occurring in radiotherapy.
If some radiation injury occurs during cardiac procedures it can be immediately detected by the cardiologist.
Radiation injury will only occur if fluoroscopy time is in the rage of 2-3 hours.

## Slide 3: Educational Objectives
Review of severity & frequency of radiation injuries in cardiology
What do these injuries teach us about the cardiologist’s role (Lessons learnt)
Points-of-view about lawsuits for severe injury
Recognizing radiation injury & effects

## Slide 4: Introduction
Fluoroscopically guided therapeutic procedures in cardiology fall in the generic category of “interventional” procedures.
They are complex procedures that require specially designed equipment and involve considerable radiation exposure to patients.
Management of radiation during these procedures is complex, and physicians and personnel must be well trained in radiation management and safety.
Notes: Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 5: Pre-interventional era practice
Fewer patients having catheterizations
Diagnostic only
Laboratories in big centers
Radiologists involved

## Slide 6: Pre-interventional era beliefs
Risk to staff – small, and mainly cataracts and thyroid cancer
Risk to patients - incredibly small risk of cancer

## Slide 7: Interventional era started in 1977
Grüntzig, Zurich
Crude catheters
Long fluoroscopy
Many cine runs
Big increase in radiation to staff and patients

## Slide 8: Interventional era using radiation
Many new devices
Multivessel PCI
Repeat procedures
Brachytherapy
Valvuloplasties
EP ablations
Lead extractions
LV pacing leads
Radiation exposure to the staff and patient has risen considerably

## Slide 9: Radiation still increasing today
Higher expectations
Tougher cases
Use of IVUS with Angio
More labs
More patients
Cardiologists now
lead the field in use
of fluoroscopy

## Slide 10: Lab challenges remain
Old equipment
Lack of digital replay
Limited catheters
Untrained help
Poor shielding
No monitoring

## Slide 11: Background on radiation injury in cardiology

## Slide 12: Background on Radiation Injury in Cardiology
The situation:
At least two million invasive/intervention/EP studies annually worldwide.
A large percentage of practitioners not trained in radiation management
Equipment sometimes poorly suited for extended procedures
Notes: Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 13: Background on Radiation Injury in Cardiology
The situation (continued):
Overall benefit/risk of interventional cardiology high, but risks of high radiation doses exist and effects can be prolonged and debilitating
When radiation doses are high, patients not usually counselled about risks, nor followed for symptoms of injury
Notes: Instructions for the lecturer/trainer

## Slide 14: Background on Radiation Injury in Cardiology
The situation (continued):
Medical professionals do not always recognize radiation injury when it occurs, often misdiagnose the true cause of the injury, and some have denied that radiation could cause it.
Occupational doses can be very high and can be reduced by reducing unnecessary patient dose.
Notes: Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 15: Background on Radiation Injury in Cardiology
The health risks:
Serious radiation skin injuries have occurred in a very small percentage of patients, but they are numerous.
Younger patients face an increased risk of induced cancer (usually <1% incidence).
Notes: Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 16: Radiation Injuries - Cases

## Slide 17: Lessons from injured patients
Case #1:Electrophysiological and ablation procedure
Three attempts in 4 months, each with more than 100 minutes of fluoroscopy.
After first attempt erythema observed by patient, but not recognized as due to procedure.
Erythemas on back healed. Arm lesion required grafting.
Material used here previously copyrighted by Louis K Wagner 2004 or Partners in Radiation Management LTD Company 2004 and used here by permission.

## Slide 18: Lessons from injured patients
Case #1:Electrophysiological and ablation procedure
After second attempt injuries mistakenly thought caused by faulty grounding pads which were also believed to cause failure of ablation attempts.
Third attempt was performed and injury subsequently got worse.
Back lesions healed, arm required grafting.
Erythemas on back healed. Arm lesion required grafting.
Material used here previously copyrighted by Louis K Wagner 2004 or Partners in Radiation Management LTD Company 2004 and used here by permission.

## Slide 19: Lessons from injured patients
If cause of initial erythemas was correctly identified, injury would likely have been avoided because failures of initial ablations would not have been blamed on faulty equipment.
Erythemas on back healed. Arm lesion required grafting.
Material used here previously copyrighted by Louis K Wagner 2004 or Partners in Radiation Management LTD Company 2004 and used here by permission.

## Slide 20: Lessons from injured patients
Case #2:
PTCA and stent placement of RCA. Involved 63 minutes of fluoroscopy and nearly 5000 frames of cine in LAO orientation with cranial tilt.
Lesion required grafting.
Material used here previously copyrighted by Louis K Wagner 2004 or Partners in Radiation Management LTD Company 2004 and used here by permission.

## Slide 21: Lessons from injured patients
Case #2:
Dose buildup due to long fluoroscopy and fluorography with steep angle through thick chested patient not recognized
Lesion required grafting.
Material used here previously copyrighted by Louis K Wagner 2004 or Partners in Radiation Management LTD Company 2004 and used here by permission.

## Slide 22: Lessons from injured patients
Case #2:
Cause of injury initially misidentified as pressure wound due to defibrillator pad.
Lesion required grafting.
Material used here previously copyrighted by Louis K Wagner 2004 or Partners in Radiation Management LTD Company 2004 and used here by permission.

## Slide 23: Lessons from injured patients
Case #3:
PTCA : 51 minutes high-dose fluoroscopy, 74 seconds cine in 141 kg man. Dose estimated retrospectively at 22 Gy.
Lesion required grafting.

## Slide 24: Lessons from injured patients
Case #3:
Cumulative buildup of dose for steeply angled high-dose beam through large patient not recognized
Lesion required grafting.

## Slide 25: Points-of-view about lawsuits for severe injury

## Slide 26: Points-of-view about lawsuits involving injury :
Severe injury involved (surgical intervention necessary);
Diagnosis delayed by misidentification of etiology;
Responsible physician not able to recognize injury, often initially denies possibility;
Pain and suffering are extensive, with major impact on lifestyle – lost work, reduced mobility, financial indebtedness, chronic care;
Patient perplexed over lack of medical knowledge about such events;

## Slide 27: Lessons from lawsuits – Physician’s point-of-view
Provided benefits necessary to save patient’s life;
Used only enough radiation to complete the procedure;
Followed all medical procedures properly;
Did not know how much radiation the machine was delivering;
Never taught that radiation could cause injury;
FDA advisory on the issue is not regulatory;
Was not informed of FDA advisory;
Believed injury likely due to other factor – e.g., electrical or thermal injury from grounding pad.

## Slide 28: Lessons from lawsuits – Hospital’s point-of-view
Physician and personnel met all qualifications for training;
Machine designed for these procedures;
Maintenance of equipment was proper;
FDA advisory not regulatory;
FDA advisory posted or sent to department.

## Slide 29: Lessons from lawsuits – The dilemma:
Physician performed procedure in good faith;
Physician not taught that injury could occur;
Physician not taught about radiation management for the patient as part of training;
FDA Advisory largely ignored by hospitals and many professional societies;
Patient’s angina cured by procedure;
Patient has severe, debilitating and extremely painful injury;
Patient never advised of the possibility of the injury;
Patient has lost faith in the medical care;
Patient’s life and life of those close to patient changed forever.

## Slide 30: Complex fluoroscopically guided therapeutic cardiology
Procedures in cardiology that potentially involve high-risk radiation exposure include:
Complex and/or mutli-vessel percutaneous coronary interventions (PCI)
Electrophysiological ablation
Biventricular pacing
Notes: Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 31: Experiences with patient radiation exposure

## Slide 32: Experiences with patient radiation exposure
About two hundred cases of injury have been reported in peer-reviewed journals and public information sources, such as court records.
A wide severity in injuries has been identified spanning the range of mild erythema to deep tissue necrosis.

## Slide 33: Experiences with patient radiation exposure
Patients often unaware of symptoms of injury associated with radiation delivery
Symptoms usually delayed by 1-3 weeks
Combining factors in items 3 and 4 result in a lack of association between lesion and previous cardiologic procedure
Many medical doctors often unfamiliar with this form of skin injury and many have misdiagnosed the etiology

## Slide 34: Experiences with patient radiation exposure
Injuries often associated with prolonged procedures, although short procedures have caused severe effects in a small number of cases
Injuries often associated with X ray transmission through thick body masses (large patients and/or steep beam angulations)
Injuries often associated with high-dose rate modes of operation (high-level control, prolonged cine, high magnification, etc.)

## Slide 35: Experiences with patient radiation exposure
Injuries sometimes associated with multiple procedures
Injuries sometimes associated with unnecessary body parts in beam (arms and breasts)
Injuries infrequently associated with patients who have certain conditions that may sensitize them to radiation injury
For all published injuries, real-time dose monitoring not employed.

## Slide 36: Recognizing radiation injury and effects

## Slide 37: Recognizing radiation injury and effects
When do symptoms of radiation injury occur?
Symptoms may occur promptly, delayed by hours or days, or delayed by weeks.
Delay depends primarily on radiation absorbed dose in skin and the mechanism of and the type of injury.
Most common experience is that one to three weeks pass before patient is aware of symptoms. 
Delay may be exaggerated if symptoms are not visually evident (i.e., they occur in area not visible to patient)

## Slide 38: Thermal burns versus X ray injuries
Heat invades relatively slowly from the outside and must globally affect many exterior cells before reaching deeper tissues.
X rays cause random and very localized damage internally in cells and does not affect surrounding cells.
Heat is sensed before injury occurs.
Symptoms of injury are prompt, progress quickly, and extent of injury is assessed in short order.
Response to treatment becomes apparent in a matter of days.
X rays cause subcellular damage not sensed by organ systems.
Symptoms are often delayed by days or weeks and progress slowly over months.
Treatment is often ineffective and wound progresses until latent damage is manifest.
Reproduced with permission from Louis K. Wagner

## Slide 39: Recognizing radiation injury and effects
When do symptoms of radiation injury occur?
Prompt erythema:
Thought to be due to activation of histamine-like substances that result in dilation of capillaries
Reddening occurs in area of X irradiation
Reddening develops within hours, then promptly fades
Threshold for this effect is about 2 Gy to the skin
Not commonly observed

## Slide 40: Recognizing radiation injury and effects
When do symptoms of radiation injury occur?
Prompt erythema (continued) –
There are anecdotal reports of the following occurring:
Sharp pain on day after procedure
Pain during procedure
Erythema when removed from table (not associated with grounding pads)

## Slide 41: Recognizing radiation injury and effects
Characteristics of radiation injury
Hair loss (depilation):
The threshold dose for hair loss is ~3 Gy.
Permanent hair loss can occur at doses ~ 7 Gy
If hair returns, it may be thinner and of a different shade of color.

## Slide 42: Recognizing radiation injury and effects
Characteristics of radiation injury
Delayed Erythema:
Most frequent experience is skin rash that develops days to weeks after procedure at site of beam entry.
Threshold dose is about 6 Gy.
Initiated by the depletion of cells in the basal layer of the epidermis.
Depending on dose, rash may fade or become prolonged.
When doses are high, treatment of symptoms proves difficult and lesion progresses into more severe stages.

## Slide 43: Recognizing radiation injury and effects
Characteristics of radiation injury
Severe effects:
If erythema prolonged, lesion may progress after additional weeks into desquamation and blistering. Threshold ~ 14 Gy.
In severe cases, ulceration develops typically about 2 -3 months after exposure.
Ulceration then may progress after several months into ischemic dermal necrosis. Threshold dose ~ 18 Gy.
Necrosis is caused by injury to the epithelial lining of dermal vasculature.

## Slide 44: Recognizing radiation injury and effects
Characteristics of radiation injury
If initial lesion heals, long-term effects may result, depending on dose; these include:
      dermal atrophy
      induration
      hyper- or hypo-pigmentation
      telangiectasia
      recurring erythema with possible progression to ulceration and necrosis.

## Slide 45
Effect | Threshold (Gy) | Single-dose Onset
Early transient erythema | 2 | Hours
Main Erythema | 6 | ~10 d
Temporary hair loss | 3 | ~3 wk
Permanent hair loss | 7 | ~3 wk
Dry desquamation | 14 | ~4 wk
Moist desquamation | 18 | ~4 wk
Secondary ulceration | 24 | >6 wk
Late erythema | 15 | ~6 – 10 wk
Ischemic dermal necrosis | 18 | >10 wk
Dermal atrophy (1st phase) | 10 | >14 wk
Dermal atrophy (2nd phase) | 10 | >1 yr
Induration (Invasive Fibrosis) | 10 | 
Telangiectasia | 10 | >1 yr
Late dermal necrosis | >12? | >1 yr
Skin cancer | not known | >5 yr
Recognizing radiation injury and effects
Characteristics of radiation injury

## Slide 46
Both patients with Discoid Lupus Erythematosis received skin doses believed insufficient to cause such effects in normal skin.
Reproduced with permission from Gironet et al, 1998, Ann Dermatol Venerol, 125, 598 - 600
Reproduced with permission from Wagner et al, 1999, Radiology, 213, 773 - 776
27
Radiation-sensitive patients
In right photo, G = graft, R = rib

## Slide 47: PATIENT DOSE MANAGEMENT IN INTERVENTIONAL FLUOROSCOPY
Before the procedure:
Review medical history for sensitizing health factors.
Review history for previous procedures and do physical exam to look for signs of previous low-grade injury.
Counsel patient on risks (don’t forget size and complexity).
During the procedure:
Know and apply good technique (non-trivial).
Manage well fluoro time and fluorography use / monitor dose.
When appropriate, seek assistance / consider postponing procedure / use different projection.
After procedure:
Advise patient if procedure was long or dose was high.
Ask patient to report skin changes to you, the physician.

## Slide 48
Skin rashes that infrequently may develop, and on very rare occasions could degenerate into tissue breakdown and possibly into severe ulcers. This depends on the difficulty of the procedure and your sensitivity to radiation.
PATIENT COUNSELLING -- Potential radiation risks to patients include:
A slightly elevated risk for cancer two or more years later in life, but this risk is typically low in comparison to the normal incidence of human cancer;
Hair loss: If the patient has hair loss on the back, this is usually temporary, but regrowth of hair might not be complete.

## Slide 49: RADIATION DOSE MANAGEMENT IN INTERVENTIONAL RADIOLOGY
Management of radiation through design of equipment;
Management of radiation through maintenance and quality control;
Management of radiation through quality of use.

## Slide 50
Common operational factors associated with radiation injuries
Long procedures with fluoroscopy on-times over the same skin area;
Fluoroscopy through thick body parts (steeply angled projections and/or large patients);
High dose rate modes of operation;
No dose monitoring devices.
Other operational factors associated with radiation injuries
Multiple procedures;
Poorly designed equipment;
Unnecessary body parts in direct radiation field;
Radiation-sensitive patients.

## Slide 51
Both of these publications directly address the broad aspects of radiation management in fluoroscopy. The one on the left specifically addresses interventional work. The one on the right addresses interventional and other fluoroscopy topics.

## Slide 52
If you use too much radiation …
you will get your ass in trouble
RULE 1

## Slide 53
RULE 2
Excess radiation
can catch you
unawares
A fluoroscope can be
a dangerous beast

## Slide 54: Answer True or False
At skin dose of about 2 Gy, it is certain that the patient will land up with skin injury.
Skin injuries can be detected within 1-3 hours after irradiation.
There are occasions when the patient should be counselled about radiation skin injury?

## Slide 55: Answer True or False
Skin injuries are often associated with high-dose rate modes of operation (high-level control), prolonged cine, high magnification, etc.
Threshold for early transient erythema is about 10 Gy to the skin.
If the patient has hair loss on the back, this is usually temporary, but re-growth of hair might not be complete.