## Slide 1: Occupational exposure and protective devices
L 7

## Slide 2: Answer True or False
The occupational dose limit for occupational is 100 mSv/year (effective dose).
A lead apron equivalent to 0.35 mm lead typically can absorb a 50% of the scattered radiation.
Collimation of the radiation field has no influence on the scatter dose.

## Slide 3: Educational objectives
How effective are individual protective items in cath. Lab?
How to monitor personnel dose?
How to estimate personnel effectiveness?

## Slide 4: Outline
Dose limits
Basis for protection, radiation risk and ICRP recommendations
Influence of patient size and operation modes
Personal dosimetry
Protection tools
Some experimental results
Practical advice

## Slide 5: Limits on Occupational Doses (ICRP)*
 | Annual Dose Limit (mSv)
Effective dose, worker | 20
Equivalent dose to lens of eye | 150
Equivalent dose to skin | 500
Equivalent dose to hands and feet | 500
Effective dose to embryo or fetus | 1
Effective dose, public | 1
*Please follow the recommendations as prescribed by your national authority
Notes: A dose of 1 mSv to the fetus during pregnancy means that a badge worn under the apron at abdominal level can exceed this value somewhat since overlying tissue provides some absorption, except in the latter months of pregnancy.

## Slide 6: Limits on Occupational Doses (ICRP)
Effective dose of 20 mSv per year— averaged over a period of 5 years
Should not exceed 50 mSv in any one year
Equivalent skin dose of 500 mSv per year—Limit is set on basis of avoiding deterministic effects
Dose limits do not apply to radiation dose employee receives as part of personal healthcare
Notes: Employees should not wear their personal radiation monitor when have personal x-ray examinations for medical or dental purposes.

## Slide 7: Basic Radiation Protection
Time (T), Distance (D), and Shielding (S)
Time– minimize exposure time
Distance– increasing distance
Shielding– use shielding effectively; portable and pull-down shields; protective aprons; stand behind someone else
Notes: The first two of these radiation protection steps are cost effective– they cost nothing to implement. Using shielding that is in the room effectively also is cost free! Each of these can reduce staff doses by factors from 2 to 20 or more.

## Slide 8: Minimize Exposure Time
Everything you do to minimize exposure time reduces radiation dose!!
  Minimize fluoro and cine times
  Whenever possible, step out of room
  Step behind barrier (or another person) during fluoro or cine
  Use pulsed fluoroscopy– minimizes time X ray tube is producing X rays
Notes: Whatever reduces patient doses reduces staff dose– minimizing fluoro and cine times, and using pulsed fluoroscopy can reduce doses by factors of 2 to 8 or greater.
Stepping out of the room or behind a barrier, especially during a cine run, will reduce staff doses dramatically.

## Slide 9: Maximize Distance – Inverse Square Law
Radiation dose varies inversely with the square of the distance
If you double your distance from source of X rays, your dose is reduced by a factor of 4, i.e., it is 25% of what it would have been!

## Slide 10: Inverse Square Law Helps Protect You
Move from 20 cm to 40 cm, or 1 m to 2 m, from patient, dose rate decreased 4X or to 25%!!
The patient is the source of scattered radiation!!
Do not stand next to patient during fluoro
Step back during cine runs

## Slide 11: Maximize and Optimize Shielding
Leaded shielding reduces doses to 5% or less!
Shielding must be between the patient and the person to be protected
  If back is to patient, need
  protection behind individual
Coat aprons protect back and
  help distribute apron weight
Everyone in the procedure room
must wear a protective apron
Notes: Most aprons available today contain 0.5 mm Pb equivalent. This should be the minimum for use by staff in a cath lab.
Another very important point– Everyone in the procedure room must wear a protection apron!!

## Slide 12: High radiation risk
Occupational doses in interventional procedures guided by fluoroscopy are the highest doses registered among medical staff using X rays.
If protection tools and good operational measures are not used, and if several complex procedures are undertaken per day, radiation lesions may result after several years of work.
Notes: <number>
<footer>
<header>
<date/time>

## Slide 13: ICRP report 85 (2001): Avoidance of Radiation Injuries from Interventional Procedures
Cataract in eye of interventionalist after repeated use of old X ray systems and improper working conditions related to high levels of scattered radiation.

## Slide 14
1- 5 mSv/h
0.5 – 2.5 mSv/h
2- 10 mSv/h
Notes: <number>
<footer>
<header>
<date/time>

## Slide 15: Radiation units used
Dose rates indicated in the slide are “personal dose equivalent” values.
Personal dose equivalent, typically referred in personal dose records as Hp(10) is the dose equivalent in soft tissue, at 10 mm depth and it is measured in Sieverts (Sv).
It is a common practice in RP to directly compare Hp(10) with the annual limit of effective dose (ICRU report 51. Quantities and Units in Radiation Protection Dosimetry. International Commission on Radiation Units and Measurements. Bethesda, MD, USA. 1993).

## Slide 16
Influence of patient thickness and operation modes in scatter dose rate

## Slide 17
Influence of patient thickness: from 16 to 24 cm, scatter dose rate could increase in a factor 5
(from 10 to 50 mSv/h during cine acquisition)

## Slide 18
Influence of operation modes: from low fluoroscopy to cine, scatter dose rate could increase in a factor of 10
(from 2 to 20 mSv/h for normal size)

## Slide 19
Isodose curves for scatter radiation for typical operation conditions and typical patient size

## Slide 20
DETERMINISTIC LENS THRESHOLD AS QUOTED BY ICRP
OPACITIES THRESHOLD
>0.1 Sv/year CONTINUOUS ANNUAL RATE
>0.15 Sv/year CONTINUOUS ANNUAL RATE
CATARACT

## Slide 21
UP TO 2 mSv IN LENS COULD BE RECEIVED IN A SINGLE PROCEDURE
if protection tools are not used
WITH 3 PROCED./DAY IT IS POSSIBLE TO RECEIVE 1500 mSv/year
IN FOUR YEARS WILL BE POSSIBLE TO HAVE LENS OPACITIES

## Slide 22
Patient and staff doses are not always correlated

## Slide 23
Different C-arm angulations, involve very different scatter dose rates (Philips Integris 5000)

## Slide 24
Measuring entrance dose, scatter dose and image quality
Scatter dose detector (lens of the interventionalist position)
Test object to measure image quality, at the isocenter
Flat ionization chamber to measure patient entrance dose

## Slide 25
For scatter dose the orientation of the C-arm is dominant in comparison with the entrance patient dose rate.
Notes: <number>
<footer>
<header>
<date/time>

## Slide 26
Different C-arm angulations can modify the scatter dose rate by a factor of 5
Notes: <number>
<footer>
<header>
<date/time>

## Slide 27
Scatter dose values at the left shoulder of the cardiologist without extra shielding (experimental results from E. Vano)

## Slide 28: Personal dosimetry

## Slide 29: Personal dosimetry ICRP report 85 (2001) states ...
Paragraph 66: The high occupational exposures in interventional radiology require the use of robust and adequate monitoring arrangements for staff.
A single dosimeter worn under the lead apron will yield a reasonable estimate of effective dose for most instances. Wearing an additional dosimeter at collar level above the lead apron will provide an indication of head (eye) dose.

## Slide 30: Personal dosimetry ICRP report 85 (2001) states ...
In addition, it is possible to combine the two dosimeter readings to provide an improved estimate of effective dose (NCRP-122; 1995).
Consequently, it is recommended that interventional radiology departments develop a policy that staff should wear two dosimeters.

## Slide 31: Types of Personal Radiation Monitors
Film
Thermoluminescent dosimeters (TLDs)
Optically stimulated luminescence (OSL) dosimeters
Electronic personal dosimeters
Notes: Each type of dosimeter has its specific characteristics and features.

## Slide 32: Advantages and Disadvantages of Personal Radiation Monitors
Film– sensitive to heat, provides permanent record, minimum dose 0.1 mSv, fading problem, can image (detect motion), maximum monthly readout, film can be re-read after processing
TLDs– some heat sensitivity, no permanent record, minimum dose 0.1 mSv, some fading, no imaging, maximum quarterly readout, no re-read capability
OSL– insensitive to heat, provides permanent record, minimum dose 0.01 mSv, no fading, image capability, quarterly to annual readout, can be re-read during use period

## Slide 33: Advantages and Disadvantages of Personal Radiation Monitors
Electronic dosimeters— insensitive to heat, no permanent record, minimum dose > 0.1 mSv, no imaging capability, calibration can be difficult, must rely on employee for care of device (somewhat delicate), employee must read-out dosimeter and record results, weekly or monthly readout
Notes: Electronic dosimeters are ideal for training new staff and when developing new procedures.

## Slide 35
The use of electronic dosimeters to measure occupational dose per procedure helps in the optimization

## Slide 36: Protection tools

## Slide 37: Personal protective equipment
Registrants and licensees shall ensure that workers are provided with suitable and adequate personal protective equipment.
Protective equipment includes lead aprons, thyroid protectors, protective eye-wear and gloves.
The need for these protective devices should be established by the RPO.

## Slide 38
Weight: 80 grams
Lead equivalent: 0.75mm front and side shields leaded glass
Lead apron typically attenuates >90%
Vest-Skirt Combination distributing 70% of the total weight onto the hips leaving only 30% of the total weight on the shoulders.
Option with light material reducing the weight by over 23% while still providing 0.5 mm Pb protection at 120 kVp

## Slide 39
THYROID PROTECTOR
Protection tools

## Slide 40: Protective Surgical Gloves
Minimal effectiveness
Transmission on the order of 40% to 50%, or more
Costly ($40 US), not reusable
Reduces tactile sensitivity
Dose limit for extremities is 500 mSv
Hands on side of patient opposite of X ray tube so dose rate is already low compared to entrance side
Lead-containing disposable products are environmental pollutants
Notes: The dose reduction offered by radiation protective surgical gloves is not significant. This is particularly true when one considers that the maximum allowable dose for the extremities is 500 mSv/y.

## Slide 41
Attenuation 55% (16 - 24 cm PMMA as phantom).
Attenuation 45 % (16 - 24 cm PMMA as phantom).
Attenuation 15% (16 - 24 cm PMMA as phantom).
But, patient dose increase in a 30%.
Position
(hand, radiation field, and AEC area)
Relative values of hand dose
(for gloves with 0.03 mm Pb)

## Slide 42: Radiation Protection of Hands
Best way to minimize dose to fingers and hand:
Keep your fingers out of the beam!!!
Dose rate outside of the beam and on side of patient opposite X ray tube:
Very low compared to in the beam!!!

## Slide 43
Sometimes your hands could be inside the direct X ray beam
Notes: <number>
<footer>
<header>
<date/time>

## Slide 44
This RP material shall be submitted to a quality control and cleaned with appropriate instructions

## Slide 45
Expensive light protective apron sent to the cleaning hospital service without the appropriate instructions

## Slide 46
Expensive light protective apron sent to the cleaning hospital service without the appropriate instructions

## Slide 47
Expensive light protective apron sent to the cleaning hospital service without the appropriate instructions
Before
After (a bad) cleaning … 1,000$ lost!!

## Slide 48
0.25 mm lead
60 kV; 100%
2 - 3 %
100 kV; 100%
8 - 15 %
Attenuation measured with lead aprons
X ray beam filtration has a great influence!!
Measurements at San Carlos Hospital, Madrid

## Slide 49
0.50 mm lead
60 kV; 100%
< 1 %
100 kV; 100%
3 - 7 %
Attenuation measured with lead aprons
X ray beam filtration has a great influence!!
Measurements at San Carlos Hospital, Madrid

## Slide 50: Ceiling suspended screen
Typically equivalent to 1mm lead
Very effective if well positioned
Not available in all the rooms
Not used by all the interventionalists
Not always used in the correct position
Not always used during all the procedure

## Slide 51
Measures to reduce occupational doses

## Slide 52: Practical advice for staff protection
Increase distance from the patient.
Minimize the use of fluoroscopy and use low fluoroscopy modes.
Acquire only the necessary number of images per series and limit the number of series.

## Slide 53: Practical advice
Use suspended screen and other personal shielding tools available.
Consider the size of the patient and the position of the X ray tube (C-arm angulation).
Collimate the X ray beam to the area of interest.

## Slide 54: Optimization of Radiation Protection
Minimization of dose to patient and staff should not be the goal
Must optimize dose to patient and minimize dose to staff
First: optimize patient dose rate assuring that there is sufficient dose rate to provide adequate image quality
If image quality is inadequate, then any radiation dose results in needless radiation dose!
Notes: In medical imaging the goal is to provide the image quality necessary for the physician to make a diagnosis. Consequently, a certain level of radiation dose is required and it is not possible to minimize patient radiation exposure.

## Slide 55
Be aware of the radiological protection of your patient and you will also be improving your own occupational protection
General recommendation:

## Slide 56: Answer True or False
The regulatory occupational dose limit for the eye lens is 150 mSv/year.
The intensity of the scatter dose (measured at a few cm from the scattering material) in comparison with the direct beam is less than 1/1,000.
LAO projections are more irradiating for staff than the RAO projections.
If the X ray tube is under the table, the scatter dose for the operator is higher for the lens in comparison with the ankles.

## Slide 57: Answer True or False
Typical dose values read from the personal dosimeter (worn under the lead apron) for cardiology staff should be not more than 0.4 mSv/month.
If you are using two personal dosimeters, one under the apron and the second one over the apron, the lens dose could be estimated from the under apron dosimeter.
If you are using two personal dosimeters, the one over the apron, could arrive to measure 1-2 mSv/procedure.

## Slide 58: Additional information

## Slide 59: Some experimental results

## Slide 61
Shoulder dose 0.3 – 0.5 mGy per procedure (without protective screen).
This represents approx. 1 mSv/100 Gy.cm2
High X ray beam extra filtration may represents a 20% reduction.
Ceiling mounted screens represent a reduction factor of 3 (screens are not used during all the procedure or not always in the correct position).

## Slide 62
Vañó et al.
Br J Radiol 1998; 71:954-960
Interventional cardiologist
Interventional
radiologist

## Slide 63: Radiation Monitoring Badge
Plastic filter
Metal filters
Open windows
Open window
Notes: Badges are designed to detect x rays and gamma rays, plus beta particles. With additional components can measure neutrons. However, betas and neutrons are not of interest in interventional imaging.
Explain the use of plastic and metal filters, i.e., to provide ability to discriminate the energy of the radiation being monitored. The open window allows for measurement of beta particles and extremely soft x rays.

## Slide 64
E = 0.5 HW + 0.025 HN
E = Effective dose
HW = Personal dose equivalent at waist or chest, under the apron.
HN = Personal dose equivalent at neck, outside the apron.
If under apron, 0.5 mSv/month, and over apron, 20 mSv/month, E = 0.75 mSv/month

## Slide 66
Conclusion: Use of 0.5 mm lead caps attenuates scatter dose in a factor of 2000 of baseline.

## Slide 67
SUGGESTED ACTION LEVELS FOR STAFF DOSE
Body 0.5 mSv/month
Eyes 5 mSv/month
Hands/Extremities 15 mSv/month
Suggested action levels in staff exposure in interventional radiology
(Joint WHO/IRH/CE workshop 1995)