## Slide 1: Talking about Radiation Dose
L 2
Notes: This sections covers the topics noted but limited to doses in the cardiac catheterization laboratory. It does not address issues and doses associated with nuclear cardiology.

## Slide 2: Answer True or False
The radiation dose a patient gets in catheterization procedure can be measured.
Same amount of radiation falling on the person at level of breast, head or gonads will have same biological effects.
2 mSv/year from natural background radiation represents effective dose.
1 Gy mentioned in relation to PTCA means effective dose.

## Slide 3: Educational Objectives
How radiation dose can and should be expressed, merits and demerits of each quantity for cardiology practice
How representative fluoroscopy time, cine time are for dose to the patient and the staff
Simplified presentation of dose quantities

## Slide 4
20 mg of beta blocker
  Dose outside (in drug) is same as dose inside the patient body
Not so in case of radiation
  Depends upon the absorption
Different expressions for radiation intensity outside (exposure units), absorbed dose [called Dose] in air, in tissue
Absorbed dose
In tissue
In air
Difficult to measure dose inside the body
Measure dose in air, then convert in tissue

## Slide 5: Dose quantities and Radiation units
Dose quantities outside the patient’s body
Dose quantities to estimate risks of skin injuries and effects that have threshold
Dose quantities to estimate stochastic risks
Notes: <number>
<footer>
<header>
<date/time>

## Slide 6: Why so many quantities?
1000 W heater giving off heat (IR radiation) - unit is of power which is related with emission intensity
Heat perceived by the person will vary with so many factors: distance, clothing, temperature in room…
If one has to go a step ahead, from perception of heat to heat absorbed, it becomes a highly complicated issue
This is the case with X rays - can’t be perceived

## Slide 7: Quantities and units
Exposure and exposure rate (R and R/s)
Absorbed dose and KERMA (Gy)
Mean Absorbed Dose in a tissue (Gy)
Equivalent dose H (Sv)
Effective Dose (Sv)
Related dosimetry quantities (surface and depth dose, backscatter factor…..)

## Slide 8: Radiation quantities
Used to describe a beam of x-rays:
  Quantities to express total amount of radiation
  Quantities to express radiation at a specific point
Radiation at a specific point
Photon fluence
Absorbed dose
Kerma
Dose equivalent
Total radiation
Total photons
Integral dose
Notes: <number>
<footer>
<header>
<date/time>

## Slide 9: Exposure: X
Exposure is a dosimetric quantity for ionizing electromagnetic radiation, based on the ability of the radiation to produce ionization in air.
This quantity is only defined for electromagnetic radiation producing interactions in air.

## Slide 10: Exposure: X
Before interacting with the patient (direct beam) or with the staff (scattered radiation), X rays interact with air
The quantity “exposure” gives an indication of the capacity of X rays to produce a certain effect in air
The effect in tissue will be, in general, proportional to this effect in air

## Slide 11: Exposure: X
X = dQ/dm
The exposure is the absolute value of the total charge of the ions of one sign produced in air when all the electrons liberated by photons per unit mass of air are completely stopped in air.

## Slide 12: Exposure: X
The SI unit of exposure is Coulomb per kilogram [C kg-1]
The former special unit of exposure was Roentgen [R]
1 R = 2.58 x 10-4 C kg-1
1 C kg-1 = 3876 R

## Slide 13: Exposure rate: X/t
Exposure rate (and later, dose rate) is the exposure produced per unit of time.
The SI unit of exposure rate is the [C/kg] per second or (in old units) [R/s].
In radiation protection it is common to indicate these rate values “per hour” (e.g. R/h).

## Slide 14: Radiation quantities
X ray beam emitted from a small source (point):
  constantly spreading out as it moves away from the source
  all photons that pass Area 1 will pass through all areas (Area 4)  the total amount of radiation is the same
  The dose (concentration) of radiation is inversely related to the square of the distance from the source (inverse square law)
  D2=D1*(d1/d2)2
Area = 1 Dose = 1
Area = 4 Dose = 1/4
d1=1
d2=2
Notes: <number>
<footer>
<header>
<date/time>

## Slide 15: Dose quantities and radiation units
Absorbed dose
The absorbed dose D, is the energy absorbed per unit mass
  D = dE/dm
      SI unit of D is the gray [Gy]
      Entrance surface dose includes the scatter from the patient ESD  D * 1.4
Notes: <number>
<footer>
<header>
<date/time>

## Slide 16: Absorbed dose, D and KERMA
The KERMA (kinetic energy released in a material)
K = dEtrans/dm
  where dEtrans is the sum of the initial kinetic energies of all charged ionizing particles liberated by uncharged ionizing particles in a material of mass dm
The SI unit of kerma is the joule per kilogram (J/kg), termed gray (Gy).
In diagnostic radiology, Kerma and D are equal.
Notes: <number>
<footer>
<header>
<date/time>

## Slide 17: Absorbed dose in soft tissue and in air
Values of absorbed dose to tissue will vary by a few percent depending on the exact composition of the medium that is taken to represent soft tissue.
The following value is usually used for 80 kV and 2.5 mm Al of filtration :
Dose in soft tissue = 1.06 x Dose in air
Notes: <number>
<footer>
<header>
<date/time>

## Slide 18: Mean absorbed dose in a tissue or organ
The mean absorbed dose in a tissue or organ DT is the energy deposited in the organ divided by the mass of that organ.
Notes: <number>
<footer>
<header>
<date/time>

## Slide 19
Detriment
Radiation exposure of the different organs and tissues in the body results in different probabilities of harm and different severity
The combination of probability and severity of harm is called “detriment”.
In young patients, organ doses may significantly increase the risk of radiation-induced cancer in later life
3 - Dose quantities for stochastic risk
Notes: <number>
<footer>
<header>
<date/time>

## Slide 20: 3 - Dose quantities for stochastic risk
Equivalent dose (H)
The equivalent dose H is the absorbed dose multiplied by a dimensionless radiation weighting factor, wR which expresses the biological effectiveness of a given type of radiation
H = D * wR
the SI unit of H is the Sievert [Sv]
For X-rays is wR=1
For x-rays H = D !!
Notes: <number>
<footer>
<header>
<date/time>

## Slide 21
Mean equivalent dose
in a tissue or organ
The mean equivalent dose in a tissue or organ HT is the energy deposited in the organ divided by the mass of that organ.
3 - Dose quantities for stochastic risk
Notes: <number>
<footer>
<header>
<date/time>

## Slide 22: Tissue weighting factor
To reflect the detriment from stochastic effects due to the equivalent doses in the different organs and tissues of the body, the equivalent dose is multiplied by a tissue weighting factor, wT,
Notes: <number>
<footer>
<header>
<date/time>

## Slide 23
Effective dose, E
The equivalent doses to organs and tissues weighted by the relative wT are summed over the whole body to give the effective dose E
E = T wT.HT
wT : weighting factor for organ or tissue T
HT : equivalent dose in organ or tissue T
3 - Dose quantities for stochastic risk
Notes: <number>
<footer>
<header>
<date/time>

## Slide 24: Dose measurement (I)
Absorbed dose (air kerma) in X ray field can be measured with
  Ionization chambers,
  Semiconductor dosimeters,
  Thermoluminescent dosimeters (TLD)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 25: Dose measurement (II)
Absorbed dose due to scatter radiation in a point occupied by the operator can be measured with a portable ionization chamber
Notes: <number>
<footer>
<header>
<date/time>

## Slide 26: Quantities and units (as shown by the X-ray systems)
Dose Area Product (or Kerma Area Product) (Gy•cm2).
Cumulative skin dose (or cumulative entrance air kerma) (mGy).

## Slide 27: Dose area product (I)
DAP = D x Area
the SI unit of DAP is the Gy•cm2
Area = 1 Dose = 1
Area = 4 Dose = 1/4
d1=1
d2=2
Notes: <number>
<footer>
<header>
<date/time>

## Slide 28: DAP (II)
DAP is independent of source distance:
  D decrease with the inverse square law
  Area increase with the square distance
DAP is usually measured at the level of tube diaphragms
Area = 1 Dose = 1
Area = 4 Dose = 1/4
d1=1
d2=2
Notes: <number>
<footer>
<header>
<date/time>

## Slide 29
In room dosimetric indications (during fluoroscopy or cine runs, dose rate is shown)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 30
The cumulative dose (or cumulative air kerma) is the sum of the dose (air kerma) at the interventional reference point during all segments of an interventional procedure. Typically it is measured in mGy.
Cumulative dose

## Slide 31
Interventional reference point

## Slide 32: Interventional procedures: skin dose
In some procedures, patient skin doses approach those used in radiotherapy fractions
In a complex procedure skin dose is highly variable
Maximum local skin dose (MSD) or peak skin dose is the maximum dose received by a portion of the exposed skin.
Notes: <number>
<footer>
<header>
<date/time>

## Slide 33: Methods to measure PSD*
  Point measurements: thermoluminescent detectors (TLD)
  Area detectors: radiotherapy portal films, radiochromic films, TLD grid
    Large area detectors exposed during the cardiac procedure: between tabletop and back of the patient
Example of dose distribution in a cardiac procedure shown on a radiochromic film as a grading of color
*Peak Skin Dose (PSD) or Maximum skin dose (MSD)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 34: Methods to measure PSD
  Use of films:
    Dose distribution is obtained through a calibration curve of Optical Density vs. absorbed dose
    Slow films:
      require chemical processing
      maximum dose 0.5-1 Gy
    Radiochromic detectors:
      do not require film processing
      immediate visualization of dose distribution
      dose measurement up to 15 Gy
Notes: <number>
<footer>
<header>
<date/time>

## Slide 35: Other related dose parameters
  Fluoroscopy time:
    has a weak correlation with DAP
    But, in a quality assurance programme it can be adopted as a starting unit for
      comparison between operators, centres, procedures
      for the evaluation of protocol optimization
      and, to evaluate operator skill
  Number of acquired images and no. of series:
    Patient dose can be a function of total acquired images
    But dose/image can have big variations
    There is an evidence of large variation in protocols adopted in different centres
Notes: <number>
<footer>
<header>
<date/time>

## Slide 36: Reference levels
3rd level
“Patient risk”
2nd level
“Clinical protocol”
1st level
“Equipment performance”
Dose rate and dose/image
(BSS, CDRH, AAPM)
Level 1 + No. images + fluoroscopy time
Level 2 + DAP + Maximum Skin Dose (MSD)
Reference levels (indicative of the state of the practice): an instrument to help operators to conduct optimized procedures with reference to patient exposure
Required by international (IAEA) and national regulations
For complex procedures reference levels should include:
more parameters
and, must take into account the complexity of the procedures.
(European Dimond Consortium recommendations)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 37
Procedure: | CA | PTCA
DAP (Gycm2) | 57 | 94
Fluoroscopy time (min) | 6 | 16
No. of frames | 1270 | 1355
Reference levels in interventional cardiology
(European proposal 2003)
DIMOND EU project. E.Neofotistou, et al, Preliminary reference levels in interventional cardiology, J.Eur.Radiol, 2003
Notes: <number>
<footer>
<header>
<date/time>

## Slide 38: Quantities and units for staff exposure
Personal dosimetry services typically provide monthly estimates of Hp(10) (mSv), the dose equivalent in soft tissue at 10 mm depth. This is in most of the cases used to estimate the effective dose.
Sometimes, Hp(0.07) (mSv) is also reported: the dose equivalent in soft tissue at 0.07 mm depth).

## Slide 39: Staff dosimetry methods
Exposure is not uniform:
  with relatively high doses to the head, neck and extremities
  much lower in the regions protected by shielding
Dose limits (regulatory) are set in terms of effective dose (E):
  no need for limits on specific tissues
  with the exception of eye lens, skin, hands and feet
The use of 1 or 2 dosemeters may provide enough information to estimate E

## Slide 40
E = 0.5 HW + 0.025 HN
E = Effective dose
HW = Personal dose equivalent at waist or chest, under the apron.
HN = Personal dose equivalent at neck, outside the apron.
If under apron, 0.5 mSv/month, and over apron, 20 mSv/month, E = 0.75 mSv/month

## Slide 41: Personal dosimetry methods
Single dosimeter worn
  above the apron at neck level (recommended) or under the apron at waist level
Two dosimeters worn (recommended)
  one above the apron at neck level
  another under the lead apron at waist level
Lens dose,
optional
Finger dose,
optional
Second dosemeter
outside and above the apron
at the neck, optional
Personal dose
dosemeter behind the lead apron
X-ray
tube
Image
intensifier
Patient
Radiation
protection
measures
Dose limits
of occupational exposure
(ICRP 60)
Effective dose
20 mSv in a year
averaged over a period of 5 years
Anual equivalent dose in the
lens of the eye
150 mSv
skin
500 mSv
hands and feet
500 mSv

## Slide 42: Re-cap
Different dose quantities are able:
  to help practitioners to optimize patient exposure
  to evaluate stochastic and deterministic risks of radiation
Reference levels in interventional radiology can help to optimize procedure
Staff exposure can be well monitored if proper and correct use of dosimeters is routinely applied
Notes: <number>
<footer>
<header>
<date/time>

## Slide 43: Answer True or False
Fluoroscopy time and number of cine frames are enough to estimate patient radiation dose.
Organ doses measured in mSv are similar to entrance patient dose in mGy.
Effective dose can be directly measured with external dosimeters.
Dose area product values are lower if measured far from the X ray tube focus.

## Slide 44: Answer True or False
Reference levels in cardiology should be understood as a limit of dose for patients.
Cumulative dose (as presented by the X ray system) is an indication of the maximum skin dose (peak skin dose).
Personal dosimetry services typically provide monthly estimates of the most irradiated organ doses for the staff.
An increase of approximately 30-40% is observed when comparing skin dose measured in air (without the patient) with the “real” skin dose measured with the patient, due to the backscatter factor.

## Slide 45: Additional information

## Slide 46: Patient dose variability in general radiology
Patient dose distribution in EU survey 1992; lumbar spine Lateral projection
1950s ‘Adrian survey’, UK
  measures of gonadal and red bone marrow dose with an ionization chamber;
  first evidence of a wide variation in patient doses in diagnostic radiology (variation factor: 10,000)
1980s, European countries
  measure of ESD with TLDs and DAP for simple and complex procedures (variation factor: 30 between patients; 5 between hospitals)
1990s, Europe
  trials on patient doses to support the development of European guidelines on Quality Criteria for images and to assess reference levels
  (variation factor: 10 between hospitals)
2000s, NRPB, UK
  UK; National database with patient dose data from 400 hospitals
  (variation factor: 5 between hospitals)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 47: Patient doses in interventional procedures
Also in cardiac procedures, patient doses are highly variable between centres
Need for patient dose monitoring
www.dimond3.org
Notes: <number>
<footer>
<header>
<date/time>

## Slide 48: Staff doses in interventional cardiology
Large variability in staff exposure
Need for staff dose monitoring

## Slide 49: Example 1: Dose rate at different distances
 measured dose rate (air kerma rate) at FSD=70 cm: 18 mGy/min
 dose rate at d= 50 cm: using inverse square law = 18 * (70/50)2 = 18 * 1.96 = 35.3 mGy/min
Image Intensifier
FDD
FSD
d
FDD = focus-detector distance FSD = focus-skin distance
Fixed FOV=17 cm & patient thickness=24 cm
Pulsed fluoro LOW 15pulses/s; 95 kV, 47 mA,
Notes: <number>
<footer>
<header>
<date/time>

## Slide 50: Example 2: Dose rate change with image quality (mA)
  1. pulsed fluoro LOW  47 mA,  dose rate = 18 mGy/min Dose rate at the patient skin including backscatter (ESD=Entrance Surface Dose): ESD= 18 * 1.4 = 25.2 mGy/min
  2. pulsed fluoro NORMAL  130 mA,  dose rate = 52 mGy/min Dose rate at the patient skin including backscatter (ESD=Entrance Surface Dose): ESD= 18 * 1.4 = 73 mGy/min
Image Intensifier
FDD
FSD
d
FDD = focus-detector distance FSD = focus-skin distance
Fixed FOV=17 cm & patient thickness=24 cm
15 pulse/s, FSD=70 cm, 95 kV
Notes: <number>
<footer>
<header>
<date/time>

## Slide 51: Example 3: Dose rate change with patient thickness
  Patient thickness 20 cm,  Dose rate at the patient skin including backscatter ESD = 10 mGy/min
  Patient thickness 24 cm,
   Dose rate at the patient skin including backscatter ESD = 25.2 mGy/min
  Patient thickness 28 cm,  Dose rate at the patient skin including backscatter ESD = 33.3 mGy/min
Image Intensifier
FDD
FSD
d
FDD = focus-detector distance FSD = focus-skin distance
Fixed FOV=17 cm; pulsed fluoro= Low, 15 p/s
Notes: <number>
<footer>
<header>
<date/time>

## Slide 52: Example 3: Patient Thickness (contd.)
Entrance dose rates increase with:
  image quality selected & patient thickness
Notes: <number>
<footer>
<header>
<date/time>

## Slide 53: Example 4: Equipment type
Notes: <number>
<footer>
<header>
<date/time>

## Slide 54: Example 1 DAP
  Patient thickness 24 cm, FOV=17 cm, FDD=100 cm, pulsed fluoro LOW  95 kV, 47 mA, 15 pulse/s  Dose in 1 min @ FSD=70 cm: 18 mGy  Area @ 70 cm: 11.9*11.9=141.6 cm2 DAP= 18 * 141.6 = 2549 mGycm2 = 2.55 Gycm2
   Dose in 1 min @ FSD=50 cm: 18 * (70/50)2 = 18 * 1.96 = 35.3 mGy  Area @ 50 cm: 8.5*8.5=72.2 cm2 DAP= 35.3 * 72.2 = 2549 mGycm2 = 2.55 Gycm2
   DAP is independent of focus to dosimeter distance X ray beam)
Image Intensifier
FDD
FSD
d=50
FDD = focus-detector distance FSD = focus-skin distance
17
11.9
8.5
Notes: <number>
<footer>
<header>
<date/time>

## Slide 55: Example 2: DAP
  Patient thickness 24 cm, FOV=17 cm, FDD=100 cm pulsed fluoro LOW  95 kV, 47 mA, 15 pulse/s  Dose in 1 min @ FSD=70 cm: 18 mGy  Area @ 70 cm: 11.9*11.9=141.6 cm2 DAP= 18 * 141.6 = 2549 mGycm2 = 2.55 Gycm2
   Area @ 70 cm: 15*15=225 cm2 DAP= 18 * 225 = 4050 mGycm2 = 4.50 Gycm2 (+76%)
   If you increase the beam area, DAP will increase proportionately
Image Intensifier
FDD
FSD
d=50
FDD = focus-detector distance FSD = focus-skin distance
17
11.9
8.5
Notes: <number>
<footer>
<header>
<date/time>

## Slide 56: Reference levels
PTCA procedures
Coronary Angiography procedures
DIMOND trial: third-quartile values of single centre data set (100 data/centre)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 57: 2 – Methods for maximum local skin dose (MSD) assessment
On-line methods:
  Point detectors (ion chamber, diode and Mosfet detectors)
  Dose to Interventional Radiology Point (IRP) via ion chamber or calculation
  Dose distribution calculated
  Correlation MSD vs. DAP
Off-line methods
  Point measurements (thermo luminescent detectors (TLD)
  Area detectors (radiotherapy portal films, radiochromic films, TLD grid)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 58: Skin Dose Monitor (SDM)
Zinc-Cadmium based sensor
Linked to a calibrated digital counter
Position sensor on patient, in the X ray field
Real-time readout in mGy

## Slide 59: 2 – Methods for MSD (contd.): on-line methods (I)
  Point detectors (ion chamber, diode and Mosfet detectors)
  Dose to Interventional Radiology Point (IRP) via ion chamber or calculation
15 cm
Isocenter
IRP
15 cm
Isocenter
IRP
Notes: <number>
<footer>
<header>
<date/time>

## Slide 60: 2 –Methods for MSD (contd.): on-line methods (II)
  Dose distribution calculated by the angio unit using all the geometric and radiographic parameters (C-arm angles, collimation, kV, mA, FIID, …)
  Correlation MSD vs. DAP:
    Maximum local skin dose has a weak correlation with DAP
    For specific procedure and protocol, installation and operators a better correlation can be obtained and MSD/DAP factors can be adopted for an approximate estimation of the MSD
Example of correlation between ESD and DAP for PTCA procedure in the Udine cardiac centre
Notes: <number>
<footer>
<header>
<date/time>

## Slide 61: 2 – Methods for MSD: off-line (III)
  Area detectors: TLD grid
    Dose distribution is obtained with interpolation of point dose data
Notes: <number>
<footer>
<header>
<date/time>

## Slide 62: 2 – Methods for MSD: off-line (III)
Dose distribution in a PTCA procedure
Dose distribution for a RF ablation
  Area detectors: TLD grid
    Example of dose distributions
Notes: <number>
<footer>
<header>
<date/time>

## Slide 63: Exercise 1: Evaluation of MSD
  A PTCA of a patient of 28 cm thickness, 2000 images acquired, 30 min of fluoroscopy:
  System A: 2000*0.4 mGy/image=0.8 Gy 30 min * 33 mGy/min=0.99 Total cumulative dose = 1.79 Gy
  System B: 2000* 0.6 mGy/image=1.2 Gy 30 min* 50 mGy/min= 1.5 Gy Total cumulative dose = 2.7 Gy
   Cumulative skin dose is a function of system performance or image quality selected
Notes: <number>
<footer>
<header>
<date/time>

## Slide 64: Exercise 2: Evaluation of MSD
  A crude estimation of MSD during the procedure can be made from the correlation between MSD and DAP in PTCA procedure:
  Example:
  A PTCA with DAP= 125 Gycm2
  MSD= 0.0141*DAP = = 0.0141*125= 1.8 Gy
  (with linear regression factor characteristic of the installation, procedure and operator)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 65: Effective dose assessment in cardiac procedures
Organ doses and E can be calculated using FDA conversion factors (FDA 95-8289; Rosenstein) when the dose contribution from each X ray beam used in a procedure is known
Complutense University (Madrid) computer code allows to calculate in a simple manner organ doses and E (Rosenstein factors used)
Notes: <number>
<footer>
<header>
<date/time>

## Slide 66: Example 1
Effective dose quantity allows to compare different type of radiation exposures:
Different diagnostic examination
Annual exposure to natural background radiation

## Slide 67
For a simple evaluation, E can be assessed from total DAP using a conversion factor from 0.17 to 0.23 mSv/Gycm2 (evaluated from NRPB conversion factors for heart PA, RAO and LAO projections)
Example:
CA to a 50 y old person performed with a DAP=50 Gycm2
Effective dose E = 50 * 0.2 = 10 mSv
Stochastic risk: R=0.01 Sv *0.05 deaths/Sv = 0.0005 (5/10000 procedures)
Compare with other sources: Udine cardia center: CA: mean DAP=30 Gycm2  E = 6 mSv PTCA:mean DAP=70 Gycm2  E = 14 mSv MS-CT of coronaries  E  10 mSv
Example 2: Effective dose assessment in cardiac procedures
Notes: <number>
<footer>
<header>
<date/time>

## Slide 68: Staff doses per procedure
High variability of staff dose/cardiac procedure as reported by different authors
Correct staff dosimetry and proper use of personal dosimeters are essential to identify poor radiation protection working conditions

## Slide 69
Staff dosimetry methods (comments)
Assessment of E is particularly problematic due to the conditions of partial body exposure
Use of dosimeter worn outside and above protective aprons results in significant overestimates of E.
On the other hand, positioning the monitor under the protective apron significantly underestimates the effective dose in tissues outside the apron.
Multiple monitors (more than 2) are too costly and impratical.

## Slide 70: Protective devices influence
Protective devices:
  Lead screens: suspended, curtain
  Leaded glasses
  Lead apron
  Collar protection,
  influence radiation field.
   Only proper use of personal dosimeter allows to measure individual doses
Lens dose,
optional
Finger dose,
optional
Second dosemeter
outside and above the apron
at the neck, optional
Personal dose
dosemeter behind the lead apron
X-ray
tube
Image
intensifier
Patient
Radiation
protection
measures
Dose limits
of occupational exposure
(ICRP 60)
Effective dose
20 mSv in a year
averaged over a period of 5 years
Anual equivalent dose in the
lens of the eye
150 mSv
skin
500 mSv
hands and feet
500 mSv

## Slide 71: Exercise 1: annual staff exposure
Operator 1: 1000 procedures/year
  20 Sv/proc
  E= 0.02*1000 = 20 mSv/year = annual effective dose limit
Operator 2: 1000 proc/year
  2 Sv/proc
  E= 0.002*1000=2 mSv/year = 1/10 annual limit