## Slide 1
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
Other medical specialities that use fluoroscopy L06C

## Slide 2: Objectives
To examine published data describing standard practice and reference values for fluoroscopy use in medical specialties
To become familiar with patient and staff doses and fluoroscopy times during specialist procedures and identify specific opportunities for occupational and patient dose reduction
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 3: Urology - Typical Doses by Procedure
Procedure | Approx Patient Dose* per procedure (mSv/procedure) | Typical Screening Time
extracorporeal shockwave lithotripsy (ECSL) | 1.6 mSv/procedure (Ref 1) | 204 sec fluoro 4 spot films
voiding cystourethrography (VCUG). | 4.3 mSv/procedure (Ref 2) | 36 sec fluoroscopy 12 spot films
1 RADIATION DOSES TO PATIENTS FROM EXTRACORPOREAL SHOCK WAVE LITHOTRIPSY. Health Physics. 90(6):583-587, June 2006.
RADIATION EXPOSURE DURING STANDING VOIDING CYSTOURETHROGRAPHY IN WOMEN UROLOGY 67: 269–274, 2006.
NOTE: Background radiation dose ~ 3 mSv/yr, so these doses are about ½ & 1.4 x annual background dose.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 4: Urology - Typical Doses by Procedure
Procedure | Mean effective dose (mSv) values
Abdomen radiography (AP) | 0.7
Intravenous Urogram (6 films) | 2.5
Micturating cysto urethrogram | 1.2
Cystography | 1.8
Lithotripsy | 1.3
Nephrostomy | 3.4
Percutaneous nephrolithotomy (PCNL) | 4.5
Ureteric stenting | 4.7
CT abdomen | 10
Renal angiogram | 2 to 30
Kidney stent insertion | 12.7
<number>
https://rpop.iaea.org/RPOP/RPoP/Content/InformationFor/HealthProfessionals/6_OtherClinicalSpecialities/Urology/index.htm

## Slide 5: Typical Urology Procedures Times*
 | Average Screening Time (sec) | Number of Film Images
IVU | 38 | 6
MCU | 156 | 5
Nephrostogram | 245 | 3
Nephrostomy | 413 | 1
*Doses to Patients from Medical X ray Examinations in the UK – 2000 Review NRPB-W14 (National Radiological Protection Board)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 6: Occupational Exposure – Urology
* 20 cases monitored for 6 weeks: 8 percutaneous nephrolithotomies, 7 retrograde pyelograms and insertions of ureteric stents, and 5 ureteroscopies for calculous disease. * Radiation doses measured outside & beneath thyroid shield with TLD. * Doses multiplied by thyroid weighting factor (0.05) to convert to effective dose (mSv). * Results: Surgeon no thyroid shield ~ 0.46 mSv, with thyroid shield 0.02 mSv ~ background (0.01 mSv).
RADIATION EXPOSURE DURING FLUOROSCOPY: PROTECTING OUR THYROIDS? V. TSE et al Aust. N.Z. J. Surg. (1999) 69, 847–848
Lesson – Wear thyroid shield and use distance to reduce exposure
Effective dose due to exposure @ thyroid for Surgeon & Nurse
= no thyroid shield
= with thyroid shield
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 7: Urologic Surgery Radiation Shield
Yang R et al, Radiation Protection during Percutaneous Nephrolithotomy (PCNL): A New Urologic Surgery Radiation Shield. J of Endourology 2002;16, 727-731
Shield reduced doses to urologists by 96% at 25 cm & 71% at 50 cm from source
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 8: Gynecology
Hysterosalpingography (HSG). Contrast enhanced fluoroscopy of the uterus
May be followed by fluoroscopy guided therapeutic procedure
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 9: Typical Radiation Characteristics
Hysterosalpingogram
NRPB Survey Published 2000
  22 Hospitals 49 Rooms, 1338 Patients
    Mean time = 56 sec
    Mean number films = 2
  Variations in Mean Dose-Area Product (Gy-cm2)
    Mean of all rooms = 3.5 Gy-cm2
    Min = 0.4 Gy-cm2
    Max = 15.7 Gy-cm2
  Factor of 39 between min and max means
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 10: Typical Variations in Mean Dose -Hysterosalpingogram
NRPB-W14 Doses to Patient from Medical X ray Examinations in the UK – 2000 Review
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 11: Anaesthesia & Venous Line Placement
 | Typical Mean Screening Times (Sec) | Number of Film Images
Hickman line*1 | 104 sec | 1
PICC*2 | 96 sec | 
*1 - NRPB-W14 Doses to Patient from Medical X ray Examinations in the UK – 2000 Review
*2 - Radiation Doses from Venous Access Procedures - Radiology: Volume 238: Number 3—March 2006 Erik S. Storm et al
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 12: Venous Line placement – Risk of skin injury
Dose data analyzed for 1010 instances of 6 different venous access placement procedures.
Results: No procedure yielded cumulative dose > 950 mGy or a peak skin dose of more than 760 mGy. All < 2000 mGy threshold for skin injury. Therefore, no risk of skin injury from venous line placement.
Radiation Doses from Venous Access Procedures - Radiology: Volume 238: Number 3—March 2006 Erik S. Storm et al
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 13: Line Placement - Variability
Note highly skewed distribution of fluoroscopy time typical of many fluoroscopy procedures.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 14: Summary
There is high degree of variability in average doses between different medical clinics. This implies there is significant opportunity for patient dose reduction if the top quartile of facilities were able to operate at doses similar to their peers.
There is a highly skewed distribution of screening times/doses at any facility. Average doses are normally well below the limits that deterministic effects are likely to occur. Nevertheless, when an extended procedure is performed the threshold for deterministic skin effects (2000 mGy) may be reached if bad radiation protection practices are employed.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 15: Dose reduction for patients
<number>
https://rpop.iaea.org/RPOP/RPoP/Content/Documents/Whitepapers/poster-patient-radiation-protection.pdf

## Slide 16
<number>
Dose reduction for patients
https://rpop.iaea.org/RPOP/RPoP/Content/Documents/Whitepapers/poster-patient-radiation-protection.pdf

## Slide 17
<number>
Dose reduction for patients
https://rpop.iaea.org/RPOP/RPoP/Content/Documents/Whitepapers/poster-patient-radiation-protection.pdf

## Slide 18
<number>
Dose reduction for patients
https://rpop.iaea.org/RPOP/RPoP/Content/Documents/Whitepapers/poster-patient-radiation-protection.pdf

## Slide 19
9. Minimize number of frames and cine runs to clinically acceptable level. Avoid using the acquisition mode for fluoroscopy
10. Use collimation
<number>
Dose reduction for patients
https://rpop.iaea.org/RPOP/RPoP/Content/Documents/Whitepapers/poster-patient-radiation-protection.pdf

## Slide 20: New Developments in Dose Reduction
Collimation Without Radiation View Last image hold (LIH) & adjust collimation with graphical overlay on image.
Patient Positioning Without Radiation Position patient via graphical display showing central beam location & edges of field on LIH. (Central beam indictor moves on display as table is moved).
Automatic Beam Filtration Adds filtration to X ray beam to decrease patient dose. Amount of filtration based on patient (e.g. 0.9 mm Cu for small patient, 0.2 mm Cu for large patient.)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 21: Collimation without Radiation
Last Image Hold with superimposed graphical collimation
Image after pressing fluoro pedal
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 22: Patient Positioning without Radiation
+
+
+
LIH shows “+” at center of field (1) and new position (2) after moving table towards patient’s feet
Image after pressing fluoro pedal. Note new center position.
(1)
(2)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 23: Interventions to Reduce staff Dose
Personal Protection
  Lead Apron with lead equivalence of 0.25-0.5 mm (>90% protection from scattered radiation)
  Leaded upper body shields (>90% protection from scattered radiation)
  Eyewear
  Thyroid
  Effective but not universally utilized
    Thyroid shield 42% “never”
    Eyewear 95% “never”
    Monitor badge 32% “never”
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06B. Radiation Exposure in Gastroenterology
Notes: <number>
  Rajoriya N, Mee A. A UK National Survey into Safety Awareness and Practice in Diagnostic and Therapeutic ERCP. 2008. Gastrointest Endosc; 67: AB231

## Slide 24: Protection tools
CURTAIN
THYROID
SCREEN AND GOGGLES
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06B. Radiation Exposure in Gastroenterology
LEAD APRON
Notes: <number>

## Slide 25: Personal dosimetry
Several personal dosemeters are recommended
From: Avoidance of radiation injuries from interventional procedures. ICRP draft 2000
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06B. Radiation Exposure in Gastroenterology
20 mSv*
*ICRP Statement on Tissue Reactions. Approved by the Commission on April 21, 2011
Notes: <number>

## Slide 26: Factors affecting staff doses
16.2: Optimization of protection in fluoroscopy
<number>
Tube undercouch position reduces, in general, high dose rates to the specialist’s eye lens
THE BEST
CONFIGURATION
INTENSIFIER UP
X-RAY TUBE DOWN
SAVES A FACTOR OF
3 OR MORE IN DOSE
IN COMPARISON
TO:
X-RAY TUBE UP
INTENSIFIER DOWN

## Slide 27: Example of dose rate around mobile C-arm
16.2: Optimization of protection in fluoroscopy
<number>
Patient
Image Intensifier
100 cm 50 cm 0
Scale
1.2
3
6
12
X-ray tube
All Contour values in µGy/min

## Slide 28: Staff protection
<number>
More information: https://rpop.iaea.org/RPOP/RPoP/Content/Documents/Whitepapers/poster-staff-radiation-protection.pdf

## Slide 29: New versus Old Equipment
Skin Surface Fluoro Exposure Rates for Abdomen Study New vs. Old fluoro units @ University of Michigan “Patient” = Modified ANSI Abdomen phantom:
New Fluoro System Old Fluoro System 12.5 pps, 0.2 mm Cu 3.75 pps, No Cu
Mode kVp mA mGy/min kVp mA mGy/min
Normal
MAG 1
MAG 2
65
74
81
1.1
1.6
2.2
2.6
4.9
8.0
79
84
92
1.8
2.5
3.5
15.8
23.6
35.6
New equipment can reduce doses by factors of ~2 to 6
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy

## Slide 30: Special groups: Children
Question: Are there specific protocols for children?
Yes. If children or young adults are subjected to CT scanning, paediatric protocols or optimised protocols (reduced exposure parameters compared to adult protocol) are recommended. It is important that children get as little radiation exposure as possible because they are more sensitive to radiation than adults and they have a longer life expectancy.
<number>
https://rpop.iaea.org/RPOP/RPoP/Content/SpecialGroups/2_Children/index.htm

## Slide 31: Thank you
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06C. Other medical specialities that use fluoroscopy