## Slide 1: Understanding radiation units L02
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy

## Slide 2: Educational Objectives
How radiation dose can and should be expressed?
Radiation quantities useful for dose to doctors that use fluoroscopy
Radiation quantities useful for patient doses from fluoroscopic procedures
Individual monitoring
The inverse square law
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
<number>

## Slide 3
500 mg of anti-inflammatory drug.
  Dose outside (in drug) is same as dose inside the patient’s body
Is the situation same with radiation?
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 4
Not so in case of radiation.
  It depends upon the absorption
Absorbed dose
In tissue
on body surface
(in air)
Difficult (rather not possible) to measure dose inside the body
Solution: measure dose in air, then estimate/calculate dose in tissue
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 5: Examples of radiation measures
Watt
Attenuation by paper or metal object
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 6: Why so many quantities?
1000 W heater giving heat (IR radiation) - unit is infra red power which is related with emission intensity
Heat perceived by the person will vary with so many factors: distance, clothing, temperature in room…
If one has to go a step ahead, from perception of heat to heat absorbed, it becomes a highly complicated issue
This is the case with X rays – They can’t even be perceived
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 7
Whole body dose in mSv
Organ dose in mGy
Dose falling on the body in mGy
How to define mGy, mSv?
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 8
mGy
Equivalent dose
to
whole body
(mSv)
mGy
mGy
mGy x wt=mSv
Effective dose (mSv)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 9: Understanding of the Radiation Units
How much radiation do we get from natural sources?
  Around 1-3 mSv (which dose is this?)
How much radiation does a patient get from a chest radiograph?
  around 0.02 mSv (effective dose)
Staff dose obtained from radiation badge?
  Effective dose
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
https://rpop.iaea.org/RPOP/RPoP/Content/InformationFor/Patients/information-public/index.htm#

## Slide 10: Absorption & biological effects
Different tissues have different sensitivities to ionizing radiation
Same amount of radiation will have different effect in different tissues
The higher the weighting factor, the more sensitive the tissue
Tissue | wT
Breast, Bone marrow, colon, Lung, stomach, remainder tissues | 0.12
Gonads | 0.08
Bladder, Esophagus, liver, thyroid | 0.04
Bone surface, brain , salivary gland, skin | 0.01
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 11
Typical effective doses from radiological examinations
(expressed as equal number of chest X rays)
number of chest X rays
Relative Dose Received
0
50
100
150
200
Arm, head,ankle & foot (1)
Head & Neck (3)
Head CT (10)
Thoracic Spine (18)
Mammography, Cystography (20)
Pelvis (24)
Abdomen, Hip, Upper & lower femur (28)
Ba Swallow (30)
Obsteric abdomen (34)
Lumbo-sacral area (43)
Cholangiography (52)
Lumber Myelography (60)
Lower abdomen CT male (72)
Upper Abdomen CT (73)
Ba Meal (76)
Angio-head, Angio-peripheral (80)
Urography (87)
Angio-abdominal (120)
Chest CT (136)
Lower Abd. CT fem. (142)
Ba enema (154)
Lymphan. (180)
mSv
.05
0.15
0.49
0.92
1.0
1.22
1.4
1.5
1.7
2.15
2.59
3.0
3.61
3.67
3.8
4.0
4.36
6.0
6.8
7.13
7.69
9.0
With modern systems
eff dose for chest =0.02 mSv
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 12: Staff exposure- which Units?
Effective dose-overall risk, personal dose equivalent Hp(10)
Skin dose-injuries (fingers, legs)
Organ dose-cataract, thyroid, gonads
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 13: Limits on Occupational Doses (ICRP)*
 | Annual Dose Limit (mSv)
Effective dose, worker | 20
Equivalent dose to lens of eye | 201
Equivalent dose to skin | 500
Equivalent dose to hands and feet | 500
Effective dose to embryo or fetus | 1
Effective dose, public | 1
*Please follow the recommendations as prescribed by your national authority
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
1In a recent statement ICRP (Statement on Tissue Reactions Approved by the Commission on April 21, 2011) the ICRP proposed a limit of 20 mSv/yr, averaged over 5 years, not exceeding 50 mSv at any single year.
Notes: <number>
A dose of 1 mSv to the fetus during pregnancy means that a badge worn under the apron at abdominal level can exceed this value somewhat since overlying tissue provides some absorption, except in the latter months of pregnancy.

## Slide 14: Patient doses
Peak skin dose
Cumulative air kerma
kerma Area Product (KAP or DAP)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 15: Peak skin dose
Example of dose distribution in a Coronary angiography procedure shown using self-darkening film.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Dose is determined by measuring
darkening and relating it to the films
Characteristic curve
After scanning the
Film and using the dose
curve, a dose map can
Be created
1
2
3
Notes: <number>

## Slide 16: Mean absorbed dose in a tissue or organ
SI unit of D is the gray [Gy]
Absorbed dose can not be measured directly during a fluoroscopic procedure.
The absorbed dose (D) is important for describing the risk of radiation effects. It relates to how much radiation energy gets deposited into a given target mass (e.g., skin, eye, thyroid gland).
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>

## Slide 17: Commonly used patient dose descriptors
Radiography and fluoroscopy:
  Entrance Surface Dose (ESD) [mGy]
    by measurements (TLD or film)
    calculated by tube output factors and exposure geometry
  Dose Area Product (DAP) [Gy-cm²]
    by real-time measurement with DAP meter
Computed Tomography:
  Computed Tomography Dose Index (CTDI) [mGy]
  Dose length product (DLP) [mGycm]
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 18: Dose Area Product (DAP) or air-Kerma Area Product (KAP)*
The dose-area product (DAP) quantity is defined as the dose in air in a plane, integrated over the area of interest. DAP is expressed in Gy-cm².
DAP is automatically measured in real-time on most recent fluoroscopy systems
*preferred term
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 19: Fluoroscopy Kerma- area product
KAP = D x Area
the SI unit of KAP is the Gy.cm2
Current name KAP
KAP is independent of distance from the source
Area = 1 Dose = 1
Area = 4 Dose = 1/4
d1=1
d2=2
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>

## Slide 20
Real-time KAP values displayed on monitor
KAP
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 21: Entrance surface dose (ESD)
The ESD is measured on the surface of a patient or a phantom
It can also be calculated by using the tube output and geometrical description of the procedure
It can be used as a metric to estimate risk of deterministic skin effects
TLD or film
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 22: Using KAP and ESD in practice
Endovascular aneurysm repair (EVAR)
Repair of abdominal aortic aneurism (AAA)
R Weerakkody et al, BJS 2008
Example from vascular surgery procedure
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>
Interventional neuroradiology is increasingly used for treatment of vascular lesions in the brain or spinal cord and their surrounding tissues. Minimal invasive treatment, less traumatic fro the patient than surgery.less aexpensive, shorter stay. Advances in the design of catheters and guidewires have allowed increasingly complex and time consuming procedures to be carried out. Mooney
AVM=arteriovenous malformation

## Slide 23: Using KAP and ESD in practice
Exposure factors for four projections commonly used in orthopaedic surgery
From N Theocharopoulos et al; Journal of Bone and Joint Surgery 2003
Example from orthopaedic surgery procedure
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>
Interventional neuroradiology is increasingly used for treatment of vascular lesions in the brain or spinal cord and their surrounding tissues. Minimal invasive treatment, less traumatic fro the patient than surgery.less aexpensive, shorter stay. Advances in the design of catheters and guidewires have allowed increasingly complex and time consuming procedures to be carried out. Mooney
AVM=arteriovenous malformation

## Slide 24
Individual monitoring and exposure assessment
Occupational exposure
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 25: Thank you
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 26
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy L02. understanding radiation units
<number>
Extra Slides

## Slide 27
Kerma
Kinetic Energy Released in Matter
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 28: Kerma Kinetic Energy Released in Matter
KERMA is defined as the sum of the initial kinetic energies of all electrons released by X ray photons per unit mass.
It is no longer recommended to use the term “absorbed dose in air” (ICRU 2005) when speaking of ionization in air.
KERMA is a precursor to absorbed dose.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units

## Slide 29: Absorbed dose, D and KERMA
The KERMA (kinetic energy released in a mass)
The SI unit of kerma is the joule per kilogram (J/kg), termed gray (Gy).
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>

## Slide 30: Dose quantities and radiation units
Absorbed dose
The absorbed dose D, is the energy absorbed per unit mass in a medium
      SI unit of D is the gray [Gy]
      Entrance surface dose includes the scatter from the patient ESD  D * 1.4
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>

## Slide 31: Absorbed Dose
The absorbed dose D, is the energy absorbed per unit mass in a medium
      SI unit of D is the gray [Gy]
      Entrance surface dose includes the back-scatter from the patient ESD  D * 1.4
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>

## Slide 32: Mean absorbed dose in a tissue or organ
The mean absorbed dose in a tissue or organ DT is the energy deposited in the organ divided by the mass of that organ.
<number>
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>

## Slide 33
The equivalent dose H is the average absorbed dose in an organ multiplied by a dimensionless radiation weighting factor, wR which expresses the biological effectiveness of a given type of radiation
H = D * wR
      the SI unit of H is the sievert [Sv]
For X rays is wR=1
Note: This unit H, is not often used in radiological imaging and is often confused with Effective dose
Equivalent dose (H)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>

## Slide 34
The equivalent doses to organs and tissues weighted by the relative wT are summed over the whole body to give the effective dose E
E = sum of (equivalent doses x wT)
wT : weighting factor for organ or tissue T
This unit is OFTEN used in radiological imaging and can be used to compare dose efficiency of competing imaging examinations
Effective dose, E
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L02. Understanding radiation units
Notes: <number>