## Slide 1
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
Radiation protection for patients in orthopaedic surgery L06A
Notes: Radiation protection for doctors using fluoroscopy

## Slide 2
Orthopaedic Surgeons
Anesthetists
Operating room personnel
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Target audience
<number>

## Slide 3: Key topics
Why is it necessary to consider radiation protection of patients?
How do X ray technique and physical factors affect patient dose?
What is the role of the operator in patient dose management?
How to manage patient dose using physical and equipment factors?
Staff radiation protection
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
<number>
Notes: Radiation protection for doctors using fluoroscopy
<number>

## Slide 4: Why is it necessary to consider patients protection?
Patient is irradiated by the direct beam
Medical personnel is irradiated by the scatter radiation
Patients may undergo repeated radiation procedures
A patient may receive in one procedure a dose equivalent to dose the staff may receive in one year
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 5
There are no fluoroscopy time constraints.
Patient entrance dose rates constrained for fluoroscopy but not for acquisitions.
Poor fluoroscopy technique can multiply patient dose rates many times above normal (>10 times)
Implies
There is a potential for high patient doses and skin injury.
Why is it necessary to consider patients protection?
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 6: Why is it necessary to consider patients protection?
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
15 minutes of fluoroscopy at 40 mGy/min skin dose rate
cumulative skin dose: 0.6 Gy
With thick patients, the radiation dose can be quite high with the possibility of radiation injury
X ray system not optimized and operators not trained in radiation protection could increase patient dose by a factor of 10:
Skin necrosis from Coronary Angioplasty Skin Doses > 20 Gy
>100 minutes fluoro time
<number>
Notes: Radiation protection for doctors using fluoroscopy

## Slide 7: The objectives of patient radiation protection are:
To protect the patient from deterministic effects, e.g., skin burns
To optimize X ray exposure to minimize risk of stochastic effects, e.g., development of cancer
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>
To eliminate

## Slide 8: Basic principles
Justification
  avoid unnecessary exams and unnecessary images
Optimization
  choose factors and perform the exam to yield the required diagnostic information while minimizing the dose to the patient.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 9: Basic principles
Dose limitation
Keep dose to patient As Low as Reasonably Achievable (ALARA) (but, must not be so low that images become non-diagnostic)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 10: Factors affecting patient dose in fluoroscopy
Patient entrance surface dose rate
X ray beam area
Beam ON time
(Note: these same factors influence staff doses)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 11: Factors affecting patient dose in fluoroscopy
Patient dependent factors:
  body mass or body thickness in the beam
  complexity of the lesion and anatomic target structure
  previous radiation exposure
  radiosensitivity of some patients
Equipment dependent factors:
  Setting of dose rates in pulsed fluoro- and continuous fluoro mode
  appropriate quality control
  last image hold, acquisition
  virtual collimation.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 12: Factors affecting patient dose in fluoroscopy
The main procedure related factors:
  number of radiographic frames per run
  Collimation
  fluoroscopic and radiographic acquisition modes
  fluoroscopy time
  wedge filter
  Magnification
  distance of patient to image receptor (image intensifier or flat panel detector)
  distance between X ray tube and patient
  tube angulations.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
https://rpop.iaea.org/RPOP/RPoP/Content/InformationFor/HealthProfessionals/6_OtherClinicalSpecialities/Orthopedic/index.htm#ref2

## Slide 13: Factors affecting patient entrance surface dose rate
Thickness & composition of patient.
X ray beam quality (kVp, filtration)
II Mag mode (Normal, Mag 1, Mag 2, etc.)
II Dose mode (low, medium, high)
Pulse rate and pulse width for pulsed fluoro
Anti-scatter grid
Angulation
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 14: Image formation
X ray
source
Absorbed radiation
Scattered radiation
Transmitted radiation
Patient body
Primary (direct) beam
Attenuation
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>

## Slide 15: Image formation
X ray
source
Radiation pattern
Patient body
Visible image
Image
receptor – II or flat panel
1) Spatially ~ uniform beam enters patient
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>

## Slide 16: Image formation
X ray
source
2) X rays interact in patient, rendering beam non-uniform
Patient body
Image
receptor – II or flat panel
1) Spatially ~ uniform beam enters patient
3) Non-uniform beam exits patient; Pattern of non-uniformity is the image
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>

## Slide 17: Image formation
Only a small percentage (typically ~1%) penetrate through to create the image.
As beam penetrates patient, x rays ionize tissue
Beam entering patient typically 100 - 500x
more intense than exit beam
100 % in
1 % out
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>

## Slide 18: Risk of injury
Lesson: Entrance skin tissues receives highest dose of X rays and are at greatest risk of injury.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 19: Skin Entrance Dose and kVp
Use of higher kVp beams usually reduces patient skin entrance dose.
Reason: Higher kVp X ray beams are more penetrating
General rule: Increase of kVp by 15% can decrease mA by factor of 2 (for same dose at image intensifier) and this reduces skin dose by 35%
Disadvantage of using higher kVp: Decreased subject contrast
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 20: Inverse Square Law
source
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 21: Inverse Square Law
X ray intensity decreases rapidly with distance from source; conversely, intensity increases rapidly with closer distances to source.
d
d/2
d/4
d/8
Intensity
64
16
4
1
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 22: Physical factors and challenges to radiation management
Lesson: Understanding how to take advantage of the rapid changes in dose rate with distance from source is essential knowledge for good radiation protection practice.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 23: Inverse Square Law & the Patient
All other conditions unchanged, moving patient toward or away from the X ray tube can significantly affect dose rate to the skin
1 unit
4 units
16 units
64 units
d
d/2
d/4
d/8
Lesson: Keep the X ray tube at the practicable maximum distance from the patient.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 24: Inverse Square Law & The Image Receptor (Film or Image Intensifier)
All other conditions unchanged, moving image receptor toward patient lowers radiation output rate and lowers skin dose rate.
4 units of intensity
Image
Receptor
2 units of intensity
Image
Receptor
Image
Receptor
Lesson: Keep the image receptor as close to the patient as is practicable for the procedure.
Remember, ABC adjusts dose to maintain same image brightness
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 25
Backscatter from thigh- high dose to operator
Position prevents close positioning of II
Forward scatter towards the operator is attenuated by mass of thigh
Patient at edge, allows close positioning of II
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 26: Scatter Levels Hip Lat Cross Table Projection* (μSv per 1000 cGy cm2)
Distance (m) | -1 | -0.5 | 0 | 0.5 | 1
1.5 | 1 | 1 | 1 | 1 | 1
1 | 1 | 2 | 2 | 2 | 1
0.5 | 2 | 5 | 5 | 3 | 2
0 | 3 | 15 | 29 | 6 | 2
 | Feet |  |  |  | Head
0 | 73 | 252 | 1080 | 114 | 11
0.5 | 73 | 160 | 301 | 104 | 8
1 | 48 | 70 | 105 | 85 | 24
1.5 | 24 | 37 | 48 | 43 | 30
Dose rate substantially higher on X ray focus side of patient compared to Image intensifier side because of scatter from the patient
*Occupational exposure from common fluoroscopic projections used in orthopedic Surgery Nicholas Theocharopoulos et al Journal of Bone and Joint Surgery; Sep 2003; 85, 9;
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
X ray tube side
Image intensifier
side

## Slide 27
Fixation of Hip Fractures
*Scattered radiation during fixation of hip fractures
J. A. Alonso, D. L. Shaw, A. Maxwell, G. P. McGill, G. C. Hart
From Bradford Royal Infirmary, England J Bone Joint Surg [Br] 2001;83-B:815-8.
As distance from source increases
Radiation decreases rapidly
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 28: Variability Of Occupational Exposure
Procedure | Approx Surgeon Dose* per procedure (μSv/procedure) with 0.5 mm lead apron worn | Screening Time
Hip | 5 | 25 sec/patient
Spine | 21 | 2 min/patient
Kyphoplasty | 250 | 10 min/patient
*Occupational exposure from common fluoroscopic projections used in orthopedic Surgery Nicholas Theocharopoulos et al Journal of Bone and Joint Surgery; Sep 2003; 85, 9;
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 29: Scatter Levels Spine Lat Projection* (μSv per 1000 cGy cm2)
Distance (m) | -1 | -0.5 | 0 | 0.5 | 1
1.5 | 1 | 1 | 1 | 2 | 2
1 | 2 | 3 | 1 | 3 | 3
0.5 | 4 | 10 | 0 | 10 | 4
0 | 14 | 12 | 79 | 18 | 2
 | Feet |  | ↑ |  | Head
0 | 12 | 46 | 215 | 37 | 13
0.5 | 25 | 88 | 241 | 141 | 51
1 | 37 | 66 | 13 | 74 | 45
1.5 | 26 | 37 | 5 | 20 | 7
Dose rate substantially higher on X ray focus side of patient compared to Image intensifier side because of scatter from the patient
*Occupational exposure from common fluoroscopic projections used in orthopedic Surgery Nicholas Theocharopoulos et al Journal of Bone and Joint Surgery; Sep 2003; 85, 9;
X ray
Direction
Receptor
X ray
Source
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 30: Effect of Patient Size on Dose
Thicker tissue masses absorb more radiation, thus much more radiation must be used to penetrate the large patient. Risk to skin is greater in larger patients!
15 cm
20 cm
25 cm
30 cm
ESD = 1 unit
ESD = 2-3 units
ESD = 4-6 units
ESD = 8-12 units
Need ~2x more exposure for every 5 cm increase in thickness.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 31: Entrance Dose to Patient vs. Imaging Geometry
Lowest (GOOD) ---------------------------- Highest (BAD)
Image Intensifier close to patient, X ray tube far from patient
Image intensifier far from patient, X ray tube close to patient
From: J American College of Cardiology 2004; 44(11): 2259-82
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>

## Slide 32: Entrance Dose to Patient vs. Imaging Geometry
For the same dose rate at II,
Entrance skin dose for B is (80/40)2 = 4 times higher
Keep the X ray tube as far away from the patient as possible
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 33: Tissue Thickness & Dose Rate
Thicker tissue masses absorb more radiation, thus much more radiation must be used.
Higher dose to patient when imaging through steep projections
Risk to skin is greater with steeper beam angles!
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 34: Factors Affecting Patient Entrance Surface Dose Rates - Grids
Grids
  Grid is placed in front of the image detector
  A grid reduces the effect of scatter (degrading of image contrast), BUT it also attenuates the primary X ray beam (both scatter & primary hit grid strips).
  typically require a 2 times increase in patient dose rate to compensate for attenuation & maintain same X ray intensity at image intensifier as without grid.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 35: Grids in Paediatric Imaging
Small patients produce less scatter
For smaller patients & small body parts (e.g. a hand) adequate imaging may be obtained without grid
Consider removing grid for patients < 20 kg
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 36: Patient doses – collimation
Collimation to square inside image reduces
dose-area product by 36 %
Area of circle = r2
Area of square = 2r2
(r2 - 2r2)/ r2 = 36 %
All else being equal
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 37: Collimation
Why is narrowing the field-of-view beneficial?
Reduces cancer risk to patient by reducing volume of tissue irradiated
Reduces scatter radiation at image receptor to improve image contrast
Reduces ambient radiation exposure to in-room personnel
Reduces potential overlap of fields when beam is reoriented
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 38: A word about collimation
What collimation does not do –
It does NOT reduce dose to the exposed portion of patient’s skin.
Note: dose at the skin entrance site may increase if collimator blades are moved too far into image and X ray machine increases dose to try and “see” through collimator
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 39: A word about collimation
What collimation does not do –
It does NOT reduce dose to the exposed portion of patient’s skin.
Skin dose may actually increase at smaller area collimation if the automatic brightness control trys to compensate for the lower number of X rays incident upon the image receptor; image quality will still improve with smaller collimation as it reduces scatter.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 40: Dose & Dose Area Product (DAP)
Note: Dose is independent of size of area exposed:
Dose = Energy absorbed (E) / Mass
Dose = 2 E / 2 Mass = E / Mass = same dose!
Like rainfall. For example, 10 l/m2 rain in each case. Doesn’t tell you how much water fell - need to know area.
vs.
a)
b)
Dose Area Product (DAP) = dose x area exposed DAPb = 2 x DAPa
A better estimate of overall cancer risk.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>
There is a problem when we try to relate dose to risk. That problem arises because dose is a concentration unit (energy per mass). Therefore it is independent of the area exposed. For example if we collimate an x-ray beam to expose one cube as shown in part a of this slide the dose is the energy (E) divided by the mass of one cube. If we open the collimator to expose 2 adjacent cubes as in part b, we would have twice as many x-rays or energy being absorbed in twice the mass, so the dose would be 2E/2Mass which is equal to E/Mass, which is the same as for 1 cube. This is like rainfall. For example we might have 1” of rainfall in each case. This doesn’t tell us how much rain fell . We need to know the area exposed to the rain or x-rays. Obviously exposing larger areas of tissue increases the cancer risk. To account for this a new unit called the dose area product or (DAP) was developed. It is equal to the dose times the area exposed. The typical units are centi-Gray centimeter squared.. In the example shown in this slide, the DAP in part b is twice the DAP in part a. Many new fluoroscopy units display the cumulative dose area product on the control console.

## Slide 41: Dose Area Product (DAP)
Many new units display DAP
DAP = D x Area
the SI unit of DAP is the Gy.cm2
Area = 1 Dose = 1
Area = 4 Dose = 1/4
d1=1
d2=2
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Notes: Radiation protection for doctors using fluoroscopy
<number>

## Slide 42: A word about collimation
What does collimation do?
Collimation confines the X ray beam to an area of the user’s choice.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 43: Projection Angle & Peak Entrance Surface Dose
Positioning anatomy of concern at the isocenter permits easy reorientation of the C-arm but in this case the image receptor is too far away from the patient’s exit surface. This causes a high skin entrance dose.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Image Receptor
Image Receptor
Image Receptor

## Slide 44: Projection Angle & Peak Entrance Surface Dose
When isocenter technique is employed, move the image intensifier as close to the patient as practicable to limit dose rate at the entrance skin surface.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Image Receptor
Image Receptor
Image Receptor
It is acceptable to have the image receptor in contact with the patient

## Slide 45: Projection Angle & Peak Entrance Surface Dose
Lesson: Reorienting the beam distributes dose to other skin sites and reduces risk to single skin site.
Reproduced with permission from Wagner LK, Houston, TX 2004.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 46: Projection Angle & Peak Entrance Surface Dose
Lesson: Reorienting the beam in small increments may leave area of overlap in beam projections, resulting in large accumulations for overlap area (red area).
Reproduced with permission from Wagner LK, Houston, TX 2004.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 47: Projection Angle & Peak Entrance Surface Dose
Lesson: Reorienting the beam in small increments may leave area of overlap in beam projections, resulting in large accumulations for overlap area (red area).
Good collimation can reduce this effect.
Reproduced with permission from Wagner LK, Houston, TX 2004.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Very small
overlap

## Slide 48: Projection Angle & Peak Entrance Surface Dose
Lesson: Reorienting the beam in small increments may leave area of overlap in beam projections, resulting in large accumulations for overlap area (red area).
Good collimation plus adequate rotation can emilinate this effect.
Reproduced with permission from Wagner LK, Houston, TX 2004.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
No over
overlap

## Slide 49: Projection Angle & Peak Entrance Surface Dose
Conclusion:
Orientation of beam is usually determined and fixed by clinical need.
When practical, reorientation of the beam to a new skin site can lessen risk to skin.
Overlapping areas remaining after reorientation are still at high risk. Good collimation reduces the overlap area.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 50: Dose rate dependence field-of-view or magnification mode
INTENSIFIER
Field-of-view (FOV)
RELATIVE PATIENT ENTRANCE DOSE RATE
FOR SOME UNITS
12" (32 cm) 100
9" (22 cm) 177
6" (16 cm) 400
4.5" (11 cm) 700
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 51: Dose rate dependence field-of-view or magnification mode
How input dose rate changes with different FOVs depends on machine design and must be verified to properly incorporate use into procedures.
A typical rule is to use the least magnification necessary for the procedure, but this does not apply to all machines.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 52: Pulsed Fluoroscopy
Usually, the lower the pulse rate, the lower the dose.
Amount of decrease varies by machine & settings.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 53: Pulsed Fluoroscopy
Usually, the shorter the pulse duration, the lower the dose.
Amount of decrease varies by machine & settings.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 54: Pulsed Fluoroscopy
Example: Modern (2007) R&F system Phantom = Adult Abdomen; 33cm FOV, 0.2 mm Cu filtration
Measured Input Exposure Rate (mR/minute)
Note: ( ) = % decrease relative to 12.5 pps 8 pps / 12.5 pps = (-21%); 3 pps / 12.5 pps = (-76%)
Dose @ 3 pps in Fluoro 3 is almost 50% > Dose @ 12.5 pps in Fluoro 2
pulses/sec | Fluoro 1 | Fluoro 2 | Fluoro 3
12.5 | 320 | 492 | 1041
8 | 199 (-38%) | 396 (-20%) | 1007 (-3%)
3 | 76 (-76%) | 232 (-53%) | 710 (-32%)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 55: Unnecessary body parts in direct radiation field
Wagner – Archer, Minimizing Risks from Fluoroscopic X Rays, 3rd ed, Houston, TX, R. M. Partnership, 2000
Vañó et al, Br. J Radiol 1998, 71, 510-516
Patient was draped for procedure and physicians did not realize that she had moved her arm so that it was resting on the port of the X ray tube during the procedur
Injury to arm of 7-year-old girl after cardiological ablation occurred due to added attenuation of beam by presence of arm and due to close proximity of arm to the source.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 56: kV setting & Patient Dose Rate
Factors that affect patient dose rate
  kVp
  mA
  manual vs auto
  pulsed vs continuous
  last image hold
  boost
  magnification
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 57: Factors Affecting patient entrance surface dose rates
kVp / mA selection
  low kVp / high mA  high patient dose rates
  high kVp / low mA  low patient dose rates, but reduced image contrast
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 58: Design of fluoroscopic equipment for proper radiation control
Fluoroscopic dose output in modern systems is controlled by the equipment. The operator can influence the way the system works by selecting various dose rate modes.
It is not always obvious that a control adjusts the X ray dose rate and may be labeled with “Brightness”, “High Detail”, “Fluoro +”, or similar.
Boost Modes increase the II input dose rates (typically x2), and hence the patient entrance dose rate increases.
Fluoroscopic X ray Output
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 59: Other factors affecting patient dose during fluoroscopy
Screening time
  Last image hold
  Fluoro Store, Snap Shot
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 60: Monitoring doses in complex exams is complex
Exam may involve one or more of:
    Fluoroscopy
    Radiography
    Digital acquisition
During the exam the following varies
    Dose rate
    Beam size
    Beam orientation (PA, Lat., etc)
    Body Part being X rayed
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 61: Monitoring doses in Complex exams – Dose-Area Product Meters
X ray Table
Dose-area product meter
Collimator
Image Intensifier
X ray Tube
2345
Reset
cGy.cm2
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 62: Monitoring doses in Complex exams – Dose-Area Product Meters
X ray Table
Dose-area product meter
Collimator
Image Intensifier
X ray Tube
2345
Reset
cGy.cm2
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
DAP counts all photon
Including those from
Fluoro and Cine runs

## Slide 63: Dose-Area Product Meters
Units Gycm2, cGycm2…
Can be used to compare dose performance with published data
Can be used to estimate skin dose
    Via conversion tables
    Via software within X ray machine (need estimate of field size @ skin)
    Via calculation. Must estimate field size @ skin from imaging geometry (SSD & SID) & collimator size at image intensifier.
  Can be used to set action levels to prevent skin injury, but dose rather than DAP is best for this.
SID
SSD
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 64: Reference doses for X ray procedures
NOT a dose limit
The amount of radiation that, under normal circumstances, one should not need to exceed in performing an X ray procedure on an average size patient.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 65
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
F Mettler et al; Radiology 2008
<number>

## Slide 66: Reference doses for X ray procedures
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Procedure | Mean effective dose (mSv) | Equivalent number of PA chest radiographs (each 0.02 mSv)
Other extremities | 0.001 | 0.05
Knee | 0.005 | 0.25
Shoulder | 0.01 | 0.5
Sternum | 0.01 | 0.5
TM joint | 0.012 | 0.6
Skull | 0.1 | 5
Arthrography | 0.17 | 8.5
Cervical Spine | 0.2 | 10
Lumbosacral joint | 0.34 | 17
Upper extremity angiography | 0.56 | 28
Pelvis | 0.6 | 30
Hip | 0.7 | 35
Thoracic Spine | 1 | 50
Lumbar Spine | 1.5 | 75
Myelography | 2.46 | 123
Lower extremity angiography | 3.5 | 175
Thoracic aortography | 4.1 | 205
Peripheral arteriography | 7.1 | 355
https://rpop.iaea.org/RPOP/RPoP/Content/InformationFor/HealthProfessionals/6_OtherClinicalSpecialities/Orthopedic/index.htm#ref2

## Slide 67: Physical factors and challenges to radiation management
Lesson:
Actions that produce small changes in skin dose accumulation result in important and considerable dose savings, sometimes resulting in the difference between severe and mild skin dose effects or no effect.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 68: New Developments in Dose Reduction
Collimation Without Radiation View Last image hold (LIH) & adjust collimation with graphical overlay on image.
Patient Positioning Without Radiation Position patient via graphical display showing central beam location & edges of field on LIH. (Central beam indicator moves on display as table (patient) is moved).
Automatic Beam Filtration Adds filtration to decrease patient dose based on patient attenuation (e.g. 0.9 mm Cu for small patient, 0.2 mm Cu for large patient.)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 69: Staff radiation protection
Question: Can I work my full professional life with radiation in the operating room and have no radiation effects?
Yes it is possible. Under optimized conditions when
  the equipment is periodically tested and it is operating properly,
  personal protective devices (lead apron of suitable lead equivalence of 0.25 to 0.5 mm and wrap around type, protective eye wear or protective shields are used for the head/face and leg regions),
  use of personnel monitoring
  using the ALARA (as low as reasonably achievable) principle.
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 70: Staff radiation protection
Question: Is the dose to orthopaedic surgeons much higher than other interventionalists?
Answer: No. The radiation dose to orthopaedic and trauma surgeons in most routine procedures is much smaller than those performing cardiac interventions
Approximate dose to the surgeon per procedure (µSv) with 0.5 mm lead apron worn. Exposure from common fluoroscopic projections used in orthopedic Surgery.
The Journal of Bone and Joint Surgery, 85 (2003) 1698-1703
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Procedure | Dose to the Surgeon per procedure (µSv) | Screening Time
Hip | 5 | 25 sec/patient
Spine | 21 | 2 min/patient
Kyphoplasty | 250 | 10 min/patient

## Slide 71: Staff radiation protection
Question: Is there a risk of cataract after several years of work in an orthopaedic operating room?
Very unlikely. Proper use of radiation protection tools and techniques can prevent deterministic effects such as cataract and can avoid any significant increase in probability of cancer risk for many years to cover the full professional life. To date, there have been no reports of radiation induced cataract among orthopaedic surgeons, however such reports do exist for interventional radiologists and cardiologists
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 72: Summary
Keep screening times and acquisitions to a minimum
Use low dose settings as defaults
Keep the X ray tube as far away from the patient as possible
Keep the Image Intensifier close to the patient
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 73
Use magnification mode as little as possible
Collimate when possible
Use last image hold and fluoro storage if available
Remove grid for procedures on small patients
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
Summary

## Slide 74: Summary
Use low pulse rate
Use higher kVp unless it compromises image contrast
Compare procedure fluoroscopy time and dose with published values (reference levels)
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment
https://rpop.iaea.org/RPOP/RPoP/Content/InformationFor/HealthProfessionals/4_InterventionalRadiology/DiagnosticFluoroscopy.htm

## Slide 75: A final general recommendation
Be aware of the radiological protection of your patient and you will also be improving your own occupational protection
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 76: Further readings
ICRP Publication 85. Avoidance of radiation injuries from medical interventional procedures
LK Wagner. Radiation injury is a potentially serious complication to fluoroscopically-guided complex interventions. Biomed Imaging Interv J 2007; 3(2): http://www.biij.org/2007/2/e22/
IAEA http://www.rpop.org Radiation protection of patients
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment

## Slide 77: Thank you
<number>
IAEA Training Course on Radiation Protection for Doctors (non-radiologists, non-cardiologists) using Fluoroscopy
L06A. Anatomy of Fluoroscopy & CT Fluoroscopy Equipment