## Slide 1: Radiation Protection in Paediatric Radiology
Radiation Protection of Children in Digital Radiography
L04
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 2: Educational objectives
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
At the end of the program, the participants
should :
  Become familiar with different digital imaging modalities
  Understand the influence of digital radiology on image quality and patient dose
  Appreciate dose management in paediatric digital radiography
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 3: Answer True or False
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Overexposure is not obvious in digital radiography.
Use of digital detectors will always reduces the dose.
Staff training is necessary in transition from screen-film to digital radiography.
Collimation is not important in digital radiography, as an image can be cropped to what is needed.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 4: Contents
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Overview of digital equipment used in digital radiography
Advantages and disadvantages of digital imaging equipment
Dynamic range, possibility for dose increase
Dose management in digital radiology
Relationship between image quality and patient dose.
Need for additional staff training in paediatric digital radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 5: Introduction
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Screen-film radiography is now being replaced by a variety of digital technologies in many countries
Most principles for dose reduction from paediatric screen-film radiography, espacially justification, are relevant to digital systems (see L03 for more details)
However, in digital systems different scenarios apply for dose reduction and optimisation compared with screen-film radiography
This presentation is focused on essential aspects of digital radiography compared with screen-film radiography, not to digital technologies and equipment
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 6: Introduction
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Paediatric radiography differs from imaging of adults, as children are smaller and less compliant
CR and DR were introduced with a goal of reducing dose and improving image quality
Initially, results were hindered by lack of relevant exposure parameters and understanding of digital technologies
Potential for over-exposure exists in both with CR and DR
Although doses in radiography are low, there is a need for dose reduction and practice improvements while maintaining diagnostic quality
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 7
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Computed Radiography (CR)
Equipment for digital radiography
Direct Radiography (DR)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 8: Equipment for digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 9: Equipment for digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
CR involves an intermediate step in which the image is stored as a latent image, in a cassette-like device, before it is converted to electronic/digital form, using laser stimulation
In DR, the image is created immediately in electronic/digital form, in the image receptor
Images from both systems can be displayed on suitable high-resolution monitors or printed out on film
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 10: Screen-film vs digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Digital images can be numerically modified (not possible in screen-film radiography)
Digital images can be easily transmitted through networks and archived
Attention should be paid to the potential increase of patient dose due to a tendency to:
  produce more images than needed
  produce higher image quality not necessarily required for the clinical purpose
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 11: Screen-film vs digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Conventional films directly show the selection of wrong exposure parameters: images are too white or too black
Digital technology always provides user with a “good image”, since its dynamic range and tone-scaling compensates for wrong settings even if the dose is higher than necessary
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 12: What is “dynamic range”?
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Wide dose range to the detector allows “good” tone scale to be obtained at different dose levels
Digital detectors have a dynamic range of 104 (from 1 to 10,000), while a screen-film system has approximately 101.5 (from 1 to 30)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 13: Screen-film vs digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Digital detector
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 14: Screen-fim vs digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
The key advantage of CR and DR is greatly improved contrast resolution, accompanied by almost infinite possibilities with digital processing
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 15: Digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
In general, digital imaging has potential for dose reduction while improving image quality and diagnostic accuracy
But only with much attention to:
  staff training
  continuous monitoring of parameters and practices
The key issue is that, because of image processing, the tone scale will continue to look good even if the dose increases
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 16: Digital radiography
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
But, with digital systems, dose cannot be reduced indefinitely because increased noise reduces image quality
  Be aware of a tendency among technologists to avoid the possibility of repeats by working on the side of overexposure (the term exposure creep has been coined)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 17: CR, DR and dose management in practice (I)
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Justification is required for CR and DR studies just as it is for screen-film radiography
Positioning, collimation, and selection of optimized exposure factors, is just as essential as for screen-film radiography
A team approach to dose management is essential
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 18: CR, DR and dose management in practice (II)
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Team participants should include: radiologist, medical physicist, radiographer/technologist, clinical engineer from the hospital, service engineer, applications specialist and imaging scientist from the manufacturer
Training of the radiologist and radiographer/technologist in the specific operational features of the CR and DR system in use is essential
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 19: CR, DR and dose management for manufacturers (I)
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Manufacturers must provide adequate end-user training as part of the equipment supply package
Nomenclature for digital imaging processing algorithms and exposure indices must be standardized
It is essential to make provision for passing the exposure/dose information from the radiological equipment to the PACS and/or patient record in an accessible form
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 20: CR, DR and dose management for manufacturers (II)
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Dose measuring devices and dose indicators must be calibrated and must be protected from casual modification by the operator.
Manufacturers and standards organizations must give particular attention to the special issues of paediatric radiology.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 21: Exposure indicators in CR
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Exposure index: term used in relation to the absorbed dose to the phosphor plate (ICRP 93)
Method for dose monitoring
Different manufacturers have different indices
IEC 62494-1:2008 specifies definitions and requirements for the exposure index of images acquired with digital X-ray imaging systems.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 22: Exposure indicators in CR
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Manufacturer | Exposure indicator | Mean receptor air KERMA (Gy) |  | 
 |  | 5 | 10 | 20
Fuji | S | 400 | 200 | 100
Kodak | EI | 1700 | 2000 | 2300
Agfa | lgM | 1.9 | 2.2 | 2.5
Exposure indicator is related to the air KERMA at the detector and the signal produced by the detector, and correlates with entrance skin air KERMA (Vano, et al, Phys Med Biol 2008, 53:3365-80
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 23: CR, DR and optimisation
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Digital systems can be operated at almost any level of receptor dose
Receptor dose needs to be tailored to the imaging task (decision is made by a responsible radiologist)
Example: follow up scoliosis examination in children could be performed at receptor dose below 1 µGy (Kushner et al. Pediat. Radiol.1986; Kogutt et al. Pediat. Radiol.1988)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 24: CR, DR and optimisation
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Training:
It is important that radiology departments prepare well for:
  introduction of digital technology, or
  a new system involving digital technology
Well-trained staff should adopt a team approach in cooperation with the suppliers technical staff, the hospital’s medical physicist, and maintenance staff
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 25: CR, DR and optimisation
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Processing:
Appropriate digital image processing is crucial in producing an optimised paediatric CR or DR image
To date there is little standardization in the methods of digital image processing or their nomenclature
Generic categories of digital image processing:
  Exposure recognition and segmentation (autoscaling)
  Edge restoration
  Contrast enhancement
  Gradation mapping
  (Flynn, Adv. In Dig Rad, RSNA Categorical Course on Diag Rad Phys, 2003 pp 71-78)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 26
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Digital image processing with edge restoration
Original
Processed
CR, DR and optimisation
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 27: CR, DR and optimisation
Role of image processing is to ‘tailor’ the raw digital data, in order to:
  Show everything that may be relevant
  Nothing more (artifacts)
  Nothing less (masking)
Image processing
  Can compensate for under- or over-exposure covering a range of 300-400 times that of screen film systems – reduces repeats
  Training by an experienced trainer in paediatrics is essential and needed on a regular basis
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 28: CR, DR and optimisation
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
In digital systems:
  Wide dynamic range of the detector allows overexposure to be tolerated
  Overexposed images have less noise, may present motion artefacts due to long exposure time
  Underexposure produces quantum mottle
  Repeat exposures by technologists not being monitored or reported – bad electronic images are deleted
  Convenience of taking images leads to covering a larger area of a patient's body and cropping or repeating the examination
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 29
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Poor collimation
Large part of the body is irradiated
Not seen on digitally cropped image
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 30: CR, DR and optimisation
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Dose parameter monitoring is vital to prevent unnecessary exposure and for optimisation:
  Exposure indicators
  Exposure indicator values vary with technical exposure factors (kVp, mAs, SID)
  Use of Diagnostic Reference Levels
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 31: CR, DR and exposure indicators
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 32: Typical dose levels in CR
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Vano et al, Paediatric entrance doses from exposure index in computed radiography, Phys Med Biol 53 (2008):3365-3380
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 33: CR, DR and optimisation
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Operational methods for controlling dose in paediatric radiology:
  Avoid use of grid
  Increase source-to-skin distance
  Use manual setting of exposure parameters
  Reprocess, rather than repeat non-diagnostic images
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 34: CR, DR and optimisation
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Viewing environment:
Inadequate viewing environment (ambient light) has an impact on diagnostic information content
Spatial requirements for monitors (contrast, resolution)
Routine Quality Control
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 35: Summary (I)
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Digital systems have distinct advantages over screen-film systems in terms of digital image availability, flexibility, and operational speed
With proper training and quality control, dose reduction and image quality improvement are possible
Users should become aware of the possibility of unnecessarily or inadvertently high doses in all digital systems
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 36: Summary (II)
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Practice in a department should be audited periodically with respect to an index for patient dose
Unless system is fully optimised, dose reduction is not achieved and overexposure is not detected
Training on dose management for particular digital system is needed during transition from screen-film and at regular intervals later
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 37: Answer True or False
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
Overexposure is not obvious in digital radiography.
Use of digital detectors will always reduces the dose.
Staff training is necessary in transition from screen-film to digital radiography.
Collimation is not important in digital radiography, as an image can be cropped to what is needed.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 38: Answer True or False
True - Due to wide dynamic range of digital detectors, nice images can be obtained at double the dose or more and overexposures are not obvious unlike dark images in screen-film radiography.
False – Only when features are utilized, not always. See the earlier question as an example.
True – Additional consideration of image quality and dose necessitate additional training.
False - Collimation is an important tool in dose reduction as cropping image after exposure has been given does not help in patient dose.
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography

## Slide 39: References
Radiation Protection in Paediatric Radiology L04. Radiation protection in digital radiography
<number>
WILLIS, C, Optimizing digital radiography for children, Eur J Radiol2009, 72: 266-273
COMPAGNONE, G., et. Al., Comparison of radiation doses to patients undergoing standard radiographic examinations with conventional film screen radiography, computed radiography and direct digital radiography Br. J. Radiol. 79 (2006) 899-904.
ONO, K., et. Al. Comparison of a digital flat-panel versus screen-film, photofluorography and storage-phosphor systems by detection of simulated lung adenocarcinoma lesions using hard copy images. Br. J. Radiol. 78 (2005) 922-927.
HUFTON, A.P., DOYLE, S.M., CARTY, H.M., Digital radiography in paediatrics: radiation dose considerations and magnitude of possible dose reduction. Br. J. Radiol. 71 (1998) 186-199.
RAPP-BERNARDT, U., et. al., Experimental evaluation of a portable indirect flat panel detector for the pediatric chest: comparison with storage phosphor radiography at different exposures using a chest phantom, Radiology 237 (2005a) 485-491.
RAPP-BERNARDT, U., et. Al.,Portable flat-panel detector for low-dose imaging in a pediatric intensice care unit, Invest. Radiol. 40 (2005b) 736-741.
VANO, E., et al., Radiation dose and image quality for paediatric interventional cardiology, Paediatric entrance doses from exposure index in computed radiography, Phys Med Biol 53 (2008):3365-3380
INTERNATIONAL COMISSION ON RADIOLOGICAL PROTECTION, Radiation protection in digital radiology, Annals of the ICRP, Publication 93, Pergamon Press, Oxford (2000).
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>