## Slide 1: Radiation Protection in Paediatric Radiology
Radiation Protection of Children in Interventional Radiology and Cardiology
L07
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 2: Educational objectives
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
At the end of the programme, the participants should know these
  What are specific considerations for paediatric patients in interventional radiology and cardiology?
  How can dose be managed in paediatric patients?
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 3: Answer True or False
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Radiation dose to the patients can only be measured by a specialized person standing in the catheterization laboratory during the procedure.
Skin injuries are possible in interventional procedures.
Detector should be as close as possible to the patient.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 4: Contents
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Interventional procedures performed on children
Review of epidemiology of radiation effects
Typical doses to patients in paediatric interventional radiology and cardiology
Requirement for optimisation and tailoring of radiological technique to small body sizes
Measures to protect patients and staff in paediatric interventional radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 5: Introduction
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Paediatric interventional radiology (IR) is a medical field that specializes in minimally invasive diagnostic or therapeutic procedures using imaging guidance, in children:
  IR treatments can solve problems faster and with smaller incisions than can other techniques
  They can help treat medical problems which cannot be solved any other way
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 6: Some interventional procedures in pediatric radiology
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Nephrostomy
Fluid collection drainage or aspiration
Biliary stent placement
Gastrostomy or cecostomy
Angiography
Angioplasty or stent placement
Embolization, thrombolysis or sclerotherapy
Radiofrequency ablation
Biopsy – generally carried out using
ultrasound, fluorography or CT
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 7: Some interventional procedures in pediatric cardiology
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
  Balloon dilatation / stenting
    vascular stenoses
      aortic coarctation
    valvular obstructive lesions
      pulmonary stenosis
      mitral stenosis
  Transcatheter closure
    atrial septal defects (ASD)
    ventricular septal defect (VSD)
    patent ductus arteriosus (PDA)
  Electrophysiology
    ablation
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 8: Unique Considerations for Radiation Exposure in Children
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Compared to a 40-year old, a neonate is several times more likely to produce a cancer over the child's lifetime, when exposed to the same radiation dose
  Interventional procedures in children must be carefully justified
  Radiation doses used to examine young children must generally be smaller than those employed in adults
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 9: Possible radiation effects
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Stochastic effects (induction of malignant or hereditary diseases)
  Measure: effective dose
Tissue reactions (deterministic effects) is skin injury following interventional procedures
  Measure: maximum skin dose or cumulative dose to interventional reference point
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 10
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Arm of 7-year-old patient after cardiological ablation procedure. Injury to arm occurred due to added attenuation of beam by presence of arm and due to close proximity of arm to the source.
Vano E, et al. Br J Radiol 1998; 71(845):510-6
Possible radiation effects
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 11: Justification and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Interventional procedures in children are now more in demand, more sophisticated and longer
Should always be individually justified and planned
Balance the risk against the therapeutic benefit
It is important to ask the referring practitioner, the patient, and/or the family about previous procedures
Determination that the procedure is necessary relies on the radiological practitioner involved, the natural history of the untreated disease, and the risks and benefits of other therapeutic options available
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 12: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Procedures should be pre planned to minimize improper or unneeded runs:
  the number and timing of acquisitions, contrast parameters, patient positioning, suspension of respiration, sedation…
The acquisition parameters should be selected to achieve lowest dose necessary to accomplish procedure, taking account:
  dose protocol
  patient size
  frame rate
  magnification
  length of run
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 13: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Image acquisition using cinefluorography, or during digital subtraction angiography (DSA), accounts for the largest radiation doses during many interventional procedures (dose rates involved can be up few of orders of magnitude higher than fluoroscopy)
Exposure factors for cinefluorographic image acquisitions and quasi-cine runs are much higher than those for fluoroscopy
The acquisition mode should be carefully selected
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 14: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Children are not just small adults:
  imaging equipment needs to be specifically designed for use with children and the operators must be trained accordingly
With infants and small children the image intensifier will completely cover the patient: accuracy of collimation is of much greater importance than with adults
After procedure the dose records should be noted and reviewed
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 15: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Steps in optimisation depend on patient size, technical challenge and critical nature of the procedure
Overall patient safety is the most important
The goal is to minimize the dose to the patient while providing important and necessary medical care
Equipment
  Requires careful selection, maintenance and Quality Assurance
  “Child size” protocol with dose reduction options available
  Dose recording system
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 16: X-ray equipment for pediatric IR
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
The generator should have enough power to allow short exposure times (3 milliseconds).
Fluoroscopic pulsing X-rays are produced during a small portion of the video frame time. The narrower the pulse width, the sharper the image. ( “Shutter speed” in camera )
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 17: X-ray equipment for pediatric IR
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
The generator should be of high frequency (i.e can produce higher pulsed fluoroscopy) to improve the accuracy and reproducibility of exposures.
  E.g. children have faster heart rate. Coronary angiography in children is often acquired at 25-30 frames/sec, instead of the usual 12.5 – 15 frames/sec for adult patients.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 18: X-ray equipment for pediatric IR
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
The generator must provide a large dynamic range of mA and mAs levels (to minimize the range of kVp and exposure time needed to compensate for differences in thickness)
Automatic exposure control (AEC) devices should be used with caution in pediatrics
Careful manual selection of exposure factors usually results in lower doses
Tree focal spots should be available
Also: lateral imaging plane, spatial and spectral beam profiling, and a well functioning system of entrance dose regulation
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 19: X-ray equipment for pediatric IR
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Image Handling
and Display
Image Receptor
X -ray tube
High-voltage transformer
Power Controller
Primary Controls
Operator Controls
Patients
Operator
Foot
Switch
Electrical
Stabilizer
Automatic
Dose Rate
Control
  Image intensifier should have high conversion factor to reduce patient dose
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 20: Beam filtration
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
The introduction of additional filtration in the X-ray beam (commonly copper filters) reduces the number of low energy photons and, as a consequence, saves skin dose for the patients.
Additional Cu filters can reduce the skin dose by more than 70%.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 21: Anti-scatter grid
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
The anti-scatter grid in pediatrics gives limited improvement in image quality and increases patient dose given the smaller irradiated volume (and mass) the scattered radiation is less
Increase kerma air product (KAP) and skin dose typically by ≥ 2 times
Does NOT improve image quality
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 22
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Antiscatter grid
The anti-scatter grid in pediatrics gives limited improvement in image quality and increases patient dose and should be removed
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 23: Wedge filter
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Partially absorbent contoured filters are also available to control the bright spots produced by the lung tissue bordering the heart.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 24: Importance of wedge filters
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
The wedge filter has not been used to obtain this cine series. Note the important difference in contrast.
The wedge filter has been used to obtain this cine series.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 25: New modalities
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Electronic road mapping is a
must as it greatly reduces the
risk of dissection during
catheterisation of complex, narrow
or irregular vascular channels.
Digital subtraction angiography
(DSA) is another must as it
permits a great reduction in the
concentration of contrast medium
used, thus reduces the risk of
toxicity to the kidneys and spinal cord
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 26: Relevant dosimetric quantities
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
For assessment of stochastic risk:
  Kerma-area product (KAP, PKA)
For prevention of tissue reactions (deterministic effects)
  Maximum skin dose (MSD)
  Cumulative dose (CD) to Interventional Reference Point (IRP)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 27
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
X-ray room dosimetric information
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 28: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
2. Procedure
Communication between in room personnel
Plan in advance: plan number of runs, injection parameters, contrast, pump, digital subtraction angiography (DSA) frame rates and optimize patient position timing with anesthesia
Lower the number of exposures: use flouro save when possible
Use last image hold, decreasing acquisitions/exposures as much as possible when that level of detail is acceptable
Step lightly: tap on pedal and examine still image on monitor, minimize live fluoroscopic time
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 29: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
2. Procedure
Use pulse fluoroscopy when possible
  Decrease from 7.5 pulses/s to 3 pulses/s whenever possible
Collimate tightly
  Decreasing the area of patient exposure directly decreases patient dose
  Avoid dose to the eyes, thyroid and gonads whenever possible
Minimize overlap of fields in repeated acquisitions
Decrease the dose rate setting to the lowest level that provides adequate image quality
Minimize use of electronic magnification
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 30: Pulsed Fluoroscopy
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Pulsed fluoroscopy can be used as a method of reducing radiation dose, particularly when the pulse rate is reduced.
…but … pulsed fluoroscopy does not mean that dose rate is lower in comparison with continuous fluoroscopy!!.
Dose rate depends on the dose per pulse and the number of pulses per second.
High frequency
Low frequency
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 31: Collimation
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Dual-shape collimators incorporating both circular and elliptical shutters may be used to modify the field for cardiac contour collimation
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 32: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
2. Procedure
Maximize distance between source and patient throughout the procedure.
Minimize patient to detector distance
Avoid radiosensitive areas (breast, eyes, thyroid, gonads) when possible
Audible periodic fluoroscopy time alerts
Image acquisition limited only to needed (frames per second, lower dose protocols, magnification, length of run)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 33: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
3. After procedure
Review dose
Counsel if :
  Skin dose greater than or equal to 2 Gy, or
  Cumulative dose of greater than or equal to 3 Gy
Follow up:
  Notes to primary care physician about procedure, dose and possible short and long term effects
  Counsel patient and primary care to call if erythema develops at beam entrance site
  Establish follow up procedures including skin examination at 30 days
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 34: Optimisation and interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
4. Training
All persons directing and conducting interventional procedures, including radiologists and technologists, should have education and training in their discipline, radiation protection physics, radiation biology, and specialist training in its paediatric aspects
Specific training in paediatric interventional radiology improves the use of safety measures
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 35: Typical dose levels in paediatric interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Procedure | Number | Mean KAP per unit of body mass (Gy cm2 kg-1) | Effective Dose (mSv)
ASD Occlusion | 259 | 0.42 | 3.9
PDA Occlusion | 165 | 0.35 | 3.2
Balloon Dilatation | 122 | 0.48 | 4.4
Coil Embolisation | 33 | 0.50 | 4.6
VSD Occlusion | 32 | 1.3 | 12
Atrial Septostomy | 25 | 0.39 | 3.6
PFO Occlusion | 21 | 0.23 | 2.2
ASD, PDA, VSD, PFO are, respectively, atria septal defect, patent ductus, ventricular septal defect and patent foramen ovale., Onnash et al, Br. J. Radiol. 80 (2007) 177-85
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 36: Typical dose levels in paediatric interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Cumulative skin dose is well correlated with patient size and not with fluoroscopy time
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 37: Typical dose levels in paediatric interventional procedures
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Comparison of surface entrance doses of radiation A: Amplatzer et al.(atrial septal defect closure) B: Moore et al. (patent ductus coil occlusion) C: Moore et al. (pulmonary valvuloplasty) D: Wu et al. (pulmonary valvuloplasty) E: Park et al. (arhythmia ablation) F: Rosenthal et al. (arhythmia ablation)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 38: Staff doses in interventional radiology and cardiology
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
All team members should become aware of the radiation exposure issues with interventional procedures, and the means of controlling them
Most staff dose, in practice, arises from scattered radiation:
  Regime/protocol (in digital fluoroscopy, cine, digital cine like, or DSA runs, the scattered dose to staff can be several orders of magnitude larger than during fluoroscopy)
  Size of the patient
  Complexity of the procedure
  Training and experience
For a given set up, both patient and staff doses are dependent
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 39: Staff doses in interventional radiology and cardiology
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
The dominant direction for scatter is from the patient back toward the X-ray tube
The operator should be on the image receptor side and step back during injections
90º LAO 150 cm
90º LAO 100 cm
60º LAO 100 cm
30º RAO 100 cm
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 40: Reducing staff doses in interventional radiology and cardiology
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Only those necessary for conduct of the procedure should be in the room
Move personnel away from table, preferably behind protective shields during acquisitions
The operator should stand to the side of the image intensifier.
The operator should use a power injector when possible and step back from the image intensifier and/or behind a mobile lead screen during contrast injections
If manual injection is necessary, maximize the distance using a long catheter
Doses in the room and from undercouch tubes can be greatly reduced by well configured and properly used table side drapes
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 41: Reducing staff doses in interventional radiology and cardiology
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Use movable overhead shields for face and neck protection. Position these prior to procedure.
Well designed suspended shielding/viewing systems are helpful to operators who learn to become skilful in their use.
Wear well fitted, appropriate weight, protective aprons
Wear a thyroid collar and/or lead glasses with side shielding
The operator and personnel should keep their hands out of beam if possible and not between tube and patient
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 42
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
A freely movable lead glass or acrylic shield suspended from the ceiling should be used.
Its sterility may be maintained by using disposable plastic covers.
Reducing staff doses in interventional radiology and cardiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 43: Reducing staff doses in interventional radiology and cardiology
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Radioprotective gloves may be worn where appropriate, but note they can be counterproductive, reduce flexibility/dexterity and/or interfere with the AEC
Slight angulation of the beam off the hands, strict collimation and careful attention to finger positioning will help reduce operator exposure
Occupational dose measurements should include at least one badge under the lead apron to assess whole body dose
Additional monitors over the apron to evaluate thyroid, hand/arm and eye doses are advisable in some situations
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 44
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 45
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Be aware of the radiological protection of your patient and you will also be improving your own occupational protection
General recommendation:
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 46
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
http://rpop.iaea.org/RPoP/RPoP/Content/index.htm
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 47
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
http://www.pedrad.org/associations/5364/ig/
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 48: Summary
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Increased radiation risks for pediatric patients
Trend of increasing number of pediatric interventional procedures
Radiation doses can be high
Radiological technique must be optimized and tailored to small body sizes
Operators shall be trained, as...
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 49: Summary
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
correct indications
fluoroscopy time reduction
frame rate reduction
collimation/filtering
distance from X-ray source
/ image receptor
protective organ shielding
e.g gonad, thyroid
lead apron and thyroid protection
protective glasses and suspended screen
(staff)
(patient)
....patients and staff share a lot...
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 50: Answer True or False
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Radiation dose to the patients can only be measured by a specialized person standing in the catheterization laboratory during the procedure.
Skin injuries are possible in interventional procedures.
Detector should be as close as possible to the patient.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 51: Answer True or False
False - There are dose indices available on the monitor in modern machine such as DAP and cumulative air kerma at interventional reference point.
True – Peak skin doses in the range of few gray exceeding the threshold of 2 Gy are possible.
True - Detector should be as close as possible and X ray tube as far as possible form the patient. This improves image quality and reduces patient dose.

## Slide 52: References
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Axelsson B, et al., Estimating the effective dose to children undergoing heart examinations- a phantom study, BJR 72(1999), 378-383
Balter S, et al. ., Interventional Fluoroscopy: Physics, Technology, Safety, Wiley, John & Sons, (2001).
Cardella JF, Miller DL, Cole PE, et al (2003) Society of Interventional Radiology position statement on radiation safety. J Vasc Interv Radiol 14 :S387.
Hayashi, N., et.al., Radiation exposure to interventional radiologists during manual-injection digital subtraction angiography, Cardiovasc. Intervent. Radiol. 21 (1998) 240-243
Hiorns MP, Saini A, Marsden PJ, A review of current local dose-area product levels for paediatric fluoroscopy in a tertiary referral centre compared with national standards. Why are they different?, BJR, 79 (2006), 326-330
International Commission on Radiation Protection, Avoidance of Radiation Injuries from Medical Interventional Procedures, Publication 85, Elsevier, Oxford and New York
National Cancer Institute and Society of Interventional Radiology (2005) Interventional fluoroscopy; reducing radiation risks for patients and staff. NIH publication no. 05-5286.
Onnash, et al., Diagnostic reference levels and effective dose in paediatric cardiac catheterization, Br. J. Radiol. 80 (2007) 177-85.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 53: References
Sidhu, MK, Goske MJ, Coley BJ, et al. (2009) Image Gently, Step Lightly: Increasing Radiation Dose Awareness in Pediatric Interventions through an International Social Marketing Campaign. Journal of Vascular and Interventional Radiology September 2009 20(9):1115-1119
Sidhu, MK, Strauss KJ, Connolly B, et al. (2010) Radiation Safety in Pediatric Interventional Radiology. Techniques in Vascular and Interventional Radiology (in press)
Vano E, et al. Dosimetric and radiation protection considerations based on some cases of patient skin injuries in interventional cardiology. Br J Radiol 1998; 71(845):510-6
Vano, E., et.al., Staff radiation doses in interventional cardiology, correlation with aptient exposure, Pediatr. Cardiol. 30 (2009) 409-413.(2001).
Racadio JM, Connoly B. Image Gently, Step Lightly: Practice of ALARA in Pediatric Interventional Radiology. Access at: www.imagegently.org or http://spr.affiniscape.com/displayemailforms.cfm?emailformnbr=120543
Radiation Protection in Paediatric Radiology L07. Radiation protection in interventional radiology
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>