## Slide 1: Radiation Protection in Paediatric Radiology
Radiation Protection of Children in Fluoroscopy
L05
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>

## Slide 2: Educational Objectives
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
At the end of the programme, the participants
should:
  To become familiar with the application of practical radiation protection principles to fluoroscopy systems in paediatric radiology
  To appreciate that good radiation protection policy and skilled personnel are essential for patient and staff doses
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 3: Answer True or False
Pulsed fluoroscopy reduces dose.
It is necessary to use the antiscatter grid in every paediatric radiology examination.
Magnification should be always used in paediatric fluoroscopy, because of the small size of the patient.
Use large radiation fields not to miss anything.

## Slide 4: Content
Components of fluoroscopy systems
General recommendations for radiation protection in fluoroscopy
Justification in paediatric fluoroscopy
Optimisation in paediatric fluoroscopy
Operational and equipment consideration
Occupational radiation protection consideration in paediatric fluoroscopy

## Slide 5: Introduction
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Children have higher radiation sensitivity than adults and have a longer life expectancy
A pediatric radiological procedure should be planned and limited to what is absolutely necessary for diagnosis
Radiologists and radiographers should be specifically trained and the higher radio-sensitivity of the patients should be taken into account

## Slide 6: Introduction
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Fluoroscopic procedures may be classified
into:
  Conventional, long established investigations (micturating cystograms and gastrointestinal contrast studies) – treated in Part 5
  Newer interventional and more sophisticated diagnostic procedures – treated in Part 7

## Slide 7: General Recommendations
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Key areas in radiation protection in paediatric
fluoroscopy:
  Justification
  Optimisation
  Evaluation of patient dose and image quality
“Do you really need a glossy picture to
make that diagnosis”

## Slide 8: Components of Fluoroscopy Systems
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Under Couch System
Over Couch System
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 9: Components of Fluoroscopy Systems
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Conventional
Digital
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 10: To obtain the images …
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Two technologies are commonly used:
  Image intensifier
  Flat panel detector

## Slide 11
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Motorized Iris
Video Camera
Image Intensifier
DETECTOR
Photons
Cesium Iodide (CsI)
Light
Amorphous Silicon Panel
(Photodiode/Transistor Array)
Digital Data
Electrons
Read Out Electronics
Photons
Cesium Iodide (CsI)
Light
Photo-cathode
Video Signal
Electrons
Output screen
Light
CCD or PUT
Electrons
Readout Electronics
1
3,000
400
400,000
2,400
Particles #
Image
Intensifier
Flat-panel

## Slide 12: Components of Fluoroscopy Systems
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Fluorescent screen
Photocathode
Amplification
TV Camera
Display
Electrons
X-Rays
TV Camera
Fluorescent Screen
Photocathode
Light Photons
Video
Display
Light Photons
Ouput
phosphor
Image Intensifier
Electrons
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
19

## Slide 13: Automatic Brightness Control (ABC)
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
ABC devices determine the amount of radiation to be incident on the patient based on a feedback mechanism from the amount of light at the output of the image intensifier – which signals back to the generator
This may decrease or increase the incident radiation and radiation dose through the feedback system
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 14: Justification and Conventional Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Justification is required for fluoroscopy studies
Ask referring practitioner, patient, and/or family about previous procedures
Use referral guidelines where appropriate
Use alternative approaches, such as ultrasound, MRI where appropriate
Consent, implied or explicit is required for justification
Include justification in clinical audit

## Slide 15: Justification in Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Referral guidelines for radiological examinations:
  EUROPEAN COMMISSION, Referral Guidelines for Imaging, Luxembourg, Radiation Protection 118, Office for Official Publications of the European Communities, Luxembourg (2001) and Update (2008)
  THE ROYAL COLLEGE OF RADIOLOGISTS, Making the Best use of Clinical Radiology Services (MBUR), 6th edition, RCR, London (2007)
  AMERICAN COLLEGE OF RADIOLOGY (ACR) Guidelines and Appropriateness Criteria

## Slide 16: Examples of Fluoroscopy Examinations not Routinely Indicated
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Upper GI contrast studies of pyloric stenosis
Contrast enema in a child with rectal bleeding.

## Slide 17: Can low dose fluoroscopic image replace conventional radiographic examinations?
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
An image recorded on film with a high-speed cassette provides image detail
However, when high image detail is not required, for example in demonstrating esophageal distensibility, the course of the duodenum, or the progress of contrast in an enema, a stored pulsed fluoroscopic image using last-image-hold is usually diagnostic.

## Slide 18: Optimisation in Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Once exposures are justified, they must be optimised
Number of measures contributes systematic dose savings
Sustainment of good practice through a quality assurance and constancy checking programme
Selection of equipment is important, but good radiography technique is the main factor in improving quality without increasing dose

## Slide 19: Optimisation in Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Once exposures are justified, they must be optimised
A number of measures contribute to systematic dose savings
Sustainment of good practice through a quality assurance and constancy checking programme
Selection of equipment is important, but good radiographic technique is the main factor in improving quality without increasing dose

## Slide 20: Practical Optimisation Measures in Fluoroscopy (I)
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Positioning, collimation, selection of optimised exposure factors are essential in fluoroscopy.
“Child Size” the protocol and use lowest dose protocol possible for patient size, frame rate, and length of run.

## Slide 21: Practical Optimisation Measures in Fluoroscopy (II)
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
The image intensifier/receptor should be positioned over the area of interest before fluoroscopy is commenced rather than positioned during fluoroscopy.
Fields should be tightly aligned to area of interest using the light beam and your eyes rather than fluoroscopy.
Tap fluoroscopy switch and confirm position by reviewing the still Image Hold on the monitor.

## Slide 22: Practical Optimisation Measures in Fluoroscopy (III)
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Field overlap in different runs should be minimized.
Exclude eyes, thyroid, breast, gonads when possible.
Minimize use of electronic magnification, use digital zoom whenever possible.
A low attenuation carbon fibre table should be used where possible.
A removable grid should be available, and normally only used with children > 8 years.

## Slide 23: Practical Optimisation Measures in Fluoroscopy (IV)
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Added copper filtration (eg., 0.3 mm) should be used, and can be left permanently in place if the equipment is deployed solely for children.
Pulsed fluoroscopy should be available and used where possible. Many workers recommend 3.5-7.5 pulses/s as adequate for guidance/monitoring of most procedures.
Static fluoroscopic or fluorographic images, or the last image hold facility should be used to review anatomy/findings.

## Slide 24: Practical Optimisation Measures in Fluoroscopy (V)
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Acknowledge fluoroscopy timing alerts during procedure.
A calibrated KAP/rate meter should be available and used effectively
Record and review dose

## Slide 25: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Fluoroscopic systems can deliver a wide range of radiation doses to patients
This provides a large scope for dose reduction
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 26: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Patient positioning and immobilisation:
  A comfortable, relaxed child is far more likely to co-operate ( higher quality images and less screening time)
  Use of sponges, sandbags, blankets or other simple restraining devices, with the help of attendants, is helpful
  Well trained & experienced staff is invaluable in persuading children to take oral contrast medium

## Slide 27: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
  2. Collimation:
  An over couch system allows use the Light Beam Diaphragm (LBD) to position the patient (Cook JV, Imagining 13:229-38, 2001)
  Prevents use of fluoroscopy for positioning
  Collimation should be to the region of interest
  Too tight collimation should be avoided if the equipment has an unregulated ABC, as this will result in glared, overcontrasted images and unnecessarily high doses.

## Slide 28: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
  3. Focus-to-Skin Distance
  The patient should be positioned as close as possible to the image intensifier
  The X-ray tube should be as far away as possible from the patient’s table in order to avoid excessive skin dose
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 29
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Bad practice
  The image intensifier/detector should be placed as close to the patient as possible
  (< 5 cm) for better image quality and reduced dose (undercoach systems)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 30: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
4. Anti-scatter Grid
Anti-scatter grid should be removable in pediatric equipment, particularly fluoroscopic systems
No grid is recommended for small children resulting in a dose reduction up to 50%
5. Magnification
Magnification should be avoided unless necessary
  Using a field of view of less than 12 cm may result in four times the dose of a 25 cm diameter field.
Digital radiography allows post-processing magnification with no increase in dose

## Slide 31: Magnification
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Changing from a large field of view to an increased magnification increases the exposure required by the image intensifier tube
The absorbed dose to tissues within the beam is also increased
Example:
Field of view, diameter 25 cm
Dose rate= 0.3 mGy/s
Field of view, diameter 17 cm
Dose rate = 0.6 mGy/s
Field or view, diameter 12 cm
Dose rate = 1.23 mGy/s.

## Slide 32: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
6. Exposure factors
Low tube potential (50–60 kV) fluoroscopy provides better demonstration of low to moderate contrast examinations,
e.g. those with iodinated contrast medium (200–300 mmol) or dilute barium (100 mg %)
In combination with heavy tube filtration (0.25 mm copper), can improve quality and reduce dose
Tube current and beam on time are directly proportional to dose
Acknowledge fluoroscopy timing alerts during procedure
Tapiovaara MJ et.al., Phys Med Biol. 1999 44(2):537-59

## Slide 33: Equipment, practice, dose and image quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
7. Filtration
Additional tube filtration may allow dose reductions
0.1mm Cu should be incorporated into all modern systems used in a paediatric setting
Dose reduction by 20% without affecting image quality

## Slide 34: Additional Filtration
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Double-contrast colon:
Added 0.3 mm Cu reduced effective dose with 40-45% at tube voltage 100 kV
No significant detoriation of image quality
B Hansson, et. al. . Eur Radiol 7 (1997) 1117-1122
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 35: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
7. Automatic Brightness Control
  Specific kV/mA dose rate curves for automatic brightness control (ABC) should be used in fluoroscopic systems for children
  An ABC giving a pre-set controlled tube potential (kV) and variable tube current (mA) and allowing ‘‘dose hold’’ is preferred

## Slide 36: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
8. Pulsed Fluoroscopy
  All new equipment should have pulsed fluoroscopy
  Variable pulse rates are possible
  Grid controlled pulsed fluoroscopy X-ray tubes: allows very short exposures with very little prior or trailing components of relatively soft radiation

## Slide 37: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
8. Pulsed Fluoroscopy
  The lowest pulse rate will usually produce the lowest dose, depending on whether there is a compensatory increase in tube current (mA) to maintain quality
  In some systems pulse width is increased on low pulse rates and thus dose reduction is not as substantial
  Take care how system works

## Slide 38: Pulsed Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Pulse length (5-20 ms) for adults reduced to 2-10 ms for children
Pulsed fluoroscopy, as low as 3 frames/sec, allows significant patient dose reduction
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 39: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
9. Frame Grab Technique
  Image is taken directly off the image intensifier during screening and does not incur any additional dose
  Appropriate use of this technique, where detail is not diagnostically needed, is recommended

## Slide 40: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
9. Grid Controlled Fluoroscopy (GCF)
  controls the output within the X-ray tube itself
  eliminate the unnecessary soft radiation emitted by ramping and trailing components
  allow the fluoroscopy parameters (kV, mA and ms) to be adjusted within the duration of a single pulse
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 41: Grid Controlled Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Continuous fluoroscopy
Grid Controlled Fluoroscopy
  Brown PJ, Johnson LM Silberberg PJ, Thomas RD, Low dose, high quality pediatric fluoroscopy, Medica Mundi 45/1 March 2001
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 42: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
10. Shielding
Use lead gonad protection whenever possible
Repeating an examination due to overuse of shielding is poor practice
Lead apron, into beam path – reduce light output from the II – System will automatically increase amount of X-rays to achieve same light output as before!! => Increase patient dose - BEWARE
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 43: Shileding
Carefully collimate the X-ray beam to area of interest excluding other regions, especially gonads, breast, thyroid and eyes.
Use appropriate gonad, thyroid, ovary and breast shielding
10 mGy breast dose to a girl, received before 35 years of age, will increase the spontaneous breast cancer rate by 14% (Brenner DJ et al 2001, Fricke BL et al 2003, Hopper KD et al 1997)

## Slide 44: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
11. Other advantages of modern systems
Use low frame rate
  Reducing the frame rate from 15 f/s down to 3f/s reduces dose by a factor of 5
Use of last image hold and digital spot imaging reduces dose by 20-50%
The cine playback (digital) and video playback (digital/conventional fluoroscopy) may allow patient dose reductions
  System automatically saves the last cine loop in memory

## Slide 45: Equipment, Practice, Dose and Image Quality
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Default characteristic curve on fluoroscopic systems is the adult curve and is usually set at 15 f/s
Ensure application specialist sets system up correctly for paediatric imaging and that radiology staff do not become accustomed to a higher frame rate – this is unnecessary radiation exposure, and may result in blurring of rapidly moving objects
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 46
Patient dose management:
Patient dose records
  After procedure the dose records should be noted and reviewed
Modern methods of patient dose management
  A calibrated DAP/KAP meter
  Real time point dosimeters (MOSFET)
Equipment, Practice, Dose and Image Quality
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology
<number>
All one bullet

## Slide 47: Mobile Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Mobile fluoroscopy is valuable on occasions when it is impossible for the patient to come to the radiology department
It can result in
  poorer quality images
  give rise to unnecessary staff and patient exposures
Where practicable, X-ray examinations should be carried out with fixed units in an imaging department

## Slide 48: Mobile Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Mobile C-Arms should have the option of removing the anti-scatter grid
Calibrated KAP/rate meter should be used
Particular attention should be given to collimation, fluoroscopic time and displayed KAP measurements
Trained staff should operate C-arms especially in pressure environments such as theatres
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 49: Typical Dose Levels in Paediatric Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Examination | Number of studies | DAP range (cGycm2) | Mean DAP (cGycm2)
Upper GI series |  |  | 
0-1 | 235 | 0.06-41.9 | 6.4
1-7 | 376 | 0.09-117.6 | 9.5
8+ | 197 | 0.18-203.5 | 24.7
Dysphagia swallow |  |  | 
0-1 | 116 | 0.3-39.2 | 12.4
1-7 | 246 | 0.1-80.6 | 13
8+ | 84 | 1.3-76.8 | 18.9
  Hiorns MP, et al BJR, 79 (2006), 326-330
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 50: Typical Dose Levels in Paediatric Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Examination | Number of studies | DAP range (cGycm2) | Mean DAP (cGycm2)
Micturating Cystourethrography |  |  | 
0-1 | 165 | 0.02-41.9 | 3.8
1-7 | 94 | 0.7-48.8 | 8.9
8+ | 36 | 4-645.9 | 44.2
Patalal screening |  |  | 
0-1 | 15 | 1.9-5.3 | 4.4
1-7 | 118 | 0.5-31.5 | 5.5
8+ | 146 | 1.3-27.2 | 6.9
  Hiorns MP, et al BJR, 79 (2006), 326-330
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 51: Typical Dose Levels in Paediatric Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Examination | Number of studies | DAP range (cGycm2) | Mean DAP (cGycm2)
Barium follow through |  |  | 
0-1 | 18 | 0.2-73.9 | 14.7
1-7 | 85 | 0.1-42.9 | 9.9
8+ | 101 | 0.2-241.2 | 31.7
Contrast enema |  |  | 
0-1 | 59 | 0.1-42.9 | 5
1-7 | 25 | 0.7-50.5 | 10
8+ | 2 | 6.3-178.6 | 92.5
  Hiorns MP, et al BJR, 79 (2006), 326-330
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 52: Typical Dose Levels in Paediatric Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Examination | Number of studies | DAP range (cGycm2) | Mean DAP (cGycm2)
Barium enema |  |  | 
0-1 | 21 | 0.7-32.8 | 11.2
1-7 | 11 | 2.8-76.8 | 15
8+ | 3 | 20.1-49 | 30.5
Intravenous urography |  |  | 
0-1 | 3 | 3-7.7 | 5.9
1-7 | 47 | 0.9-39.2 | 10.2
8+ | 22 | 7-135.5 | 38.5
  Hiorns MP, et al BJR, 79 (2006), 326-330
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 53: Special considerations
Patients that may have disease that make them more radiation sensitive such as ataxia talengiectasia or connective tissue disease
Pregnancy in adolescent girls
  Refer to Lecture 10 for more details

## Slide 54: Occupational Radiation Protection in Paediatric Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Lead lined aprons and thyroid collars must be worn at all times
Radiologist should use lead lined goggles if they are beside the beam for lengthy periods
Radiation badges must be worn as per recommendations of a national authority

## Slide 55: Occupational Radiation Protection in Paediatric Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Individual monitoring for whole body dose and extremities, if necessary, should be provided
Only essential personnel should be in the room during examination
Comforters such as parents, and caretakers such as nurses, should also be protected as above

## Slide 56: Radiation Protection Tools in Paediatric Fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Lead apron
Thyroid collars
Lead goggles
Lead gloves?
Stay behind someone else…..
Interupt fluoroscopy when the nurse helps the patient
Then, 1910...?
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 57
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
...and now
Radiation Protection Tools in Paediatric Fluoroscopy
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 58: How effective is a lead apron?
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
0.25 mm Pb-equivalence- reduce 90 %
0.35 mm Pb-equivalence- reduce 95 %
0.50 mm Pb-equivalence- reduce 99%
As a rule of thumb one can use for scattered radiation
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 59
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Undercoach fluoroscopy system
Without apron
With apron
Radiation Protection Tools in Paediatric Fluoroscopy
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 60
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Lead curtain at the table for an under-couch system
Radiation Protection Tools in Paediatric Fluoroscopy
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 61
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Lead curtain at the table for an undercoach system
Mobile lead screen with leaded glass
Radiation Protection Tools in Paediatric Fluoroscopy
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 62: Radiation protection tools in paediatric fluoroscopy
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Lead curtain at the table for an undercoach system
Mobile lead screen with leaded glass)
Mobile leaded glass screen attached to the ceiling
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 63
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology

## Slide 64
Fluoroscopy Parameters | Radiation Dose Parameters | 
 | Patient DOSE | Staff DOSE
Large Patient Size | Increases | Increases
Higher mA | Increases | Increases
Higher Filtration | Decreases | May Increase
Increasing Focus Skin Distance | Decreases | Depends on geometry—Lateral vs. PA & staff positioning
Image Receptor-Patient Distance | Decreases | Decreases
Using Electronic Magnification with Image Intensifiers | Increases | Increases
Using Anti-Scatter Grid | Increases | Increases
Using Wide Collimator | Increases | Increases
Beam on Time | Increases | Increases
Personnel Shielding | No Effect | Decreases

## Slide 65
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
http://rpop.iaea.org/RPoP/RPoP/Content/index.htm

## Slide 66: Summary
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Components of fluoroscopy systems
Justification in fluoroscopy
Operational and equipment consideration related to dose and image quality
Example of dose reduction through optimisation process. DAP-values
Occupational radiation protection consideration in paediatric fluoroscopy
Shielding tools

## Slide 67: Answer True or False
Pulsed fluoroscopy reduces dose.
It is necessary to use the antiscatter grid in every paediatric radiology examination.
Magnification should be always used in paediatric fluoroscopy, because of the small size of the patient.
Use large radiation fields not to miss anything.

## Slide 68: Answer True or False
True – Radiation emanating only as pulses rather than continuously provides significant time when there is no radiation in a cycle and thus helps to reduce dose.
False - The use of gird increases dose without significant impact in smaller children on image quality and is therefore not recommended (e.g. age <8).
False - Magnification increase the dose and it should be used only if it is necessary.
False – Radiological examinations are not whole body or large area screening examinations. Only part of the body that requires irradiation based on clinical indication should be irradiated. Collimation and positioning are very important tools for dose reduction. This requires a skilled operators.

## Slide 69: References
Radiation Protection in Paediatric Radiology L05. Radiation protection in fluoroscopy
<number>
Hiorns MP, Saini A, Marsden PJ, A review of current local dose-area product levels for paediatric fluoroscopy in a tertiary referral centre compared with national standards. Why are they different?, BJR, 79 (2006), 326-330
Brown PJ, Johnson LM Silberberg PJ, Thomas RD, Low dose, high quality pediatric fluoroscopy, Medica Mundi 45/1 March 2001
Tapiovaara MJ, Sandborg M, Dance DR. A search for improved technique factors in paediatric fluoroscopy. Phys Med Biol. 1999 Feb;44(2):537-59
B Hansson, T Finnbogason, P Schuwert J Persliden: Added copper filtration in digital paediatric double-contrast colon examinations: effects on radiation dose and image quality. Eur radiol 7 (1997) 1117-1122
Cook, V., Radiation protection and quality assurance in paediatric radiology, Imaging, 13 (2001) 229–238.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection in Pediatric Radiology