## Slide 1: L 10
WRITTEN PROCEDURES AND ORGANIZATION

## Slide 2: Answer True or False
The radiation protection programme must be designed to ensure safety, and prevent accidents, with the proper handling of radiation, including safety/risk assessment and the establishment of procedures to be followed in emergencies
Local rules are intended to provide adequate levels of protection and safety including line of responsibility and all information required for work in the area
Among the events that could potentially be categorized as misadministrations, injection of the wrong amount of radioactivity is not considered a misadministration so long as the activity injected is for a diagnostic, rather than a therapeutic, procedure

## Slide 3: Objective
To become familiar with the radiation protection organization, safety / risk assessment, designation of areas, the written procedures and local rules to ensure the safe operation of the PET/CT unit and production facilities and emergency procedures

## Slide 4
Radiation protection organization
Safety / Risk assessment
Controlled areas
Local rules and written procedures / Radiation safety manual
Emergency procedures
Content

## Slide 5
10.1 Radiation Protection Organization

## Slide 6: Management Requirements (BSS)
To ensure a culture of safety
Ensure QA
Take account of human factors
Ensure the use of qualified experts

## Slide 7: Technical requirements (BSS)
Security of sources
Prevention of accidents
Good engineering practice

## Slide 8: Radiation Protection Program
Determine measures and resources required to meet objectives and ensure they are provided
Review measures and resources regularly
Identify failures and shortcomings and take steps to prevent them
Establish consultation and co-operation of all parties
Maintain records of activities
More details at http://rpop.iaea.org

## Slide 9: Responsibilities of a Licensee (BSS)
The licensee may appoint other people to carry out actions and tasks related to these responsibilities, but they shall retain the responsibility for the actions and tasks themselves. Registrants and licensees shall specifically identify individuals responsible for ensuring compliance with the Standards

## Slide 10: Documentation of Radiation Protection Programme (RPP)
Specifically in relation to medical exposure, the system for calibration of sources, clinical dosimetry of patients and provision of quality assurance programmes
The system for constraining the exposure of patients’ family & friends, care givers and volunteers
The education and training programme on the nature of the hazards, protection and safety

## Slide 11: Documentation of Radiation Protection Programme, cont.
The methods for periodically reviewing and auditing the performance of the radiation protection programme
The plans to be implemented in case of intervention
The health surveillance programme for workers
The requirements for the assurance of quality and process improvement

## Slide 12: Radiation Protection Committee (RPC)
The RPC shall recommend actions to the licensee to:
  Correct identified deficiencies in the RPP
  Review new uses of radiation sources that may lead to modifications of the RPP
  Identify any amendments to the licence
  Identify staff needing personal authorization
  Periodically review the training programme

## Slide 13
10.2 Safety Assessment / Prior risk assessment

## Slide 14: Prior Risk Assessment
    Radiological Risk Assessment
Before any new activity involving the use of ionizing radiation, an assessment of the risks is required to indicate what control measures and special procedures are necessary to restrict exposure

## Slide 15: Safety / Risk Assessment
Identification of the mechanisms for exposure (both routine and accidental exposure)
Realistic estimate of doses and likelihood of accidental doses
Identification of possible safety system failures
Identification of protection measures needed

## Slide 16: Categorization of Hazard Weighting factors according to radionuclide
Class | Radionuclide | Weighting Factor
A | Se-75, Sr-89,I-125, I-131 | 100
B | C-11, N-13, O-15, F-18, Cr-51, Ga-67, Tc-99m, In-111, In-113m, I-123, Tl-201 | 1.00
C | H-3, C-14, Kr-81m, Xe-127, Xe-122 | 0.01

## Slide 17: Categorization of Hazards Weighting factors according to type of operation
Type of operation or area Weighting factor
Storage 0.01
Waste handling,
waiting area 0.10
Local dispensing,
radionuclide administration 1.00

## Slide 18: Categorization of Hazards
Patient administration, 370 MBq 18F
Weighting factor, radionuclide 1
Weighting factor, operation 1
Total weighted activity 370 MBq
Weighted activity Category
< 50 MBq Low hazard
50-50000 MBq Medium hazard
>50000 MBq High hazard

## Slide 19: Safety Assessment - Patient Procedure
Patient Procedure Incident
Request and scheduling Wrong patient
Identification Wrong patient
Information Pregnancy, nursing
Administration Misadministration
of radiopharmaceutical
Waiting Contamination
Emptying bladder Contamination
Examination Contamination
Poor image quality
Leaving the department Behaviour striction
        Medical emergency
      Death of patient
Notes: These are examples of what can go wrong in the different procedures where a patient is involved. The next step is to identify the people that might be exposed and finally try to define what to do to prevent the identified accidents and incidents and to set up an emergency plan to follow in the case of an accident,

## Slide 20: Safety Assessment - Worker
Procedure Incident
Transport Transport accident
Receipt and unpacking Uncontrolled exposure
        Spread of contamination
Handling of sources Spread of contamination
Uncontrolled exposure
Radioactive waste Contamination
Care Uncontrolled exposure
for a radioactive patient and contamination

## Slide 21: Safety Assessment - Source
Procedure Incident
Transport Transport accident
Storage Loss of sources
Radioactive waste Loss of sources
Contamination
Examination Contamination

## Slide 22: Safety Assessment - General Public
Procedure Incident
Radioactive patient Uncontrolled exposure
and contamination

## Slide 23: Safety Equipment
Shields
Protective clothing
Tools for remote handling
of radioactive material
Containers for radioactive waste
Dose rate monitor with alarm
Contamination monitor
Decontamination kit
Signs and labels
Records

## Slide 24: Monitoring and Verification of Compliance (BSS)
“2.38. Monitoring and measurements shall be conducted
of the parameters necessary for verification of compliance
with the requirements of the Standards.
2.39. For the purposes of monitoring and verification
of compliance, suitable equipment shall be provided
and verification procedures introduced. The equipment
shall be properly maintained and tested and shall be
calibrated at appropriate intervals with reference
to standards traceable to national or international
standards.”

## Slide 25
10.3 Designation of areas

## Slide 26: Classification of Areas
Controlled area
Supervised area
No
entry
Unsealed
Radionuclides
Risk of
contamination
Risk from
external radiation
No entry when red light is on
Radiation
Controlled
Area
X-rays
Authorised persons only

## Slide 27: Controlled Area (BSS)
“I.23. Registrants and licensees shall:
(a) delineate controlled areas by physical means or, where this is not reasonably practicable, by some other suitable means;
…………………..
(c) display a warning symbol, such as that recommended by the International Organization for Standardization (ISO), and appropriate instructions at access points and other appropriate locations within controlled areas;
(d) establish occupational protection and safety measures, including local rules and procedures that are appropriate for controlled areas;
(e) restrict access to controlled areas by means of administrative procedures, such as the use of work permits, and by physical barriers, which could include locks or interlocks; the degree of restriction being commensurate with the magnitude and likelihood of the expected exposures;”

## Slide 28: Controlled Area (BSS) cont.
“ (f) provide, as appropriate, at entrances to controlled
areas:
(i) protective clothing and equipment;
(ii) monitoring equipment; and
(iii) suitable storage for personal clothing;
(g) provide, as appropriate, at exits from controlled areas:
(i) equipment for monitoring for contamination of
skin and clothing;
(ii) equipment for monitoring for contamination of
any object or substance
being removed from the area;
(iii) washing or showering facilities; and
(iv) suitable storage for contaminated protective clothing and equipment; and
(h) periodically review conditions to determine the
possible need to revise the protection measures or safety
provisions, or the boundaries of controlled areas.”

## Slide 29: Controlled Areas
should have appropriate instructions at access points

## Slide 30: Controlled Areas
….. display a warning symbol

## Slide 31: CT
For CT there must be an X ray sign at each entrance to the room
It must switch on automatically when the X rays are on

## Slide 32
10.4 Local Rules and Radiation Safety Manual

## Slide 33: Local Rules and Supervision (BSS)
“ I.26. Employers, registrants and licensees shall, in consultation with
workers, through their representatives if appropriate:
establish in writing such local rules and procedures as are
necessary to ensure adequate levels of protection and safety for workers and other persons;
(b) include in the local rules and procedures the values of any relevant investigation level or authorized level, and the procedure to be followed in the event that any such value is exceeded;
(c) make the local rules and procedures and the protective measures and safety provisions known to those workers to whom they apply and to other persons who may be affected by them;
(d) ensure that any work involving occupational exposure be adequately supervised and take all reasonable steps to ensure that the rules, procedures, protective measures and safety provisions be observed; and
(e) when required by the Regulatory Authority, designate a radiation protection officer.”

## Slide 34: Local Rules
Written and practiced local rules
  Restrict exposure and ensure ALARA
  Cover normal working conditions
  Cover incidents and accidents

## Slide 35: Local Rules
Local rules are intended to provide adequate levels of protection and safety through the establishment of common work procedures and other systems to be followed by all workers in the area
They should include all information required for work in the area, and be made known to all workers
Notes: The picture shows many files - the lecturer should point out that local rules should not be like this:
easily accessible and understandable
comprehensive
not filed away

## Slide 36: Local Rules
Include:
Line of responsibility
Designation of areas
System of work
Investigation levels
Emergency procedures

## Slide 37: Example Local Rules
    1. Radiation Protection Supervisor
    The Radiation Protection Supervisor is….(name)
    The Radiation Protection Advisor is…..(name)
    2. Controlled Area
The controlled area is defined by the walls, floor, ceiling and room entrance
Example of mechanisms to contact the appropriate personnel in a case of an emergency:
Use of a doorplate posted on the door to a hot lab, identifying the responsible persons and the means by which to contact them

## Slide 38: System of Work
Staff should minimise contact with the patient, consistent with adequate patient care
All doors to the X ray room MUST be correctly closed before a CT scan is initiated
Visitors & care givers required to support a patient during exposure must be informed of risks, wear protective clothing if appropriate and be positioned to keep doses ALARP. The occurrence should be recorded in the book provided
Care should be taken that the smallest volume of the patient is scanned and the lowest exposure factors selected consistent with achieving the required diagnostic information
Daily contamination monitoring must be carried out and the results recorded. Any significant contamination found must be dealt with immediately

## Slide 39: Emergency Procedures
In case of a spill, follow the emergency procedures described in the departmental protocols and inform the Radiation Protection Supervisor
If there is any reason to suspect that a patient has received a dose greater than intended the concerned technical staff (i.e. Superintendent Radiographer) and Facility Manager must be informed
All faults and irregularities with the PET/CT unit must be noted in the fault book and reported to the Superintendent Radiographer and Radiology or Nuclear Medicine Services Manager immediately.
In the case of fire, follow the procedures in the Radiation Protection Manual
An Incident Report Form must be used to record all radiation incidents

## Slide 40: Investigation Levels
Whole Body Eyes/Shoulder Single Organ
1 month 1 mSv 7 mSv 25 mSv
3 month 3 mSv 22 mSv 75 mSv
Note: Single organ is skin or thyroid

## Slide 41: Radiation Safety Manual (RSM)
The RSM is ideally a local reference book, and made available to (and possibly mandatory reading for) all radiation workers
An effective aid to training
This is now required practice in many countries
Includes responsibilities, all documented policies, procedures, protocols and forms
Notes: There are several radiation safety manuals available on the web - a search engine using the key word ‘radiation safety manual’ will bring these up. If time permits, the lecturer could visit some of the sites with participants.

## Slide 42
10.5 Contingency Plans

## Slide 43: Accident (BSS)
A accident is any unintended event, including operating errors, equipment failures and other mishaps, the consequences or potential consequences of which cannot be ignored from the radiation and safety point of view, and which can lead to potential exposure and subsequently to abnormal exposure conditions

## Slide 44: Emergency Kit
Should be kept readily available for use in an emergency. It may include the following:
  Protective clothing e.g. overshoes, gloves
  Decontamination materials for the affected areas including absorbent materials for wiping up spills
  Decontamination materials for persons
  Warning notices
  Portable monitoring equipment
  Bags for waste, tape, labels, pencils

## Slide 45: Emergency Facilities
An emergency eye-wash should be installed near the hand-washing sink
There should be access to an emergency shower in or near the dose preparation area

## Slide 46: Contingency Plans
Contamination
CT failure to terminate
Death of a patient
Medical emergency
Fire
Loss / damage of a source
Misadministration

## Slide 47: Emergencies
All people in the PET/CT department shall be trained in handling emergencies

## Slide 48: Contamination
Spill
Radioactive patient body fluids

## Slide 49: Decontamination Procedures
Use adsorbent paper on wet spill or wet absorbent
paper on dry spill
Repetitively swab the area inwards towards the center
of the spill
Place contaminated paper in a plastic bag or container
Monitor the area
Repeat the procedure until the exposure rate is below
given limits
If the decontamination is not successful, mark the
contaminated area and classify the room as a controlled
area (if not already done) until the contamination is
completely removed
These procedures should be practiced!

## Slide 50: Decontamination of Skin
If contamination of the skin occurs:
Wash the area thoroughly using mild soap and tepid (not hot) water
Particular care should be paid to cleaning under the fingernails
If this does not bring the contamination to an acceptably low level the procedure should be repeated using a decontaminating detergent
Scrub with a nail brush but take care not to break the skin

## Slide 51: CT – Failure to Terminate
Turn mains off - Press emergency ‘off’ button on console or walls
Take note of all relevant machine settings
Note off button on gantry does not terminate X rays
Remove patient
Call service
Report incident and estimate dose received

## Slide 52: Death of a Patient
No precautions are necessary unless death occurs whilst still in the unit
Prior to this small external dose rate hazard, common sense use of distance and shielding (if possible) should be used. Early Post Mortem or transfer of body off-site not advised unless individual risk assessment shows minimal risks

## Slide 53: Medical Emergency
Contact the Radiation Protection Officer for specific instructions
Medical personnel should proceed with emergency care while attempting to take precautions against spread of contamination:
  avoid direct contact with patient’s mouth,
  all members of the emergency team should wear impermeable protective gloves.
Medical personnel shall be informed and trained in how to deal with a radioactive patient

## Slide 54: Fire
The normal hospital drill should be observed and the safe evacuation of patients, visitors and staff is the most important consideration
When the fire brigade arrives, they should be informed of the presence of radioactive material
No one is allowed to re-enter the building until it has been checked for contamination

## Slide 55: Loss of Shipment
Check all possibilities in the hospital
If still not found, call the company and inform them of the failure so they can trace the shipment and find out where the radioactive material is
If not found, the loss of the material should be reported according to rules given by the Regulatory Authority

## Slide 56: Damage to Radioactive Shipment
If transport container is damaged
Do not accept the shipment
If the contents are damaged
Cordon off the area
Treat as a major spill

## Slide 57: Misadministration
Wrong patient
Wrong activity
  diagnosis > 10% of prescribed activity
Tissued injection

## Slide 58: Causes of Misadministration
Communication problems
Busy environment, distraction
Unfamiliar with written procedures
Not clearly defined responsibilities
No efficient quality assurance

## Slide 59: Consequences of Misadministration
Non-justified exposure
Increased radiation risk
Delayed diagnosis
Increased cost
Increased work load
Reduced confidence

## Slide 60: How to Avoid Accidents and Misadministrations
Culture of safety
Safety assessments to define critical procedures and emergency situations
Reporting system (When? Where? Why?)
Education and training: initial, continuing

## Slide 61: Investigation of Accidental Medical Exposure
Inform responsible PET/CT physician
Inform patient and referring physician
Calculate dose
Indicate corrective measures
Implement measures
Submit report to Radiation Protection Advisor and Regulatory Authority

## Slide 62: SUMMARY OF WRITTEN PROCEDURES AND ORGANIZATION
The radiation protection program must be designed to ensure safety, and prevent accidents, with the proper handling of radiation, including safety/risk assessments and specific procedures for handling unexpected situations
Controlled areas must have appropriate instructions and warnings posted in easily accessible locations
Local rules are intended to provide adequate levels of protection and safety, including clearly indicated lines of responsibility, and all information required for work in the area
Emergency procedures and contingency plans must be established and the staff trained in their use