## Slide 1: L 3
MEDICAL EXPOSURE- BSS Requirements

## Slide 2: Answer True or False
PET absorbed dose is influenced by amount of injected radioactivity, tracer kinetics and radionuclide physical half-life
CT absorbed dose is determined by CT scanner settings, including kVp and mAs per rotation, scan length and pitch
Typical 18F-FDG PET/CT whole body oncology imaging procedures deliver far more radiation dose from the PET scan than the negligible amount of additional effective dose delivered by the CT scan

## Slide 3: Objective
To become familiar with the BSS requirements for medical exposure in PET/CT: responsibilities, training, justification, optimization, guidance levels, and dose calculations

## Slide 4: Content
Responsibilities
Training
Justification
Optimization
Guidance Levels
Absorbed Dose

## Slide 5: International
International Basic Safety Standard (BSS)
“…marks the culmination of efforts that have continued over the past several decades towards the harmonization of radiation protection and safety standards internationally”
This BSS is now undergoing revision

## Slide 6
3.1 Responsibilities

## Slide 7: Responsibilities (BSS)
(a) no patient be administered a diagnostic or therapeutic medical exposure unless the exposure is prescribed by a medical practitioner;
(b) medical practitioners be assigned the primary task and obligation of ensuring overall patient protection and safety in the prescription of, and during the delivery of, medical exposure;
(c) medical and paramedical personnel be available as needed, and either be health professionals or have appropriate training adequately to discharge assigned tasks in the conduct of the diagnostic or therapeutic procedure that the medical practitioner prescribes;
II.1. Registrants and licensees shall ensure that:

## Slide 8: Responsibilities (BSS) cont.
(d) for therapeutic uses of radiation (including teletherapy and brachytherapy), the calibration, dosimetry and quality assurance requirements of the Standards be conducted by or under the supervision of a qualified expert in radiotherapy physics;
(e) the exposure of individuals incurred knowingly while voluntarily helping (other than in their occupation) in the care, support or comfort of patients undergoing medical diagnosis or treatment be constrained as specified in Schedule II; and
(f) training criteria be specified or be subject to approval, as appropriate, by the Regulatory Authority in consultation with relevant professional bodies.

## Slide 9: Responsibilities (BSS) cont.
II.2. Registrants and licensees should ensure that for diagnostic uses of radiation the imaging and quality assurance requirements of the Standards be fulfilled with the advice of a qualified expert in either radiodiagnostic physics or nuclear medicine physics, as appropriate.
II.3. Medical practitioners shall promptly inform the registrant or licensee of any deficiencies or needs regarding compliance with the Standards with respect to protection and safety of patients and shall take such actions as may be appropriate to ensure the protection and safety of patients.

## Slide 10
3.2 Training

## Slide 11: Reporting of Scans
Who can report PET/CT scans?
Issues
  Nuclear medicine physicians may not have CT training
  Radiologists may not have nuclear medicine training
How much cross-training is needed?

## Slide 12: UK
Nuclear medicine physician training includes CT component
The ‘Standards for delivering a PET service within the UK’ suggests medical practitioners should experience reporting 300 PET/CT scans per year
Administration of Radioactive Substances Advisory Committee (ARSAC) also recommend 600 PET/CT scans and 3 months training in a PET/CT centre to obtain a certificate
Notes: ARSAC - Administration of Radioactive Substances Advisory Committee

## Slide 13: UK
Radiologists can train in radionuclide radiology or dual accreditation in radiology and nuclear medicine
To hold an ARSAC certificate, radiologists must have either radionuclide radiology or nuclear medicine training with the same PET training as a nuclear medicine physician
Notes: ARSAC - Administration of Radioactive Substances Advisory Committee

## Slide 14: Germany
A nuclear medicine physician is allowed to perform low dose CTs in the "attenuation correction mode" (effective dose below 2 mSv)

## Slide 15: Society of Nuclear Medicine
‘Procedure Guidelines for Tumour Imaging with PET/CT’ recommend 150 supervised PET/CT reports and a further 500 CT reports if the clinician is not trained in CT

## Slide 16: Training
Who can operate the CT unit?
  Can a nuclear medicine trained member of staff operate a CT unit?
  What if a contrast agent is to be given?
  How much training is needed?
Who can inject radioactivity?
  Can a CT trained member of staff handle radioactivity without supervision?
  How much training is needed?

## Slide 17: American Society of Radiologic Technologists and Society of Nuclear Medicine
Consensus statement
Any registered radiographer (RT(R)) or nuclear medicine technologist (RT(N) or CNMT) may operate PET-CT equipment after obtaining appropriate additional training and demonstrate competency
Lists a curriculum for cross-training

## Slide 18: UK
Standards for delivering a PET service within the UK - A report of the Intercollegiate Standing Committee on Nuclear Medicine
Representing
    The Royal College of Physicians of London
    The Royal College of Physicians and Surgeons of Glasgow
    The Royal College of Physicians of Edinburgh
    The Royal College of Pathologists
    The Royal College of Radiologists, and
    The British Nuclear Medicine Society
States the need for adequate training

## Slide 19: Who Else is Involved?
Medical Physics Expert (MPE)?
  Dosimetry experience in Nuclear Medicine and CT?
  Radiation protection experience in both?
  Training in Nuclear Medicine / CT?
Notes: MPE - Medical Physics Expert

## Slide 20
3.3 Justification

## Slide 21: Justification
Every medical exposure must be justified
Often a trained and medically qualified expert who makes this decision
Balances (potential) benefit and risk

## Slide 22: Justification of Patient Examination
Diagnostic benefit Risk of medical exposure
Notes: Justifaction of medical exposure means ensuring a proper balance between the diagnostic objective and radiation burden for the patient

## Slide 23: Justification
Applies to the decision to proceed with an investigation
Justification applies to all the steps thereafter to optimise the exposure
Optimization may result in a higher or lower dose but usually is applied as the lowest dose consistent with diagnostic purpose

## Slide 24
3.4 Optimization

## Slide 25: Factors Affecting Medical Exposure
Request
Examination or treatment
Biomedical research
Treatment
Examination
Administration of
radiopharmaceutical
Patient identification
and information
Child ?
Lactating woman ?
Pregnant woman ?
Safe preparation of
prescribed radio-
pharmaceutical and
activity
Dose constraints
Method
Quality of equip-ment
Correct use of equipment
Dose constraints to comforters and visitors
Accidental medical exposure
Who is responsible?
Notes: This image is aimed as an overview of the different factors affecting medical exposure. It should be used as an introduction to a discussion about who should be responsible.

## Slide 26: Optimization of Medical Exposure
Administered
activity
Radiation risk
Quality of
Equipment, etc
Effective
dose
Image quality
Radionuclide
Biokinetics
Patient
Notes: This image shows in more detail the two branches originating from the administered activity, the radiation risk and the image quality.

## Slide 27: Optimization of Medical Exposure in PET
BSS II.17. Registrants and licensees shall ensure that:
(a) the medical practitioners who prescribe or conduct diagnostic applications of radionuclides:
  (i) ensure that the exposure of patients be the minimum required to achieve the intended diagnostic objective;
  (ii) take into account relevant information from previous examinations in order to avoid unnecessary additional examinations; and
  (iii) take into account the relevant guidance levels for medical exposure;

## Slide 28: Optimization (BSS II.17 cont)
(b) the medical practitioner, the technologist or other imaging staff, as appropriate, endeavour to achieve the minimum patient exposure consistent with acceptable image quality by:
  (i) appropriate selection of the best available radiopharmaceutical and its activity, noting the special requirements for children and for patients with impairment of organ function;

## Slide 29: Optimization (BSS II.17 cont)
(c) administration of radionuclides for diagnostic or radiotherapeutic procedures to women pregnant or likely to be pregnant be avoided unless there are strong clinical indications;

## Slide 30: Optimization (BSS II.17 cont)
(d) administration of radionuclides to children for diagnostic procedures be carried out only if there is a strong clinical indication, and the amount of activity administered be reduced according to body weight, body surface area or other appropriate criteria.

## Slide 31: Optimization of dose from CT
Use factors consistent with minimizing dose while maintaining adequate image quality, e.g. mA, kVp, rotation time, pitch
Use all dose reduction features available, e.g. tube current modulation, automatic exposure control, adaptive collimation, ECG-pulsing for cardiac imaging, if appropriate
Only scan the relevant areas
Minimize repeat scans with contrast

## Slide 32
3.5 Guidance Levels

## Slide 33: Guidance Levels (BSS)
2.27. Guidance levels for medical exposure shall be established for use by medical practitioners. The guidance levels are intended:
(a) to be a reasonable indication of doses for average sized patients;
(b) to be established by relevant professional bodies in consultation with the Regulatory Authority following the detailed requirements of Appendix II and the guidance levels given in Schedule III;
(c) to provide guidance on what is achievable with current good practice rather than on what should be considered optimum performance;
(d) to be applied with flexibility to allow higher exposures if these are indicated by sound clinical judgment; and
(e) to be revised as technology and techniques improve.

## Slide 34: Guidance Levels (BSS)
II.24. Registrants and licensees should ensure that guidance levels for medical exposure be determined as specified in the Standards, revised as technology improves and used as guidance by medical practitioners, in order that:
(a) corrective actions be taken as necessary if doses or activities fall substantially below the guidance levels and the exposures do not provide useful diagnostic information and do not yield the expected medical benefit to patients;
(b) reviews be considered if doses or activities exceed the guidance levels as an input to ensuring optimized protection of patients and maintaining appropriate levels of good practice; and
(c) for diagnostic radiology, including computed tomography examinations, and for nuclear medicine examinations, the guidance levels be derived from the data from wide scale quality surveys which include entrance surface doses and cross-sectional dimensions of the beams delivered by individual facilities and activities of radiopharmaceuticals administered to patients for the most frequent examinations in diagnostic radiology and nuclear medicine respectively.

## Slide 35: Guidance Levels (BSS)
II.25. In the absence of wide scale surveys, performance of diagnostic radiography and fluoroscopy equipment and of nuclear medicine equipment should be assessed on the basis of comparison with the guidance levels specified in Schedule III, Tables III-I to III-V. These levels should not be regarded as a guide for ensuring optimum performance in all cases, as they are appropriate only for typical adult patients and, therefore, in applying the values in practice, account should be taken of body size and age.

## Slide 36: Guidance Levels - PET
Notes: From ARSAC (Administration of radioactive substances Advisory committee, UK)

## Slide 37
3.6 Dose Quantities and Units

## Slide 38: Absorbed Dose
Energy absorbed in a material per unit mass
SI unit is gray (1 Gy = 1 J/kg)
In CT can be specified as Computed Tomography Dose Index, CTDI, in mGy in air or in a phantom

## Slide 39: Equivalent Dose
Absorbed dose weighted for the type of radiation and averaged over the whole organ or tissue (except in the case of skin)
SI unit is Sievert (Sv)
Weighting factor for electrons, positrons, gamma and X ray is 1
Equivalent dose (Sv) = Absorbed dose (Gy) for a weighting factor = 1

## Slide 40: Effective Dose
Weighted sum of equivalent doses to individual organs and tissues
SI unit is Sievert (Sv)
Weighting factors take into account relative susceptibility of different tissues to radiation damage
Allows for the assessment of long-term risk of harm from low-level exposure

## Slide 41: Dose Length Product
A quantity displayed on the CT scanner
Equals dose (CTDIw) multiplied by scan length
Unit is mGy•cm
Is used to describe reference levels for CT scans

## Slide 42: Absorbed Dose to an Organ in Nuclear Medicine is determined by:
Radionuclide
Activity administered
Activity and time in the organ
Size and shape of the organ
Activity in other organs
Kinetics of radiopharmaceutical
Quality of radiopharmaceutical

## Slide 43: Absorbed Dose in CT is determined by:
kVp and mAs per rotation
Scan length
Pitch
Scan width / volume
Characteristics of the CT scanner – filter, type of reconstruction kernel, etc.
Size and shape of the organ
Notes: CTDI – computed tomography dose index

## Slide 44: Equilibrium Absorbed Dose Constants Gy/MBq/h
0.709
0.570
0.139
F-18
0.328
0.219
0.109
I-131
0.082
0.072
0.01
Tc-99m
Total
Penetrating
Non-penetrating
PET radionuclide - Energy deposition similar to therapy rather than diagnostic radionuclides

## Slide 45: Absorbed Dose Estimates (FDG)
0.021
Uterus
0.028
Brain
0.062
Heart
0.16
Bladder
Absorbed Dose
(mGy/MBq)
Organ
ICRP Publication 53. Radiation Dose to Patients from Radiopharmaceuticals

## Slide 46: Typical Effective Doses - PET
Notes: From ARSAC (Administration of radioactive substances Advisory committee, UK)

## Slide 47: European Diagnostic Reference Levels CT
570
Pelvis
780
Abdomen
650
Chest
1050
Brain
Dose length product (mGy•cm)
Scan
Notes: Dose length product is displayed on many scanners. These levels are likely to change as they were not based mainly on multi slice scanners

## Slide 48: Typical Effective Dose CT
10
Pelvis
10
Abdomen
8
Chest
2
Brain
Effective Dose
(mSv)
Scan
Notes: Note these are for standard diagnostic scans. Doses in PET CT can be 50 to 80% of these values

## Slide 49: SUMMARY OF MEDICAL EXPOSURE
The BSS details the responsibilities of all parties involved in the appropriate actions involved in all aspects of the ordering, performance, and reporting of results of PET/CT scans
CT absorbed dose is determined by kVp and mAs per rotation, scan length and pitch and the size and shape of the scanned organ
Typical 18F-FDG PET/CT whole body oncology imaging procedures involve the delivery of 5-8 mSv of effective dose from the PET scan and an additional 8-10 mSv for the CT scan