## Slide 1: L 5
FACILITY DESIGN

## Slide 2: Answer True or False
Radiation Protection in PET/CT
<number>
Medical cyclotrons require extensive internal shielding to adequately protect occupationally exposed workers
Adequate structural shielding is needed for the PET scanner whereas the requirements are less for the CT scanner
Building materials should be used in the design of PET/CT facilities that are easily decontaminated on a daily basis in all areas where liquid radiopharmaceuticals are handled

## Slide 3: Objective
Radiation Protection in PET/CT
<number>
Considerations to minimize staff doses when designing a new PET/CT and/or cyclotron facility, including shielding and layout issues

## Slide 4: Content
Radiation Protection in PET/CT
<number>
Cyclotron design
PET/CT department design
Structural shielding
Building requirements

## Slide 5
5.1 Cyclotron Design

## Slide 6: Example 1 of technical features of a cyclotron
Radiation Protection in PET/CT
<number>

## Slide 7: Example 2 of technical features of a cyclotron
Radiation Protection in PET/CT
<number>
18 MeV proton beam
In vault
150 µA dual beam
9 MeV deuteron beam with 40 µA intensity
8 independent targets
Possible upgrades:
  Double proton ion sources
  Additional targets to produce 124I, 123I, 64Cu

## Slide 8: Cyclotrons - Radiation
Radiation Protection in PET/CT
<number>
Prompt radiation
  Radiation exposure – primarily gamma
  On shield surface near targets and seams between shield blocks the neutron dose = 10-50% of total measured dose
Room door closed during bombardment to prevent casual entry
Residual radiation
  Low levels after cool down (could be 2 days)
  Cyclotron servicing: Survey before work

## Slide 9
Radiation Protection in PET/CT
<number>

## Slide 10: PET Cyclotron - Technical Consideration for Radiation Safety
Radiation Protection in PET/CT
<number>
Cyclotron: Self-shields vs. Vault
Room shielding
Activation components: Protons & neutrons
Safety interlocks
Cyclotron ON lights
Room radiation monitors
Preventive maintenance (PMS)
  Surveys
  Pocket dosimeters
  Action levels
Scram buttons
Target rebuilds
Activated components-storage
Waste disposal: long-lived

## Slide 11
Radiation Protection in PET/CT
<number>
Technical considerations features of a typical cyclotron
15 cm steel cylindrical magnet acts as primary shield
Cyclotron enclosed in cylindrical shielding system consisting of 68 cm thickness of boron-doped water
Wall of vault is 60 cm thick concrete

## Slide 12
Radiation Protection in PET/CT
<number>
Technical considerations features of typical self-shielded cyclotrons

## Slide 13
Radiation Protection in PET/CT
<number>
Examples of cyclotron shielding
Example 1
Example 2
Notes: Note not all “self shielded” cyclotrons have sufficient protection depending on use of surrounding areas

## Slide 14
Radiation Protection in PET/CT
<number>
Some typical cyclotron gamma exposure rates
Notes: Doses to engineers and staff working with cyclotrons can be substantial

## Slide 15
5.2 Department Design

## Slide 16: Design Aspects to Consider
Radiation Protection in PET/CT
<number>
Delivery of radiopharmaceutical
Storage of radioactive material
Dose preparation
Administration
Resting rooms
Lavatory facilities
Scanning room
Control room
Post-scan requirements
Accompanying persons

## Slide 17: Typical Patient Instantaneous Dose Rates
Radiation Protection in PET/CT
<number>
For dose rates measured at 0.1 m and 1 m immediately after injection
Dose rate (
µ
Sv/hr/MBq)
0.1m
1m
Bone
0.27
0.02
FDG
2
0.22

## Slide 18: Air Kerma Rate Constants (µGym2/Bqh)
Radiation Protection in PET/CT
<number>
C-11 140
N-13 140
O-15 140
F-18 140
Tc-99m 14
I-131 53

## Slide 19: The Radioactive Patient (95th percentile immediately after injection)
Radiation Protection in PET/CT
<number>
0.1 0.5 1 2 m
0.8 0.3 0.09 0.04 mSv/h
Contamination External
saliva
perspiration
breath
urine
400 MBq 18F
Benatar NA, Cronin BF, O’Doherty M. Radiation dose rates from patients undergoing PET: implications for technologists and waiting areas. Eur J Nucl Med 2000: 27: 583-9

## Slide 20: Layout of a Nuclear Medicine Department
Radiation Protection in PET/CT
<number>
From high to low activity
Notes: This image demonstrates the general layout of a nuclear medicine department trying to seåarate high activity levels facilities from waiting area etc

## Slide 21: Shielding
Radiation Protection in PET/CT
<number>
Much cheaper and more convenient to shield the source, where possible, rather than the room or the person
Structural shielding is generally not necessary in a nuclear medicine department, but becomes necessary with PET-CT
However, more extensive and heavier shielding usually is required in facilities that use 18F versus those that do not

## Slide 22: Differences for a Facility using 18F versus One that Does Not
Radiation Protection in PET/CT
<number>
Higher energy gamma rays are more penetrating - standard lead/concrete protection is not adequate
Dose rates are higher than those for 99mTc
Staff should be outside the scanning room (in a control room as with CT scanning), not inside the PET scanning room during acquisitions

## Slide 23: Other Considerations
Radiation Protection in PET/CT
<number>
Resting phase requires patients to be within facility for many hours
All rest rooms may be occupied all day for a high-volume facility
Post-scan patients are hungry and may require refreshment before being sent home
Separate areas for patients not yet injected, and those accompanying patients, are likely to be required

## Slide 24: Areas of Concern
Radiation Protection in PET/CT
<number>
Staff whole body dose can be significantly higher than with conventional nuclear medicine
Staff extremity doses can approach dose limits without good technique and shielding
Public dose limits can be exceeded in surrounding areas if structural shielding is not adequate
Multislice CT scanners may need protection to full ceiling height

## Slide 25: Shielding Design Issues
Radiation Protection in PET/CT
<number>
Construction, breeze blocks/plasterboard partitions/single course of brick cladding
Building shared with non-radiation workers
Buildings/areas very close to scanner suite
Areas above and below scanner

## Slide 26: Preconstruction Design Issues
Radiation Protection in PET/CT
<number>
Dose constraints for staff and public must be adopted in designing the facility
Layout of department should be considered. Direct lines of sight between resting areas and staff areas should be eliminated
Shielding should be calculated taking into account all radiation sources
Allowance should be made for the short half life of the radionuclides to avoid over- protection

## Slide 27: Postconstruction Design Issues
Radiation Protection in PET/CT
<number>
Following construction, if actual measured exposure levels are too high, shielding must be increased or other corrective measures taken
Diligent monitoring of staff and public exposure levels must be performed
Any changes with time, such as significant increase in the number of patients handled per day, may necessitate increased shielding or other corrective measures to remain in compliance

## Slide 28: Layout of a Standard Department – Inadequate for PET Imaging
Radiation Protection in PET/CT
<number>
X
Defects:
Direct line of sight from resting patient
No control room – inadequate protection for operators
High dose rate to in vivo counting
Notes: This image demonstrates the general layout of a nuclear medicine department trying to separate high activity levels facilities from waiting area etc

## Slide 29: Good Design (1)
Radiation Protection in PET/CT
<number>
Notes: Cold waiting area to left. Short route for delivery of radionuclide. Staff protected from resting patients. Lavatory facilities close to rest and scan rooms

## Slide 30: Good Design (2)
Radiation Protection in PET/CT
<number>

## Slide 31: Inadequate Trailer Design Resulting in High Operator Dose
Radiation Protection in PET/CT
<number>

## Slide 32: Inadequate Trailer Design Resulting in High Operator Dose
Radiation Protection in PET/CT
<number>

## Slide 33
5.3 Shielding

## Slide 34: Shielding
Radiation Protection in PET/CT
<number>
incident
radiation
transmitted
radiation
Barrier thickness

## Slide 35: Definitions
Radiation Protection in PET/CT
<number>
Dose rate constant
The dose rate (μSv/h) at 1 m from a point source of activity = 1 MBq
TVL
Tenth value layer, which is the thickness of a material that reduces the number of incident photons by a factor of 10.

## Slide 36: 18F Physical Data and Attenuation Characteristics
Radiation Protection in PET/CT
<number>
511 keV gamma
TVL 17 mm lead (Delacroix Rad. Prot. Dos. 1998)
TVL 150 mm concrete (2350 kg/m3)
TVL 176 mm solid concrete blocks (2000kg/m3)

## Slide 37: Structual Shielding
Radiation Protection in PET/CT
<number>
The absorbed dose is determined by factors such as:
  source strength
  length of exposure
  distance from the source
  transmission through the protective barrier

## Slide 38: Sample Design Criteria
Radiation Protection in PET/CT
<number>
Assume typical 400 MBq injected activity
Resting phase 1 hour
Scanning phase 30 mins
Workload supplied by hospital
Dose constraint for all areas outside resting/scanning rooms 300 Sv
Occupancy factors included in some areas (fraction of time a given room is occupied)

## Slide 39: Dose Rate from Patients - 18F
Radiation Protection in PET/CT
<number>
65 Sv/h predicted from point source calculation
33 mSv/h at 5 cm from unshielded syringe with 555 MBq of 18F
max 70 Sv/h at 1m after injection
AAPM Task Group 108: PET and PET/CT Shielding Requirements
Med. Phys. 33, Issue 1, January 2006;
DOI: 10.1118/1.2135911

## Slide 40: Comments
Radiation Protection in PET/CT
<number>
Standard building material may not afford sufficient protection for PET studies
Each facility individually needs to be analyzed carefully
Generally, 300 mm concrete appears to be conservative and is considered “safe”
There is a need to consider shielding for patient’s administration room and if regulations require for patient waiting area

## Slide 41: CT Scatter Plot
Radiation Protection in PET/CT
<number>

## Slide 42: (PET/) CT Scatter Plot
Radiation Protection in PET/CT
<number>

## Slide 43: Room Shielding
Radiation Protection in PET/CT
<number>
CT unit needs separate control area
Operator cannot sit in the room with the patient
Use CCTV to watch, and an intercom to communicate with patient
AAPM Task Group 108: PET and PET/CT Shielding Requirements
Med. Phys. 33, Issue 1, January 2006;
DOI: 10.1118/1.2135911

## Slide 44
5.4 Building requirements

## Slide 45: Building Requirements
Radiation Protection in PET/CT
<number>
Category Structural shielding Floors Worktop surfaces
of hazard walls, ceiling
Low no cleanable cleanable
Medium no continuous cleanable
sheet
High possibly continuous cleanable
one sheet
folded to
walls
The use of the room should be taken into account,
e.g. a waiting room as opposed to a control room.

## Slide 46: Floors
Radiation Protection in PET/CT
<number>
Impervious material
Washable
Chemical-resistant
Curved to the walls
All joints sealed
Glued to the floor
NOTE: No carpet!

## Slide 47: Walls and Ceiling
Radiation Protection in PET/CT
<number>
Should be finished in a smooth and washable surface with joints being sealed, wherever practicable. Walls should be painted with washable, non-porous paint (e.g. glossy paint)
The use of the room should be taken into account,
e.g. a waiting room as opposed to a control room
Notes: This image should be used to illustrate that categorization of hazard and safety assessments not solely should used to determine the furnishing of a room but also its use.

## Slide 48: Worktop Surfaces
Radiation Protection in PET/CT
<number>
  Worktop surfaces must be finished in a smooth, washable and chemical-resistant surface with all joints sealed
  Open shelving should be kept to a minimum to prevent dust accumulation
  Services (e.g. gas, electricity, vacuum) should not be mounted on top of the bench, but on walls or on panels for this purpose
  Light fixtures should be easy to clean and of an enclosed type in order to minimize dust accumulation

## Slide 49: Worktop Surfaces
Radiation Protection in PET/CT
<number>
Structural reinforcement may be necessary, since a considerable weight of lead shielding may be placed on work tops
Notes: The lecturer can ask the students if there is something wrong in the design of the lead shields although the image basically is an illustration of the possible need of reinforcement of the bench.

## Slide 50: Sinks
Radiation Protection in PET/CT
<number>
If the Regulatory Authority allows the release of aqueous waste to the sewer, a special sink shall be used
Local rules for the discharge shall be available
The sink shall be easy to decontaminate
Special flushing units are available for diluting the waste and minimizing contamination of the sink

## Slide 51: Washing Facilities
Radiation Protection in PET/CT
<number>
The wash-up sink should be located in the dose preparation area adjacent to the work area
Taps should be operable without direct hand contact and disposable towels or hot air dryer should be available

## Slide 52: Patient Toilet
Radiation Protection in PET/CT
<number>
A separate toilet room for the exclusive use of injected patients
The patient washing facilities SHOULD NOT be used by hospital staff, as it is likely that the floor, toilet seat and sink faucet handles will be contaminated frequently
Sited so that staff do not have to accompany patient

## Slide 53: Patient Toilet
Radiation Protection in PET/CT
<number>
The facilities should:
  Include a sign requesting patients to flush the toilet well and wash their hands
  Include a wash-up sink as a normal hygiene measure
  Be finished in materials that are easily decontaminated
  Consider wall mounted sanitary ware so that floor is completely clear

## Slide 54: Rest Room
Radiation Protection in PET/CT
<number>
CCTV to monitor patient
Be finished in materials that are easily decontaminated
Lights that can be dimmed
Quiet area
Separate area for each patient

## Slide 55: Dispensing Area
Radiation Protection in PET/CT
<number>
Be finished in materials that are easily decontaminated
Be tidy!

## Slide 56: Emergency Facilities
Radiation Protection in PET/CT
<number>
An emergency eye-wash should be installed near the hand-washing sink
There should be access to an emergency shower in or near the dose preparation area

## Slide 57: SUMMARY OF FACILITY DESIGN
Radiation Protection in PET/CT
<number>
Because cyclotrons accelerate particle beams at high energy for the production of positron emitters, it is important for them to have adequate shielding to protect occupationally exposed workers
Adequate structural shielding is needed to maintain exposure rates below established acceptable limits due to the radiotracers used for PET imaging as well as the X ray flux involved with CT imaging
It is necessary that the facility be designed so as to minimize dose both to occupationally exposed personnel and to the public at large, and this includes the use of building materials that are easily decontaminated on a daily basis in all areas where liquid radiopharmaceuticals are handled