## Slide 1: L 11
QUALITY CONTROL

## Slide 2: Answer True or False
QC procedures are required to assess the radiopharmaceutical purity, not the chemical purity, of the agent to be injected into the patient
QC tests must be performed not only for PET alone, but also for CT alone
It is important that QC be performed for PET and for CT machines by established protocols in accordance with pre-determined daily, weekly, and quarterly schedules

## Slide 3: Objective
To consider the QC needed on the production of the radiopharmaceutical and optimization of each PET and CT scanner, and their combined usage

## Slide 4: Content
Radiopharmaceutical
Calibrator
PET/CT scanner

## Slide 5
11.1 Radiopharmaceutical

## Slide 6: QC on FDG
20-30 minutes following production
QC for sterility, pyrogens, radiochemical purity, radionuclidic purity, chemical purity, pH, clarity and particulates
Notes: QC on FDG
20-30 additional minutes
QC for sterility, pyrogens, radiochemical purity, radionuclidic purity, chemical purity, pH, clarity and particulates
PETNet meets or exceeds USP standards

## Slide 7
Radiochemical purity
  Fraction of a specified radionuclide present in the desired chemical form and in the specific molecular position
Chemical purity
  Fraction of the compound in the specified molecular form
  Appropriate chemical purity is mandatory to avoid adverse reactions, pharmacological or toxic effects and avoidable irradiation of other organs

## Slide 8: Chemical and Radiochemical Analysis
Chromatography
  Technique used to separate components within a sample
  Many types but all work on the principle that different components migrate at different rates
  Stationary phase and mobile phase
  Gas Chromatography (GS)
  Thin Layer Chromatography (TLC)
  High Performance Liquid Chromatography (HPLC)

## Slide 9: Use of Reference Standard
100% radiochemical
purity
Spike with known standard to identify peak
FDG
Glucose
FDG

## Slide 10
11.2 Calibrator

## Slide 11: Calibrator
Daily sensitivity/consistency check
Linearity
Dependence on the position of source within the calibrator
Certified calibration sources of Cs-237, Ge-68 or Na-22 to perform QC tests
IAEA-TRS-454 Quality Assurance for Radioactivity Measurement in Nuclear Medicine 2006 , IAEA QA for PET and PET/CT systems (in press),
IAEA QA for SPECT systems (in press)

## Slide 12
11.3 PET/CT Scanner

## Slide 13: PET/CT scanner QC
Acceptance test
  Very important to set reference values against which routine tests values will be compared
Constancy tests
  Daily tests
  Weekly tests
  Quarterly tests
  Yearly tests

## Slide 14: PET/CT scanner QC
Daily Checks
  Blank scan
  CT calibration
Weekly
  Registration
  Spatial resolution
Quarterly
  Distortion
  Sensitivity
  Set up and normalization
  CT, HU-physical density calibration
Annually
  - CT QC

## Slide 15: Daily QC
Blank scan using 68Ga line source
  looking for block malfunction
  5 min acquisition time
CT warm up and calibration
CT number accuracy (water)
CT image noise
Good blank scan
Bad block

## Slide 16: Weekly QC
Single gain update
High count rate blank scan
  similar to daily check, but typically takes 30 minutes
Image registration (PET and CT)
Spatial resolution
(mainly CT)
Notes: Singles Gain Uptake - calibrate the gain of the detector modules

## Slide 17: Quarterly QC
Distortion (PET and CT)
Well counter calibration
  Ensures accurate SUV data
2D and 3D calibration
Set up
  PM tube gains adjusted to give similar efficiencies (Hardware correction)
Normalization
  Residual variations in sensitivity are corrected using a sensitivity map (software correction)
HU - calibration
Notes: 2D and 3D calibration - calibration of the individual slices since each plane is different

## Slide 18: Annual QC
CT number accuracy
CT number uniformity
CTDI
Irradiated slice thickness (dose profile)
High contrast spatial resolution
Imaged slice thickness
Light and scan plane alignment accuracy
Table top incrementation
Couch travel accuracy (spiral scan)
ImPACT, IPEM 77

## Slide 19: Example Gamma Camera PET/CT
Weekly
  High energy uniformity correction
Periodically
  Fusion of FDG and CT image

## Slide 20: Standards from the industry for quality assurance
Performance Measurements of Positron Emission Tomographs, NEMA Standards Publication No. NU2, National Electrical Manufacturers Association (NEMA), Washington, DC, 2001(spatial resolution, sensitivity, count rate performance, accuracy of count losses, and random coincidence correction and image quality)
International Standard: Radionuclide imaging devices- Characteristics and test conditions - Part I: Positron Emission Tomographs, International Electrotechnical Commission (IEC), 61676-1, Geneva, Switzerland, 1998

## Slide 21: Maintenance
Documented systems and procedures
Equipment
Contamination monitors
Calibrators
Personal monitors

## Slide 22: SUMMARY OF QUALITY CONTROL
QC on 18F-FDG must be performed with regard to both radiochemical and chemical purity
Because PET scans will require CT images for attenuation correction, QC must be performed for PET alone, and for CT alone, and for the fusion of the two separate imaging modalities
In PET/CT, it is critically important to verify by established QC techniques the correct registration between the PET and CT scans
IAEA publication QC of PET-CT systems (in press)
Nichols K, et al, J Nucl Cardiol 2006;13:e25-e41