## Slide 1: L 7
PERSONAL & WORK PLACE MONITORING

## Slide 2: Answer True or False
It is important that PET/CT clinicians diligently use their personal monitoring devices in order to avoid excessive radiation dose
The choice of the particular type of monitor to be used is determined by the type of radiation, its energy, and whether it is more likely to involve high or low amounts of activity
In the event of skin contamination it is important to remove activity but not to break the skin in the attempt to do so

## Slide 3: Objective
To consider both personal monitoring and monitoring of the workplace: type of monitors and where, who and when to monitor, decontamination procedures

## Slide 4: Content
Personal monitoring
Type of monitors
Work place monitoring
Decontamination procedures

## Slide 5: Monitoring
Personal
(effective dose, extremity dose & contamination)
Workplace
(external dose rate & contamination)

## Slide 6
7.1 Personal Monitoring

## Slide 7: Individual Monitoring (BSS)
I.33. For any worker who is normally employed in a controlled area, or who occasionally works in a controlled area and may receive significant occupational exposure, individual monitoring shall be undertaken where appropriate, adequate and feasible……….
I.34. For any worker who is regularly employed in a supervised area or who enters a controlled area only occasionally, individual monitoring shall not be required but the occupational exposure of the worker shall be assessed. This assessment shall be on the basis of the results of monitoring of the workplace or individual.

## Slide 8: Individual Monitoring (BSS) cont.
I.35. The nature, frequency and precision of individual monitoring shall be determined with consideration of the magnitude and possible fluctuations of exposure levels and the likelihood and magnitude of potential exposures.
I.36. Employers shall ensure that workers who may be exposed to radioactive contamination, including workers who use protective respiratory equipment, be identified and shall arrange for appropriate monitoring to the extent necessary to demonstrate the effectiveness of the protection provided and to assess the intake of radioactive substances or the committed doses, as appropriate.

## Slide 9: Personal Monitoring Instruments

## Slide 10: Individual Monitoring
Control of radiation exposure
Identifies high doses
Assessment of working practices

## Slide 11: Personnel Monitoring
TLD film badges - monthly record
Electronic monitor (with alarm) for immediate record. This helps to remind operator to keep a distance from the radioactive patient
Finger stalls/rings for extremity monitoring should be adequate in terms of envelope thickness, calibration & to measure dose from electrons
Contamination monitoring

## Slide 12: Contamination Monitoring
Monitor before leaving unit
Monitor hands (especially finger tips and nails) after handling unsealed radionuclides
Monitor if contamination suspected

## Slide 13: Who Should be Monitored?
Those who are working with radioactive substances, including:
  preparation and administration of radiopharmaceuticals
  those who are performing patient examinations
  those who are performing quality control of the equipment such as:
      PET/CT clinicians
      PET/CT technologists
      Medical physicists
      Radiopharmacists

## Slide 14: Suggested Investigation Levels
Site | Investigation level for a Four-Week Period (mSv) | Investigation level for a Thirteen-Week Period (mSv)
Body | 0.5 | 2
Eye | 5 | 20
Individual organs/ extremities | 15 | 50

## Slide 15: Individual Monitoring (BSS)
A formal investigation is required by the Regulatory Authority whenever:
(a) the individual annual effective dose exceeds the investigation
level
(b) any of the operational parameters subject to periodic quality
control are out of the normal range established for operational
conditions
(c) any severe accident or error takes place
(d) any other event or unusual circumstance that causes, or has
the potential to cause, a dose in excess of the regulatory
limits.

## Slide 16: Investigations and Follow-up (BSS)
“IV.18. Registrants and licensees shall conduct formal investigations as specified by the Regulatory Authority if:
(a) a quantity or operating parameter related to protection or safety exceeds an investigation level or is outside the stipulated range of operating conditions; or
(b) any equipment failure, accident, error, mishap or other unusual event or circumstance occurs which has the potential for causing a quantity to exceed any relevant limit or operating restriction.
IV.19. The investigation shall be conducted as soon as possible after the event and a written report produced on its cause, with a verification or determination of any doses received or committed and recommendations for preventing the recurrence of similar events.
IV.20. A summary report of any formal investigation relating to events prescribed by the Regulatory Authority, including exposures greater than a dose limit, shall be communicated to the Regulatory Authority as soon as possible and to other parties as appropriate.”

## Slide 17
7.2 Types of Environmental Monitors

## Slide 18: Environmental Monitoring Instruments
Ionization chambers
Proportional counters
GM-tubes
Scintillation detectors

## Slide 19: Monitoring Instruments
Count rate meters for contamination detection and measurement
Dose rate meters for dose rate measurements

## Slide 20: Choice of Monitoring Instrument…
… depends on...
  High or low levels?
  Particles or photons?
  Energy of photons?
  Required accuracy?

## Slide 21: Choice of Detector Type…
… may be important for required sensitivity
… otherwise NOT so important !
The design of the instrument to fulfill the requirements is more important!

## Slide 22: Monitoring Positron Emitters
Scintillation monitors are not suitable for positron detection, however:
  can be used to detect 511 keV gamma rays

## Slide 23: Contamination Detectors
All with an appropriate design!
β - emitters | GM-tube Proportional counter
 - emitter (< 50 keV) | GM-tube Proportional counter Nal(TI) scintillation detector
 - emitter (high energy) | Proportional counter Nal(TI) scintillation detector

## Slide 24: Dose-rate Meters
 - emitter Ionization chamber
Scintillation detector
Often designed to meet the requirements to measure one of the operational dose quantities defined in ICRU 47

## Slide 25
7.3 Workplace Monitoring

## Slide 26: Workplace Monitoring (BSS)
“I.38. The nature and frequency of monitoring of workplaces
shall:
(a) be sufficient to enable:
(i) evaluation of the radiological conditions in all workplaces;
(ii) exposure assessment in controlled areas and supervised
areas; and
(iii) review of the classification of controlled and supervised
areas; and
(b) depend on the levels of ambient dose equivalent and activity concentration, including their expected fluctuations and the likelihood and magnitude of potential exposures. ”

## Slide 27: Monitoring Instruments - Dose-rate Meters

## Slide 28: Monitoring Instruments - Contamination Monitors
Notes: Proportional counters for betas. Scintillation detectors to detect 511 keV photons

## Slide 29: Making Measurements
Switch on before entering radiation area
Move monitor slowly
Change scale if necessary
Check batteries

## Slide 30: Workplace Monitoring Procedures
Identify the quantity to be measured
Specify location and frequency
Specify procedures
Identify reference levels
Specify record keeping
Specify procedure if contamination found

## Slide 31: Workplace Monitoring Procedures
Monitor at end of each day
Record (even if background)
Contamination monitor must be checked on annual basis

## Slide 32: Example Monitoring Worksheet
Action levels typically = 100 cps, but depends on monitor

## Slide 33: Contamination Monitoring
Where to monitor?
Injection area (especially any trolley)
Syringe loading areas
Patient toilet
Waste bins (contaminated items can be placed by mistake)
Floor in areas where unsealed sources are used

## Slide 34: Derived Limits for Surface Contamination
18F in class B
Radio-toxicity Class | Area |  | 
 | Controlled (Bq/cm2) | Supervised (Bq/cm2) | Body (Bq/cm2)
A | 30 | 3 | 3
B | 300 | 30 | 30
C | 3000 | 300 | 300

## Slide 35
7.4 Decontamination Procedures

## Slide 36: Contamination
Skin dose rate for 1 kBq 18F in 0.05 ml droplet is 0.8 mSvh-1

## Slide 37: To Minimize Contamination Risks
Adopt clean operating conditions
Adopt good laboratory practices - do not eat, drink, smoke etc…
Use protective gloves and clothing

## Slide 38: General Decontamination Procedures
Use adsorbent paper on wet spill or wet absorbent
paper on dry spill
Repetitively swab the area inwards towards the center
of the spill
Place contaminated paper in a plastic bag or container
Monitor the area
Repeat all procedures until the exposure rate is below
given limits
If the decontamination is not successful, mark the
contaminated area and classify the room as a controlled
area (if not already done) until the contamination is
completely removed

## Slide 39: Decontamination of Skin
If contamination of the skin occurs:
Wash the area thoroughly using mild soap and tepid (not hot) water
Particular care should be paid to cleaning under the fingernails
If this does not bring the contamination to an acceptably low level the procedure should be repeated using a decontaminating detergent
Scrub with a nail brush but take care not to break the skin

## Slide 40: Small Amounts of Radioactive Spills
Use protective clothing and disposable gloves
Quickly blot the spill with an absorbent pad to keep it from spreading
A plastic bag to hold contaminated items shall be available as well as some damp paper towels
Remove the pad from the spill
Wipe with a towel from the edge of the contaminated area toward the centre
Dry the area and monitor
Continue the cycle of cleaning and monitoring until the area is below the contamination limits to indicate that the spill is cleaned
The procedures should be practiced!

## Slide 41: Large Amounts of Radioactive Spills
The procedures should be practiced!
The Radiation Protection Officer should immediately be informed and directly supervise the clean-up
Absorbent pads may be thrown over the spill to prevent further spread of contamination
All people not involved in the spill should leave the area immediately
All people involved in the spill should be monitored for contamination when leaving the room
If clothing is contaminated it should be removed and placed in a plastic bag labeled ’RADIOACTIVE’
If contamination of skin occurs, the area should immediately be washed
If contamination of eye occurs, flush with large quantities of water

## Slide 42: Emergency Kit
Should be kept readily available for use in an emergency. It may include the following:
  Protective clothing, e.g. overshoes, gloves
  Decontamination materials for the affected areas including absorbent materials for wiping up spills
  Decontamination materials for persons
  Warning notices
  Portable monitoring equipment
  Bags for waste, tape, labels, pencils

## Slide 43: SUMMARY OF PERSONAL & WORK PLACE MONITORING
It is important that PET/CT clinicians, technologists and medical physicists diligently use their ring badges, whole body film badges and/or TLD badges to avoid excessive radiation dose
The choice of the particular type of monitor to be used (ionization chambers, proportional chambers, GM-tubes or scintillation detectors) should be determined by the type of radiation, its energy, and whether it is more likely to involve high or low amounts of activity
Decontamination procedures are aided substantially by the diligent routine use of protective gloves and clothing and absorbent disposable pads in areas where spills of radioactive liquids are likely to occur, such as at injection tables. In the event of skin contamination, it is important to remove activity but not to break the skin in the attempt to do so