## Slide 1: L 2
PET/CT TECHNOLOGY

## Slide 2: Answer True or False
Cyclotrons accelerate protons to strike 18O, thereby producing a neutron and the positron emitter 18F
PET scanners work by detecting the amount of gamma rays originated as a result of annihilation positrons and transmitted through the body of the patient at different angles from internally located cyclotron-generated positron sources
CT scanners work by detecting the amount of X rays that are generated by an external X ray tube and transmitted through the body of the patient at different angles

## Slide 3: Objective
To become familiar with the basic PET/CT technology including cyclotron, PET scanners, CT scanners and the merging of the two technologies into PET/CT

## Slide 4
Cyclotrons
PET scanners
CT scanners
PET/CT scanners
Content

## Slide 5
2.1 Cyclotrons

## Slide 6: Cyclotrons

## Slide 7: Cyclotrons
  Self-shielded or
  in a vault
  Applications:
  all PET radioisotopes:
  F18-, C11, N13, O15 and 18F2
  ‘new’ PET radioisotopes:
  I124, I123, Cu64, Y86, Br76 …

## Slide 8: Cyclotrons
CLASSIFIED BY:
  Particles
    Single/Dual
    Proton/Deuteron
  Energy
    7 to 18 or even 70 MeV
  Bombardment capabilities
    Single/Dual beam
  Number of Targets
    Quantity of radioactivity
    Chemical form

## Slide 9
Dees
Beam extractor
Magnetic coil
Target
Ion
Source
Notes: Magnetic dees keep ions in a circular path

## Slide 10: Manufacture of 11C
Proton is accelerated
Strikes 14N target
Merges with 14N
Alpha particle is ejected

4
2
11
6
1
1
14
7



C
p
N

## Slide 11: Manufacture of 18F
Proton is accelerated
Strikes 18O target
Merges with 18O
Neutron ejected
n
F
p
O
1
1
18
9
1
1
18
8




## Slide 12: Manufacture of FDG
Bombardment of the target material with the ion beam yields 18F
Bombardment could typically be 2 hours (one half-life)
18F then sent to a chemistry module (synthesis module) to react with a number of reagents to produce fluorinated deoxyglucose
Synthesis module performs a number of steps such as heating, cooling, filtering, purifying, etc.
FDG synthesis typically adds another hour
Notes: Is it Soup yet?
At the end of bombardment of the target material with the ion source beam, we have NOT made FDG …only Fluorine 18 (18[F])
Bombardment could typically be 2 hours (one halflife)
The 18[F] is then sent to a chemistry module to be synthesized with a number of chemicals and reagents that are necessary to produce fluorinated 2-deoxyglucose.
The chemistry module (known in the past as the “black box”) or CPCU, operated by a computer program, will perform a number of steps such as heating, cooling, filtering, purifying, etc. that will allow the 18[F] to be tagged as Fludeoxyglucose
FDG synthesis typically adds another hour

## Slide 13
1
2
3
18F synthesis system

## Slide 14
FDG Module

## Slide 15
2.2 PET scanners

## Slide 16: Coincidence Detection
Detector
Detector

## Slide 17
E = mc²
= 9.11 x10-31kg x (3x108)² m/sec
= 8.2 x10-14 J
= 8.2 x10-14 J ÷ (1.6x10-19 J/eV)
= 511 keV
Notes:  - neutrino

## Slide 18: Detection of Emissions
PET radionuclides are positron emitters
PET can detect
  beta particles
  or Brehmsstrahlung
  or annihilation gammas
Brehmsstrahlung not considered significant
Most detection systems detect 511keV gammas

## Slide 19: Configurations
Full ring
Partial ring
  rotated continuously
Flat panel detectors
  reduced number of PM tubes
Gamma camera
  2 heads rotate through 180o (rarely used now)

## Slide 20: Scintillators
Na(Tl) I works well at 140 keV. Poor efficiency at 511 keV
BGO, LSO and LYSO are common scintillators used in PET scanners
Density
(g/cc)
Z
Decay
time
(ns)
Light
yield
(% NaI)
Atten
.
length
(mm)
Na(Tl)
I
3.67
51
230
100
30
BGO
7.13
75
300
15
11
LSO
7.4
66
47
75
12
GSO
6.7
59
43
22
15
Notes: BGO - Bismuth Gremanate
LSO - Lutetium Oxyorthosilicate
GSO - Gadolinium Oxyorthosilicate

## Slide 21: Scanner Detectors
Lightguide
PMT

## Slide 22: Full Ring System
Block detectors

## Slide 23: Randoms and Scatter
· Annihilation event
Gamma ray
----- Line of response

## Slide 24
Scatter
  Patient dependent
  Correction applied using CT data
Randoms
  Number of randoms can exceed ‘true’ events
  Correct by
    reducing coincidence window
    measuring randoms ( delayed coincidence window)

## Slide 25
Siemens
Randoms and scatter degrade image both qualitatively and quantitatively
trues
randoms
&
scatter
Typical coincidence image*
containing a high percentage
of randoms and scatter
trues
randoms
&
scatter
Same image with same number of
counts but a positive change in the
ratio of trues to randoms & scatter
Notes: 7
42

## Slide 26: 2D and 3D
2D
Intersliced septa
Low randoms and scatter
3D
Remove intersliced septa
High sensitivity (x10)
High randoms and scatter
Susceptible to ‘out of field’ activity
2D mode
3D mode

## Slide 27: Standard Uptake Value (SUV)
SUV = Activity in ROI (MBq) / vol (ml)
Injected activity (MBq)/patient weight (g)
Areas with higher than average uptake will have SUV’s >1.
Higher the SUV, greater the risk of disease
Compare SUVs to monitor therapy
Cannot be used as an absolute number
before chemotherapy
SUV = 17.2
chemotherapy day 7
SUV = 3.9
chemotherapy day 42
SUV = 1.8
ROI

## Slide 28: Gamma Camera PET
1” NaI crystal is scored
  12.5 mm deep
  5940 squares at 7x7 mm
  Reduce light scattering in the crystal
  Reflect light towards the PM-tubes
PMT
PMT
low energy
high energy
1”

## Slide 29
2.3 CT scanners

## Slide 30: Computed Tomography
Computed Tomography (CT) imaging provides high quality images which reproduce transverse cross sections of the body.
Tissues are therefore not superimposed on the image as they are in conventional projections
The technique offers improved low contrast resolution for better visualization of soft tissue, but with relatively high absorbed radiation dose

## Slide 31: Computed Tomography
CT uses a rotating X Ray tube, with the beam in the form of a thin slice (about 1 - 10 mm)
The “image” is a simple array of X Ray intensity, and many hundreds of these are used to make the CT image, which is a “slice” through the patient

## Slide 32: Conversion of  to CT number
Distribution of  values initially measured
 values are scaled to that of water to give the CT number

## Slide 33: A look inside a rotate/rotate CT
X Ray
Tube
Detector Array
and Collimator

## Slide 34: Helical (spiral) Scan Principle
If the X Ray tube can rotate constantly, the patient can then be moved continuously through the beam, making the examination much faster
Scanning Geometry
Continuous Data Acquisition and Table Feed
X Ray beam
Direction of patient
movement

## Slide 35: Helical CT Scanners
For helical scanners to work, the X Ray tube must rotate continuously
This is obviously not possible with a cable combining all electrical sources and signals
A “slip ring” is used to supply power and to collect the signals

## Slide 36: A Look Inside a Slip Ring CT
X Ray
Tube
Detector
Array
Slip Ring
Note:
how most
of the
electronics is
placed on
the rotating
gantry

## Slide 37: Multi Slice Scanners
Single axial slices replaced by 2 slice in 1990s
In 2006 2-, 4- and 8-slice scanners superseded by 16-slice and 64-slice scanners, with better z axis resolution and allowing gated cardiac imaging
True cone beam CT not yet a commercial reality

## Slide 38: Multislice CT

## Slide 39: Helical (spiral) CT
Spiral CT and Spiral multislice CT:
Volume acquisition may be preferred to serial CT
Advantages:
  dose saving:
    reduction of single scan repetition (shorter examination times)
    replacement of overlapped thin slices (high quality 3D display) by the reconstruction of one helical scan volume data
    use of pitch > 1
  no data missing as in the case of inter-slice interval
  shorter examination time
    to acquire data during a single breath-holding period avoiding respiratory disturbances
    disturbances due to involuntary movements such as peristalsis and cardiovascular action are reduced

## Slide 40: Pitch ratio of the distance the table travels per rotation to the x-ray beam width
Number rotations | 10 |  | 5 |  | 2.5
Slice thickness | 10 | 10 | 10 | 10 | 10
Table movement per rotation | 10 | 15 | 20 | 30 | 40
Pitch | 1 | 1.5 | 2 | 3 | 4
Dose | 10 | 7.5 | 5 | 3.33 | 2.5

## Slide 41: Pitchx Definition = beam pitch
Pitchx = Table travel per rotation
Slice width (or beam width)
10
15
= 1.5
10
20
= 2.0

## Slide 42: Pitchd Definition (multislice)
Pitchd = Table travel per rotation
detector width
2.5
15
= 6.0 !!
This definition is no longer used by manufacturers
Notes: Pitch with multslice scanners is somewhat counterintuitive

## Slide 43: State of the Art of CT in 2008
1/3 sec tube rotation time
10-30 sec whole body scans
0.4-0.6 mm isotropic spatial resolution
64-320 multi-detector slices
> 1000 mm scan range
3-20 mSv doses (mean = 10 mSv)

## Slide 44
2.4 PET/CT

## Slide 45: PET/CT
Accurate registration
CT data used for attenuation (and scatter) correction
Applications
Anatomical localization
Monitor response to therapy
Radiotherapy planning

## Slide 46: PET/CT Scanner
PET scanner
CT unit

## Slide 47: Attenuation of 511 keV gamma photons
Vast majority of interactions of gamma-rays with tissue occur via Compton scatter
Attenuation factor across chest may be as high as 50
Reduces visibility of deep lesions
Reduces quantitative accuracy

## Slide 48: Attenuation Correction
Radioactive sources
  Germanium-68 rod sources
  Caesium-137 point sources
X ray source
  Quicker to acquire than radioactive sources
  Lower noise than radioactive sources
  Higher patient dose
a)
b)
c)
a) 68Ge
b) 137Cs
c) CT
Notes: Rod sources were used with PET scanners without a CT facility

## Slide 49: Attenuation Correction
Attenuation map applied to the emission images during iterative reconstruction
Emission
Transmission
Corrected

## Slide 50: Attenuation Correction with CT
CT - 120 kV (effective mean energy 70keV)
But, attenuation maps are energy dependent, so…
…need to adjust map from CT kV to 511 keV

## Slide 51: PET/CT
CT
PET
Survey
scan
CT
Reconstruction algorithm
Attenuation correction
PET
Fused Image

## Slide 52: Scan Process
CT scanogram performed first
Full CT performed second
Patient moved further into scanner and PET scan acquired third

## Slide 53: Patient Timings / Workflow
Injection
0
60
mins
Survey scan& CT
PET scan
(2 to 3 mins /bed position)
50
Patient empties bladder
65
100
Patient gets dressed and rehydrates
Rest
In modern systems, the full scan is completed in less than 20 min

## Slide 54: SUMMARY OF PET/CT TECHNOLOGY
Cyclotrons are used for producing positron emitters by accelerating protons to strike 18O, thereby producing a neutron and the positron emitter 18F
PET scanners work by simultaneous detection of two 511 keV gamma rays
CT scanners work by detecting the amounts of X rays generated by an external X ray tube that is transmitted through the body of the patient at different angles
PET/CT scanners have a PET scanner immediately after a CT scanner for accurate registration of the PET scan with the CT scan, enabling attenuation correction of the PET scan by the CT scan and anatomical localization of areas of unusually high activity revealed by the PET scan