## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L15.2: Optimization of protection in radiography: Radioprotection aspects
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Introduction
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Optimization of patient radiation protection requires periodic evaluation of doses and image quality.
Operators of the X Ray system should be aware of the interdependence between technical factors, dose, and image quality
Procedures should be established for each examination to ensure proper use of the equipment.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 3: Topics
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
  Practical rules to protect patients
  Generators and X Ray production related factors
  Imaging devices related factors (film, intensifying screens…)
  Examination procedures related factors (number of radiographic projections, technique settings..)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 4: Part 15.2: Optimization of protection in radiography
Topic 1: Practical rules to protect patients
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 5: Existing past images
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Before beginning an examination, it is advisable to compile the existing past images of similar examinations in or outside the institution, in order to minimize the number of radiological examinations for the patient.

## Slide 6: Periodic measurement of the entrance patient dose
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The periodic measurement (at least once a year) of the entrance patient dose (or similar quantity) and the comparison with the diagnostic reference levels will permit the detection of changes due to equipment or radiographers
When the entrance doses (or other dosimetric values) are clearly in excess of the diagnostic reference levels (or of those obtained previously ), it is necessary to investigate, and correct, the causes of the higher doses.

## Slide 7: Corrective actions
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
It is advisable to record corrective measures (in a log book) when higher doses are detected.
The log book should contain complete information as to the problem, dose increase, corrective action, and the date

## Slide 8: Source to skin distance and others
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
In general radiography (except dental) and fluoroscopy with X Ray mobile equipment, the source-to-skin distance should not be less than 30 cm
In radiography and fluoroscopy with fixed equipment, the skin-focus distance should not be lower than 45 cm
Fluoroscopy equipment without image intensifiers must be replaced, or upgraded with an image intensifier

## Slide 9: Practical rules to protect patients
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Equipment should be evaluated whenever it is suspected that patient doses have increased above the diagnostic reference levels.
This evaluation should include both image quality and patient dose

## Slide 10: Part 15.2: Optimization of protection in radiography
Topic 2: Generators and X Ray production related parameters
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 11: X-Ray Tubes
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
All x-ray equipment must conform to applicable standards of the International Electrotechnical Commission (IEC) and the ISO or to equivalent national standards
Adequate filtration (minimum 2.5 mm Al, in general radiology), significantly reduces the patient dose due to low energy X Rays which do not contribute to the image formation
For x-ray tubes equipped with removable aluminum filters (added filtration), it is essential to verify whether the filters are still in place after maintenance by measuring the half-value layer.

## Slide 12: X-Ray Generators and Tubes
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
In mammography, very low voltages are used and the filtration requirements are different (0.03 mm thickness of Molybdenum recommended, in equipment with anode of the same material).
Measurement of the half-value layer is necessary to confirm the appropriate filtration.

## Slide 13: X-Ray Generators and Tubes
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
For fluoroscopic examinations, a timer should be available allowing the measurement of the elapsed fluoroscopic exposure time, with audible indications at every five minutes.
Fluoroscopy should be controlled with a "dead man" switch.

## Slide 14: Mobile X-Ray Equipment
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
When using mobile radiographic equipment in intensive care units or in-patient rooms, high instantaneous electric power supply is needed.
Inadequate power will result in poor quality radiographs.

## Slide 15: Mobile X-Ray Equipment
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
For the same reason, in order to avoid retakes, it is important to assure that the battery is fully charged
For units connected to the mains, one must be sure that the required power is actually supplied.

## Slide 16: Part 15.2: Optimization of protection in radiography
Topic 3: Imaging devices related parameters
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 17: Screen-film combinations
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The use of the appropriate screen-film combination, the "fastest" compatible with the type of image needed, is recommended to assure lowest patient dose
Due to human errors, it is not advisable to use several screen-film combinations of different sensitivity classes in the same room,
An exception is the case when each combination has a different film format and the selection of radiological technique is made manually (no AEC)

## Slide 18: With reference to imaging devices (I)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The use of cassettes, grids, and tables of carbon fiber material, results in significant patient dose reductions
Intensifying screens with scratches or cassettes that do not provide the correct film-screen contact should not be used

## Slide 19: With reference to imaging devices: automatic exposure control
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Inappropriate selection of automatic exposure control settings might lead to images which are too light or too dark
Automatic exposure control devices should ALWAYS be evaluated, especially when the sensitivity of the screen-film combination has been changed.
The correct operation of the automatic exposure control device requires, for each projection, the selection of the chamber or detector closest to the area of interest, so that this area will have the appropriate density.

## Slide 20: With reference to fluoroscopic imaging devices (II)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The patient entrance air kerma rate should not exceed 50 mGy/min.
In modern image intensifiers, this value should be much lower depending on patient size and projection
The use of devices for storing and displaying the last image (last image hold) is recommended

## Slide 21: With reference to imaging devices (III)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
For both automatic and manual film processing, it is essential to change the chemicals according to the manufacturer instructions
The safelights in darkrooms do not have filters with endless life. Replacement is recommended whenever darkroom fog tests indicate a problem.
When changing to a more sensitive film it is essential to test for darkroom fog to assure the safelight filter and light bulb provide optimum results

## Slide 22: With reference to imaging devices (IV)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
It is important to have view boxes in areas with correct environmental light, enough brightness and uniformity over the surface and a high intensity light source (hot light). Periodic cleaning of the internal and external surfaces, and replacement of the fluorescent tubes are essential
Assure that the illuminated area is not larger than the film size, especially in mammography, i.e., mask to the exposed area of the film

## Slide 23: Part 15.2: Optimization of protection in radiography
Topic 4: Examination procedures related parameters
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 24: With reference to procedures (I)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
A technique chart should be placed beside each x-ray control panel for the various projections. For manual exposures, the techniques (kVp, mA, and time) should be specified as a function of body part thickness.
The body part of interest should be measured with a caliper when using manual exposures
For automatic exposure control systems, the technique chart should specify the control panel settings to be used for each projection.

## Slide 25: With reference to procedures (II)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
With any equipment (manually as well as automatically controlled), it is important to know which techniques should be selected to obtain a good image
When changes are introduced on any component in the imaging chain (generator or tube as well as other accessory devices such as film type, cassette, intensifying screen, etc.), an update of technique settings should be carried out.
A yearly check of the technique charts is essential

## Slide 26: With reference to procedures (III)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
It is advisable to use the highest kVp (and the lowest mAs) compatible with the image that one expects to obtain. In this way the patient dose will be minimized. Optimization should be used to find the proper balance between contrast and dose
Short exposure times should be used when imaging non-cooperative patients.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 27: With reference to procedures (IV)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Pediatric examinations should be carried out with three-phase or high frequency generators, capable of producing short exposure times, in order to avoid retakes due to motion blurring.
The radiographic techniques in use in each room should be compared on a regular basis with those recommended in published guidelines
Likewise, patient doses should be compared with the applicable Diagnostic Reference Levels.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 28: With reference to procedures (v)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Radiographic examinations must be prescribed by a properly accredited physician
When applicable, patient’s own medical record as well as medical indications should be available (WHO, EC)
It may be appropriate to modify the examination, e.g., modify or substitute with concurrence of the referring physician, in order to adopt the most appropriate strategy
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 29: With reference to procedures (VI)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
It is essential that the personnel operating the equipment be properly trained and accredited
Training criteria shall be specified or be subject to approval, as appropriate, by the Regulatory Authority in consultation with relevant professional bodies
Personnel should inform the patient on the correct positioning and immobilization as well as on other aspects of the examination (suspended respiration, deep inspiration, etc.)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 30: With reference to procedures (VII)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
It is important to assure that the radiological examination is "justified", taking into account the benefits and risks of available alternative techniques that do not involve ionizing radiation
The patient should wear gonadal protectors, if gonads are exposed, assuming that it does not interfere with the image
In case of possible foetal exposure it is advisable to adapt the examination procedure of pregnant women, together with the radiation protection strategy
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 31: With reference to procedures (VIII)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
In order to avoid unwanted irradiation of the fetus, it is recommended to post warnings, both at the X Ray room entrance and in the waiting ward such as:
"IF YOU THINK THAT YOU ARE PREGNANT, PLEASE TELL THE RADIOGRAPHER (RADIOLOGICAL TECHNOLOGIST) OR THE RADIOLOGIST, BEFORE THE X RAY EXAMINATION IS PERFORMED".
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 32: With reference to procedures (IX)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Female patients should be asked about the possibility of being pregnant, even the young pediatric patient. If they are pregnant, appropriate measures should be taken
When a pregnant patient undergoes a radiological examination (of the abdomen), it is advisable to evaluate the expected fetal dose.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 33: With reference to procedures (X)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The pregnant patient or worker has a right to know the magnitude and type of potential radiation effects that might result from in utero exposure
Communication should be related to the level of risk. Verbal communication may be adequate for low dose procedures
If fetal doses are above 1 mGy, a more detailed explanation may be necessary
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 34: Approximate fetal doses from conventional X Ray examinations (data from the UK 1998)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
 | Mean (mGy) | Maximum (mGy)
Abdomen | 1.4 | 4.2
Chest | <0.01 | <0.01
IVU or lumbar spine | 1.7 | 10
Pelvis | 1.1 | 4
Skull or thoracic spine | <0.01 | <0.01
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 35: Approximate fetal doses from fluoroscopic and computed tomography procedures (data from the U.K. 1998)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
 | Mean (mGy) | Maximum (mGy)
Barium meal (UGI) | 1.1 | 5.8
Barium enema | 6.8 | 24
Head CT | <0.05 | <0.05
Chest CT | 0.06 | 1.0
Abdomen CT | 8.0 | 49
Pelvis CT | 25 | 80
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 36: With reference to procedures (XI)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The most appropriate projection should be adopted, when the diagnostic information is not compromised.
  For pregnant women, PA abdominal projections are preferable to minimize uterus dose
  For skull examinations, eye lenses are better protected in PA projection
  PA projections should be used for scoliosis imaging to minimize breast dose
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 37: With reference to procedures (XII)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The smallest film and cassette size compatible with the expected image must be used together with automatic collimation. Otherwise the patient would be over-irradiated, by receiving radiation over a larger volume. Irradiating a smaller volume also minimizes the amount of scattered radiation and improves image contrast.
When using equipment without automatic X Ray beam collimation, it should be verified that the radiation field is reduced up to the smallest size compatible with the required image, even in fluoroscopy applications (the unit usually will allow for radiation field reductions both in radiography as in fluoroscopy).
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 38: With reference to procedures (XIII)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The use of the anti-scatter grid improves the image quality, but ALWAYS increases patient doses. It is advisable to evaluate whether the grid is actually necessary in equipment where its use is optional according to procedure or patient characteristics. In that case, one should check its location (rids result in an increase in patient skin dose byfactors of 2 to 3 times)
When the grid is of a focused type, it is important to confirm that the focus-film distance has been selected within the correct range
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 39: With reference to procedures (XIV)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
When a change of the usual technique is needed in order to improve or maintain the image quality, it is advisable to check the performance of the complete imaging chain. Usually, the change implies an increase in patient dose.
The patient should be visible from the operation control panel.
When possible fluoroscopy should be used in “intermittent mode”, irradiating the patient only when necessary.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 40: With reference to procedures (XV)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
The use of fluoroscopy for centering the radiation field is not appropriate as it results in patient doses many times higher than from the radiographic exposure.
Whenever possible a compression device (e.g., mammography), should be used, since it reduces dose while improving image quality.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
<number>

## Slide 41: With reference to procedures (XVI)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
In general radiology, the source-to-skin distance should be as long as possible, and the patient-image device (detector) distance as close to the patient as possible
In fluoroscopy the source-to-skin distance is usually determined by the collimator, and the patient-intensifier distance should be as short as possible.

## Slide 42: With reference to procedures (XVII)
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
CT examinations should be done with the minimum possible number of slices while obtaining the necessary information. Increasing the number of slices results in a higher patient dose over a larger volume.
In general, the radiological examination should be performed with the minimum number of images.
Fluoroscopy times should be minimized.

## Slide 43: Summary
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Practical rules to protect patients include periodic assessment of dose and image quality, with corrective action as needed
The practitioner should be aware of the influence on dose and image quality of technical parameters (field size, grid type, kV, type of projection…)
Procedures (number of images and technique factors) have to be established for each examination.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Let’s summarize the main subjects we did cover in this session. (List the main subjects covered and stress again the important features of the session)

## Slide 44: Where to Get More Information
15.2: Optimization of protection in radiography: Radioprotection aspects
<number>
Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards, Revision of IAEA Safety Series No. 115, IAEA, Vienna Austria, 2011
A practical guide on radiological protection and quality assurance in diagnostic radiology. CE, Value Programme, 1996. Vañó E, Gonzalez L, Maccia C, Padovani R. Edited by Cátedra de Física Médica, Facultad de Medicina, Universidad Complutense de Madrid, 28040 Madrid, Spain.
Radiological protection of the worker in medicine and dentistry. ICRP Publication 57, Pergamon Press, 1989.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources