## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L 22: Optimization of Protection in Dental Radiology
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Introduction
22: Optimization of Protection in Dental Radiology
<number>
Dental radiology makes use of specific types of imaging equipment. Frequent exposures, though each with relatively low dose, involve a risk for the practitioner and for the patient
Background: general principles of x-ray diagnostic imaging

## Slide 3: Topics
22: Optimization of Protection in Dental Radiology
<number>
  Dental X-ray equipment
  Radiation protection in dental radiology
  Quality control for dental equipment

## Slide 4: Overview
22: Optimization of Protection in Dental Radiology
<number>
To be able to apply the principle of radiation protection to dental radiology system including design and Quality Control.

## Slide 5
Part 22: Optimization of protection in dental radiology
Topic 1: Dental x-ray equipment
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology

## Slide 6: Types of units
22: Optimization of Protection in Dental Radiology
<number>
“Intra-Oral” units
  Standard dental tube
  Uses an intra-oral image receptor and extra-oral x-ray tube
Panoramic (Orthopantomography, OPG)
Cephalometric (Ceph)

## Slide 7: Intra-Oral Dental X-Ray Equipment
22: Optimization of Protection in Dental Radiology
<number>

## Slide 8: Modern Dental X-Ray Unit
22: Optimization of Protection in Dental Radiology
<number>

## Slide 9: Panoramic X-Ray Equipment
22: Optimization of Protection in Dental Radiology
<number>

## Slide 10: Cephalometric X-Ray Equipment
22: Optimization of Protection in Dental Radiology
<number>

## Slide 11: X-Ray Tube
22: Optimization of Protection in Dental Radiology
<number>
stationary Anode
avoid overheating
tube duty cycle:
  typical: 1:30 intaroral
  1:10 OPG
  420 mAs/hr intraoral

## Slide 12: Tube Head
22: Optimization of Protection in Dental Radiology
<number>

## Slide 13: Generator Circuit
22: Optimization of Protection in Dental Radiology
<number>

## Slide 14: Generators & Pre-Heat
22: Optimization of Protection in Dental Radiology
<number>
Medium frequency - stable waveform
Single phase (SP) - pulsed
Pre-Heat: separate circuit for heating filament
Single Phase units without a pre-heat circuit
  initial pulses of variable kV

## Slide 15: Collimator
22: Optimization of Protection in Dental Radiology
<number>
1. Lead Collimator
with central hole
2. Spacer Tube (cone, position indicating device or PID)

## Slide 16: Cones
22: Optimization of Protection in Dental Radiology
<number>
Good Bad Bad

## Slide 17: Cone (PID) Length and Collimation
Three cone (source-to-skin) distances– 8”, 12”, and 16”
  Longer distance improves image sharpness, reduces dose
Circular vs rectangular collimation
  Rectangular– smaller field irradiated
    Results in lower dose
    Less scattered radiation
    Increased contrast
    But more difficult to position
22: Optimization of Protection in Dental Radiology
<number>

## Slide 18: Cephalometric Holder
22: Optimization of Protection in Dental Radiology
<number>

## Slide 19: Intra-Oral Dental X-Ray Equipment (technical data)
22: Optimization of Protection in Dental Radiology
<number>

## Slide 20: Panoramic X-Ray Equipment (technical data)
22: Optimization of Protection in Dental Radiology
<number>
Focal spot 0.5 mm
kV 60 - 80 kV in 2 kV steps
mA 4 - 10 mA steps 4, 5, 6, 8, 10
Exposure time 12 s (standard projections) 0.16 - 3.2 s (cephalometric projections)
Flat panoramic cassette 15x30 cm (Lanex Regular screens))

## Slide 21: Image Receptors in Dental Radiology
22: Optimization of Protection in Dental Radiology
<number>
Small films (2 x 3 or 3 x 4 cm) in light-tight envelopes (no screen)
Digital intraoral sensors - compared with category F film, the radiation dose is reduced by 60%.
Intraoral Radiology
Panoramic Radiology and Cephalometry
Screen-film combination
Digital sensors - compared with screen-film sensitivity class 200, the radiation dose is reduced by 50-70%.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 22: Dental Radiology Film Types
22: Optimization of Protection in Dental Radiology
<number>
Sensitivity class D
  Good spatial resolution
  Typical delivered dose: about 0.5 mGy
  Typical exposure times: 0.3 - 0.7 s
Sensitivity class E, E-F, or F
  Good spatial resolution
  Typical delivered dose: about 0.25 mGy
  Typical exposure times: 0.1 - 0.3 s
Image quality of D, E, E-F, F films similar
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Uso di film sensibili. Sono attualmente disponibili film di due classi di sensibilità: i film di classe D devono essere esposti a una dose di circa 0.54 mGy, mentre quelli di classe E necessitano di una dose di circa 0.27 mGy (50% in meno). I film di classe E hanno una risoluzione leggermente inferiore a quelli di classe D, ma nella maggior parte delle indicazioni odontoiatriche non vi e’ differenza nell’informazione ottenuta con i due tipi di film. L’uso dei film di classe E richiede il controllo e la messa a punto del temporizzatore sui valori brevi del tempo di esposizione (0.1-0.3 s).

## Slide 23: Part 22: Optimization of Protection in Dental Radiology
Topic 2: Radiation Protection in Dental Radiology
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology

## Slide 24: Radiation Protection in Dental Radiology
22: Optimization of Protection in Dental Radiology
<number>
Facts
Very frequent examination (about 25% of all the radiological examinations)
Delivered doses may differ of a factor 2 or 10 or more (entrance doses between 0.5 and 150 mGy)
Full mouth examination requires 20 exposures
Image Quality often very low due to poor techniques and processing
Organs at risk: parathyroid, thyroid, larynx, parotid glands

## Slide 25: Radiation Protection in Dental Radiology
22: Optimization of Protection in Dental Radiology
<number>
Films must be processed using appropriate development time for the specific developer temperature.
Replenish chemicals as recommended by film manufacturer
Do not adjust development time by viewing the film
Technical hints to reduce patient doses
Quality Control of Film Processing

## Slide 26: Radiation Protection in Dental Radiology
22: Optimization of Protection in Dental Radiology
<number>
Technical hints to reduce patient doses
Lead apron and collar
Useful when the path of primary beam intercepts the protected organs (downward bite-wing projection).

## Slide 27: Radiation Protection in Dental Radiology
22: Optimization of Protection in Dental Radiology
<number>
Panoramic examination
Image quality not as good as in intra-oral films– serves different purpose
Important global information
Relatively low dose (one panoramic examination  0.50 mGy)

## Slide 28: Part 22: Optimization of Protection in Dental Radiology
Topic 3: Quality Control for Dental Equipment
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology

## Slide 29: Why Dental QC ?
22: Optimization of Protection in Dental Radiology
<number>
Widespread use of dental units
Lack of QC on most units
Dental practitioners working in the primary health care sector do not have the continuous medical physics support available in a hospital-based diagnostic imaging department
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 30: What Tests ?
22: Optimization of Protection in Dental Radiology
<number>
Processor QC– most critical
Collimation
Dose
Exposure Time
Half-Value Layer
Kilovoltage (kVp)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 31: Quality Control for Dental Equipment
22: Optimization of Protection in Dental Radiology
<number>
The recommended tests are consequently divided into:
those simple tests which can be performed by dental practice staff
those more complex tests which can be carried out by medical physicists.

## Slide 32: Quality Control for Dental Equipment
22: Optimization of Protection in Dental Radiology
<number>
Tests which can be performed by dental practice staff
Physical parameter | Tolerances | Frequency
Image Quality | ±10% reference values | Quarterly
Developer temperature and condition of processing solutions | Specified by the film manufacturer | Every time processing solutions are used
Processing | Base+Fog Mid Density Density difference | Every time processing solutions are used
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 33: Quality Control for Dental Equipment
22: Optimization of Protection in Dental Radiology
<number>
Tests performed by medical physicists
Physical parameter | Tolerances | Frequency
Tube voltage | >50 kV and error <10% | 3 yearly
Beam size | <60 mm diameter (intra-oral) <150 x 10 mm at cassette (panoramic) | 3 yearly
Dose at cone tip | 70 kV: <2.5 mGy (E speed film) | 1-3 yearly
Dose-width product for panoramic film | <75 mGy mm | 1-3 yearly
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 34: Dental QC Methods
22: Optimization of Protection in Dental Radiology
<number>
Unit
Intra-Oral Receptors
(I/O)
Cephalometric
(Ceph)
Panoramic
(OPG )
Test Method
  as for Radiology QC
  as for Radiology QC
  where possible:
  immobilise unit
  remove slit collimator
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 35: Test Equipment
22: Optimization of Protection in Dental Radiology
<number>
kVp meter
  measure kVp average
  Programmable delay ~100 ms
  Range: 50 to 120 kV
Aluminum filters
  4 x 1mm
  Grade 1100
Dosimeter
  small & large volume chambers
  2 mm wide detector for OPG
Timer
  triggering at 75% peak kV
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 36: Collimation
22: Optimization of Protection in Dental Radiology
<number>
Expose film or fluorescent screen
Measure x-ray field image
Unit | Film position | Limits of X Ray field
Standard | End of Cone | 60 mm diameter or as specified
OPG | Slit, in film cassette (2 films) | Equal slit images Vertical image ± film height
Ceph | At cassette | ±1 % of FFD on all margins
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 37: Collimator Light Intensity
22: Optimization of Protection in Dental Radiology
<number>
Ceph units:
Place external detector 1 m from focus
Measure illuminance in lux
Read each quadrant
Limit: >100 lux at 1m
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 38: “Dead man” Switch
22: Optimization of Protection in Dental Radiology
<number>
timer at 2 m from x-ray tube
set low kV, mA, long time
start exposure
release switch during exposure
Require exposure termination when switch is released. Check exposure time is less than set time
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 39: Kilovoltage
22: Optimization of Protection in Dental Radiology
<number>
Accuracy:
Set kVp meter to ~100 msec delay
Observe kVp waveform at 70 kV if poss.
Limit: measured kVp within 5% of set value
Reproducibility:
Take 5 repeat exposures
Limit: coefficient of variation ≤ 2%
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 40: Dose Evaluation
22: Optimization of Protection in Dental Radiology
<number>
Skin dose from Intraoral units:
place cone 10 mm from dosimeter
set adult bitewing technique factors
Should be (65-70 kVp):
  2-3 mGy for molar view
  < 5 mGy for any view
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 41: Output Reproducibility (1)
22: Optimization of Protection in Dental Radiology
<number>
Standard Intraoral units & Ceph units:
Dosimeter position:
  I-O units ~10 mm from cone
  Ceph units: 75 cm from focus
  or other recommended distance
Three repeat exposures
Limit: coefficient of variation ≤ 5%
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 42: Output Reproducibility (2)
22: Optimization of Protection in Dental Radiology
<number>
Optional Method for OPG units:
Align detector on film cassette slit
Measure dose rate
Take 3 repeat exposures
Limit: coefficient of variation ≤ 5%
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 43: Exposure Time Accuracy
22: Optimization of Protection in Dental Radiology
<number>
Standard Intraoral units & Ceph units:
Set timer to trigger at 75% peak kV
Test times in the normal working range
Limit: ≤10 % error for I-O units
≤ 5% error for all other units
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 44: Timer Reproducibility
22: Optimization of Protection in Dental Radiology
<number>
Standard Intraoral units & Ceph units
Place timer in beam
3 repeat exposures
Limit: coefficient of variation ≤5%
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 45: Half Value Layer (HVL)
22: Optimization of Protection in Dental Radiology
<number>
Standard Intraoral units:
Position cone facing down
Place dosimeter at ~ 40 cm from focus
Position Al filters near end of cone
Measure dose
  measure with no added filters
  with 2,3,4 mm Al added, then again with no filters
Plot on semi-log paper and find HVL
  Limit: HVL > 1.5 mm Al
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 46: HVL Set-up
22: Optimization of Protection in Dental Radiology
<number>

## Slide 47: HVL: Ceph & OPG Units
22: Optimization of Protection in Dental Radiology
<number>
Position Al filters on collimator
Cephalometric units
Position dosimeter at 75 cm from focus
OPG Units
Position dosimeter on film cassette slit
Measure dose rate, dose for fixed exposure time, or dose for full scan
NB Test kVp accuracy before measuring HVL
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 48: OPG Quality Control (kVp/HVL measurement)
22: Optimization of Protection in Dental Radiology
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 49: HVL- Minimum Values
22: Optimization of Protection in Dental Radiology
<number>
kVp HVL (mm Al)
Intraoral Ceph or OPG
60 1.5 1.8
70 1.5 2.1
80 2.3 2.3
90 2.5 2.5
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 50: Radiation Protection in Dental Radiology
22: Optimization of Protection in Dental Radiology
<number>
Control the time and temperature of the developing process.
Do not use oxidized chemicals
Regularly check processing with phantom
Quality Control of Film Processing
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Molti studi hanno evidenziato che l’esposizione del paziente odontoiatrico e’ spesso superiore al necessario per compensare processi di sviluppo impropri e che un gran numero di radiografie dentali risultano illeggibili per contrasto e densità insufficienti.
Controllo di qualità del processo di sviluppo. E’ questo il fattore essenziale nella radiologia odontoiatrica per ottenere immagini di buona qualità senza irradiare eccessivamente il paziente. Con semplici metodi possono essere tenuti sotto controllo i parametri essenziali del processo di sviluppo (tempo e temperatura dello sviluppo, stato di ossidazione del liquido di sviluppo). Cambiamenti di questi parametri hanno l’effetto di deteriorare l’immagine e possono indurre ad aumentare i tempi di esposizione per compensare le perdite di contrasto.
Non e’ pratica inusuale compensare con tempi di sviluppo piu’ brevi tempi di esposizione troppo lunghi per ottenere radiogrammi con una densità accettabile (sviluppo “a vista”). Questo grave errore procedurale puo’ aumentare la dose al paziente per radiogramma di 2-5 volte.

## Slide 51: Inexpensive and easy processor QC
22: Optimization of Protection in Dental Radiology
<number>
Simple and inexpensive device for processor QC
Sensitometer and densitometer not required
Expose film in device, process, and compare in same device
Cost approximately $40 (USD)
www.xrayqc.com

## Slide 52: Dental Phantom
22: Optimization of Protection in Dental Radiology
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 53: Dental Phantom
22: Optimization of Protection in Dental Radiology
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 54: Summary
22: Optimization of Protection in Dental Radiology
<number>
Technical characteristics of intraoral, panoramic, and cephalometric dental x-ray equipment are summarized, along with operating characteristics
Although doses are generally low, the high frequency of examinations requires radiation protection and quality control in dental radiology
Some tests are detailed for quality control of dental equipment.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 55: References (1)
22: Optimization of Protection in Dental Radiology
<number>
European guidelines on radiation protection in dental radiology. The safe use of radiographs in dental practice. Radiation Protection 136. 2004 http://ec.europa.eu/energy/nuclear/radioprotection/publication/doc/136_en.pdf
Quality assurance in dental film radiography. 2007. http://gar.carestreamdental.com/~/media/Files/GAR/N-416_Quality_Assurance_Brochure.ashx

## Slide 56: References (2)
Exposure and processing for dental film radiography http://www.carestreamdental.com/~/media/Files/FILM%20AND%20ANESTHETICS/Support/Exposure%20and%20Processing%20for%20Radiography.ashx
Guidance Notes for Dental Practitioners on the Safe Use of X-Ray Equipment. 2001 http://www.hpa.org.uk/web/HPAwebFile/HPAweb_C/1194947310610
22: Optimization of Protection in Dental Radiology
<number>