## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L 1: Overview of Radiation Protection in Diagnostic Radiology
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Introduction
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Persons are medically exposed as part of their diagnosis or treatment.
According to ICRP and BSS, the two basic principles of radiation protection justification and optimization
Dose limits are not applicable, but Diagnostic Reference Levels (DRLs) apply to patient dose levels
Investigation of doses that exceed the DRLs is strongly recommended
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 3: Topics
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Definition of medical exposure
Justification
Optimization
Diagnostic Reference Levels - practical aspects
DRls and effective doses
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 4: Overview
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
To become familiar with the BSS Safety Standards requirement for medical exposure: justification, optimization, diagnostic reference levels, and investigation of exposure.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 5: Part 1: Overview of Radiation Protection in Diagnostic Radiology
Topic 1: Definition of medical exposure
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 6
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Mr. Sharp, I am given to understand that 2 CT examinations performed on me have given me 25 mSv whereas 20 mSv is the safe dose. I want to file legal suit against the doctor. What do you feel ??

## Slide 7
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Medical exposure
versus
occupational

## Slide 8
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
My resident doctor has got 12 mSv in her last badge report as she was wearing the badge while getting her barium study. She wants off from radiation work.
?????

## Slide 9
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
While holding his child in diagnostic examination Mr. Joseph got 2 mSv.
As a member of the public with 1 mSv dose limit, he can not get any additional radiation dose this year.
???????

## Slide 10
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Dose constraints for
Comforters
under a category of
Medical exposure

## Slide 11: Three types of exposure
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Medical Exposure principally the exposure of persons as part of their diagnosis or treatment
Occupational Exposure exposure incurred at work, and practically as a result of work
Public Exposure including all other exposures

## Slide 12: Medical exposure
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Medical Exposure
  Exposure of persons as part of their diagnosis or treatment
  Exposures (other than occupational) incurred knowingly and willingly by individuals such as family and close friends helping either in hospital or at home in the support and comfort of patients
  Exposures incurred by volunteers as part of a program of biomedical research

## Slide 13: Framework of radiological protection for medical exposure
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Justification
Optimization
The use of doses limits is NOT APPLICABLE
  Dose constraints and Diagnostic Reference Levels ARE RECOMMENDED

## Slide 14: Part 1: Overview of Radiation Protection in Diagnostic Radiology
Topic 2: Justification
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 15: The justification of a practice
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
The decision to adopt or continue any human activity involves a review of benefits and disadvantages of the possible options, e.g., choosing between the use of X Rays or ultrasound
Often, the radiation detriment will be only a small part of the total detriment
Most of the assessments needed for the justification of a practice are made on the basis of experience, professional judgement, and common sense

## Slide 16: Three levels of justification
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
General level: The use of radiation in medicine is accepted as doing more good than harm
Generic level: specific procedure with a specific objective: chest radiographs for patients showing relevant symptoms
Third level: the application of the procedure to an individual patient

## Slide 17: Generic justification (I)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
It is a matter for national professional bodies, sometimes in conjunction with national regulatory authorities
The exposures to staff (occupational) and to members of the public should be taken into account
The possibility of accidental or unintended exposures (potential exposure) should also be considered
The decisions should be reviewed from time to time as new information becomes available

## Slide 18: Generic justification (II)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
The resources in a country or region should be considered, e.g., fluoroscopy for chest imaging could be the procedure chosen instead of radiography for economical reasons
The justification of diagnostic exposures for which the benefit to the patient is not the primary objective needs special consideration, e.g., radiography for insurance purposes

## Slide 19: Generic justification (III)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Any radiological examination for occupational, legal or health insurance purposes undertaken without reference to clinical indications is deemed to be not justified unless it is expected to provide useful information on the health of the individual examined or unless the specific type of examination is justified by those requesting it in consultation with relevant professional bodies.

## Slide 20: Justification for an individual patient (third level)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
To check that the required information is not already available
Once the procedure is generically justified, no additional justification is needed for simple diagnostic investigations
For complex procedures (such as CT, IR, etc) an individual justification should be taken into account by medical practitioner (radiologist, referral doctor..)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 21: Part 1: Overview of Radiation Protection in Diagnostic Radiology
Topic 3: Optimization
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 22: The optimization of protection (I)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Optimization is usually applied at two levels:
  The design and construction of equipment and installations
  Day to day radiological practice (procedures)
Reducing the patient dose may reduce the quantity as well as the quality of the information provided by the examination or may require important extra resources
Optimization means that doses should be “as low as reasonably achievable”, compatible with achieving the required image quality objectives

## Slide 23: The optimization of protection (II)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
There is a considerable scope for dose reductions in diagnostic radiology (ICRP 103)
Simple, low-cost measures are available for reducing doses without loss of diagnostic information (ICRP 103)
The optimization of protection in diagnostic radiology does not necessarily mean the reduction of doses to the patient, i.e., it may be necessary to increase some doses to obtain clinical image quality
Antiscatter grids improve the contrast of the image but increase the dose by a factor of 2-4

## Slide 24: Part 1: Overview of Radiation Protection in Diagnostic Radiology
Topic 4: Diagnostic Reference Levels—
practical aspects
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 25: Diagnostic Reference Levels (DRLs) for medical exposure (as defined by the BSS)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
A value of dose, dose rate or activity selected by professional bodies in consultation with the Regulatory Authority to indicate a level above which there should be a review by medical practitioners and medical physicists in order to determine whether or not the value is excessive, taking into account the particular circumstances and applying sound clinical judgement

## Slide 26: Diagnostic Reference Levels (DRLs) for medical exposure (as defined by the BSS)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
The Diagnostic Reference Levels are intended:
  to be a reasonable indication of doses for average sized patients
  to be established by relevant professional bodies in consultation with the Regulatory Authority
  to provide guidance on what is achievable with current good practice rather than on what should be considered optimum performance

## Slide 27: Diagnostic Reference Levels (DRLs) for medical exposure (as defined by the BSS)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
The DRLs are intended:
  to be applied with flexibility to allow higher exposures if these are indicated by sound clinical judgement
  to be revised as technology and techniques improve

## Slide 28: Diagnostic Reference Levels (DRLs) for medical exposure (as defined by the BSS)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Corrective actions should be taken as necessary if doses or activities fall substantially below the DRLs and images do not provide adequate clinical image quality

## Slide 29: Dose constraints for medical exposure
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
For medical exposure, dose constraints should only be used in optimizing the protection of persons exposed for medical research purposes, or of persons, other than workers, who assist in the care, support or comfort of exposed patients.

## Slide 30: Dose constraints
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
for medical research purposes
for individuals helping in care, support or comfort of patients, and visitors
  5 mSv during the period of the examination or treatment
  1 mSv for children visiting
maximum activity in patients discharged from hospitals
  Iodine 131-1100 MBq

## Slide 31: PUBLIC - Optimization under Constraints
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
DOSE LIMITS
effective dose of 1 mSv in a year
in special circumstances, effective dose of 5 mSv in a single year, provided that the average over five consecutive years in less than 1mSv per year
equivalent dose to lens of the eye 15 mSv in a year
equivalent dose to skin of 50 mSv in a year.

## Slide 32: Diagnostic Reference Levels
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Values of measured quantities above which some specified action or decision should be taken
The ICRP recommends the use of Diagnostic Reference Levels (DRL) for patients (Report 103), as does the IAEA in its International Basic Safety Standards (IAEA Safety Series 115, 2011)
The DRL is intended for use as
  a convenient test for identifying
  situations where the levels of patient
  dose are unusually high.

## Slide 33: Diagnostic Reference Levels Practical aspects (I)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Diagnostic Reference Levels are not dose limits
DRLs could be assimilated to investigation levels
DRL are not applicable to individual patients. Comparison with DRL shall be only made using mean values of a sample of patients
Quantities used for DRLs should be easily measured

## Slide 34: Diagnostic Reference Levels Practical aspects (II)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Quantities used for DRLs should be understood by radiologists and radiographers
DRLs should always be used in parallel with image quality evaluation (enough information for diagnosis shall be obtained)
DRLs can be based on several quantities (such as DAP) and parameters (such as fluoro time and number of images)

## Slide 35: Diagnostic Reference Levels Practical aspects (III)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
DRLs should be ‘flexible’ (tolerances should be established: different patient sizes, different pathologies, etc). DRLs are not a border line between good and bad medicine
Values BELOW DRLs may need optimization if the image quality is inadequate for clinical purposes. Values ABOVE DRLs require an investigation and optimization of X Ray system or protocols.
The main objective of DRLs is their use in a dynamic and continuous process of optimization

## Slide 36: Part 1: Overview of Radiation Protection in Diagnostic Radiology
Topic 5: Diagnostic Reference Levels and Effective Doses
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 37: Diagnostic reference levels for adult radiography
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Examination | HPA Entrance surface dose per radiograph (mGy) | NCRP Free-in-air (mGy)
Lumbar spine AP | 6 | 4.2
Lumbar spine LAT | 14 | 
Lumbar spine LSJ | 26 | 
Abdomen AP | 6 | 3.4

## Slide 38: Diagnostic references for adult radiography
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Examination | HPA Entrance surface dose per radiograph (mGy) | NCRP Free-in-air (mGy)
Pelvis AP | 4 | 
Hip joint AP |  | 
Chest PA | 0.2 | 0.15
Chest LAT | 1.0 | 

## Slide 39: Diagnostic reference levels for adult radiography
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Examination | HPA Entrance surface dose per radiograph (mGy) | NCRP Free-in-air (mGy)
Thoracic spine AP | 3.5 | 
Thoracic spine LAT | 10 | 
Dental intraoral | 2.4 (2007) | 1.6
Dental AP |  | 1.6

## Slide 40: Diagnostic reference levels for adult radiography
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Examination | HPA Entrance surface dose per radiograph (mGy)
Skull AP | 3
Skull LAT | 1.5

## Slide 41: Diagnostic reference levels for adult CT
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Examination | UK 3rd Quartile, 2006 | EC 2006 | NCRP CTDIW
Head | 66 | 60 | 82
Chest | 17 | 30 | 
Lumbar spine |  |  | 22
Abdomen | 19 | 35 | 22
(a) Derived from measurements on the axis of rotation in water equivalent phantoms, 15 cm in length and 16 cm (head) and 30 cm (lumbar spine and abdomen) in diameter. |  |  | 

## Slide 42: Diagnostic reference levels for mammography
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Average glandular dose per cranio-caudal projection
3.0 mGy (with grid)
Determined for a 4.5 cm compressed breast consisting of 50% glandular and 50% adipose tissue, for screen-film systems and dedicated Mo-target and Mo-filter mammography units.

## Slide 43: http://www.hpa.org.uk/web/HPAweb&HPAwebStandard/HPAweb_C/1195733771087 2000, last reviewed 2008 http://radiology.rsna.org/content/240/3/828.full.pdf+html CT DRLs http://www.hpa.org.uk/Publications/Radiation/HPARPDSeriesReports/HpaRpd022/ HPA dental 2.4 mGy NCRP values from draft report—to be published 2012
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Sources for DRL Values

## Slide 44: Diagnostic reference levels for adult fluoroscopy
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Operation Mode | Entrance surface dose (mGy/min) (a)
Normal | 25
High Level (b) | 100
(a) In air with backscatter (b) For fluoroscopes that have an optional 'high level' operational mode, such as those frequently used in interventional radiology | 

## Slide 45: Typical effective doses from diagnostic medical exposures
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
From: Referral Criteria For Imaging. CE, 2000.
Diagnostic procedure | Typical effective dose (mSv) | Equiv. no. of chest X rays | Approx. equiv. period of natural background radiation
Chest (single PA film) | 0.02 | 1 | 3 days
Skull | 0.07 | 3.5 | 11 days
Thoracic spine | 0.7 | 35 | 4 months
Lumbar spine | 1.3 | 65 | 7 months

## Slide 46: Typical effective doses from diagnostic medical exposures
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Diagnostic procedure | Typical effective dose (mSv) | Equiv. no. of chest X rays | Approx. equiv. period of natural background radiation
Hip | 0.3 | 15 | 7 weeks
Pelvis | 0.7 | 35 | 4 months
Abdomen | 1.0 | 50 | 6 months
IVU | 2.5 | 125 | 14 months
From: Referral Criteria For Imaging. CE, 2000.

## Slide 47: Typical effective doses from diagnostic medical exposures
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Diagnostic procedure | Typical effective dose (mSv) | Equiv. no. of chest X rays | Approx. equiv. period of natural background radiation
Barium swallow | 1.5 | 75 | 6 months
Barium meal | 3 | 150 | 16 months
Barium follow through | 3 | 150 | 16 months
Barium enema | 7 | 350 | 3.2 years
From: Referral Criteria For Imaging. CE, 2000.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 48: Typical effective doses from diagnostic medical exposures
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Diagnostic procedure | Typical effective dose (mSv) | Equiv. no. of chest X rays | Approx. equiv. period of natural background radiation
CT head | 2.3 | 115 | 1 year
CT chest | 8 | 400 | 3.6 years
CT Abdomen or pelvis | 10 | 500 | 4.5 years
From: Referral Criteria For Imaging. CE, 2000.

## Slide 49: Investigation of exposure (B.S.S. 3.46)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Registrants and licensees shall promptly investigate:
any diagnostic exposure substantially greater than intended or resulting in doses repeatedly and substantially exceeding the established diagnostic reference levels, or unusually low exposures
any equipment failure, accident, error, mishap or other unusual occurrence with the potential for causing a patient exposure significantly different from that intended.

## Slide 50: Investigation of exposure (B.S.S. 3.180)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Registrants and licensees shall:
calculate or estimate the doses received and their distribution within the patient
indicate the corrective measures required to prevent recurrence of such an incident
implement all the corrective measures that are under their own responsibility

## Slide 51: Investigation of exposure (B.S.S. 3.48)
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Registrants and licensees shall:
submit to the Regulatory Authority, as soon as possible after the investigation or as otherwise specified by the Regulatory Authority, a written report which states the cause of the incident and includes the information specified in (a) to (c), as relevant, and any other information required by the Regulatory Authority; and
inform the patient and his or her doctor about the incident.

## Slide 52: Summary
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Exposure of patients as part of their diagnosis or treatment, has to be justified
Optimization of patient exposures means keeping doses to a minimum without loss of diagnostic information
Diagnostic Reference Levels are defined to serve as a reference for medical practitioners: if a level is exceeded some specified action or decision should be taken
DRLs are not dose limits.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Let’s summarize the main subjects we did cover in this session. (List the main subjects covered and stress again the important features of the session)

## Slide 53: Where to Get More Information
1 : Overview of Radiation Protection in Diagnostic Radiology
<number>
Safety Requirements: Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards. Revision of IAEA Safety Series No.115, IAEA, September 2011.
ICRP 73. Radiological Protection and Safety in Medicine. Annals of the ICRP, 26(2), 1996.
Radiation Protection 118. Referral Guidelines for Imaging, European Commission, 2008. http://ec.europa.eu/energy/nuclear/radioprotection/publication/doc/118_update_en.pdf
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources