## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L18: Optimization of Protection in Computed Tomography (CT)
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Introduction
18: Optimization of Protection in CT Scanner
<number>
The subject matter: CT scanner and related image quality considerations
The importance of the technological improvement made in this field
The quality criteria system developed to optimize the CT procedure
Background: medical doctor, medical physicist
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 3: Topics
18: Optimization of Protection in CT Scanner
<number>
  CT equipment and technology
  Radiation protection rules and operational consideration
  Quality criteria for CT images
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 4: Overview
18: Optimization of Protection in CT Scanner
<number>
To understand the principles and the technology of CT
To be able to apply the principle of radiation protection to CT scanner including design, Quality Control and dosimetry.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 5: Part 18: Optimization of protection in CT scanner
Topic 1: CT equipment and technology
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 6: Introduction
18: Optimization of Protection in CT Scanner
<number>
Computed Tomography (CT) was introduced into clinical practice in 1972 and revolutionized X Ray imaging by providing high quality images which reproduced transverse cross sections of the body.
Tissues are not superimposed on the image as they are in conventional projections
The CT provides improved low contrast resolution for better visualization of soft tissue, but with relatively high radiation dose, i.e. CT is a high dose procedure

## Slide 7: Computed Tomography
18: Optimization of Protection in CT Scanner
<number>
CT uses a rotating X Ray tube, with the beam in the form of a thin slice (about 1 - 10 mm)
The “image” is a simple array of X Ray intensities, and many hundreds of these are used to make the CT image, which is a “slice” through the patient

## Slide 8: The CT Scanner
18: Optimization of Protection in CT Scanner
<number>

## Slide 9: A look inside a rotate/rotate CT
18: Optimization of Protection in CT Scanner
<number>
X Ray
Tube
Detector Array
and Collimator
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 10: Helical (spiral) CT
18: Optimization of Protection in CT Scanner
<number>
If the X Ray tube can rotate constantly, the patient can then be moved continuously through the beam, making the examination much faster

## Slide 11: Helical Scan Principle
18: Optimization of Protection in CT Scanner
<number>
Scanning Geometry
Continuous Data Acquisition and Table Feed
X Ray beam
Direction of patient
movement
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 12: Helical CT Scanners
18: Optimization of Protection in CT Scanner
<number>
For helical scanners, the X Ray tube rotates continuously
This is obviously not possible with a cable combining all electrical sources and signals
A “slip ring” is used to supply power and to collect the signals

## Slide 13: A Look Inside a Slip Ring CT
18: Optimization of Protection in CT Scanner
<number>
X Ray
Tube
Detector
Array
Slip Ring
Note:
how most
of the
electronics are
placed on
the rotating
gantry
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 14: New CT Features
18: Optimization of Protection in CT Scanner
<number>
The new helical scanning CT units allow a range of new features, such as:
  CT fluoroscopy, where the patient is stationary, but the tube continues to rotate
  multislice CT, where up to 128 slices can be collected simultaneously
  3-dimensional CT and CT endoscopy

## Slide 15: CT Fluoroscopy
18: Optimization of Protection in CT Scanner
<number>
Real Time Guidance (up to 8 fps)
Great Image Quality
High Dose Rate
Faster Procedures (up to 66% faster than non-fluoroscopic procedures)
Approx. 80 kVp, 30 mA
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 16: Multi slice CT collimation
18: Optimization of Protection in CT Scanner
<number>
0,5mm
1mm
2,5mm
5mm
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 17: 3D Stereo Imaging
18: Optimization of Protection in CT Scanner
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 18: CT Endoscopy
18: Optimization of Protection in CT Scanner
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 19: CT Scanner
18: Optimization of Protection in CT Scanner
<number>
  Generator
    High frequency, 30 - 70 kW
  X Ray tube
    Rotating anode, high thermal capacity: 3-7 MHU
    Dual focal spot sizes: about 0.8 and 1.4
  Gantry
    Aperture: > 70 cm of diameter
    Detectors: gas or solid state; > 600 detectors
    Scanning time: <1 s, 1 - 4 s
    Slice thickness: 1 - 10 mm
    Spiral scanning: up to 1400 mm

## Slide 20: Image processing
18: Optimization of Protection in CT Scanner
<number>
Reconstruction time:
  0.5 - 5 s/slice
Reconstruction matrix: 256x256 – 1024x1024
Reconstruction algorithms:
  Bone, Standard, High resolution, etc
Special image processing software:
  3D reconstruction
  Angio CT with MIP
  Virtual endoscopy
  CT fluoroscopy

## Slide 21: Spiral (helical) CT
18: Optimization of Protection in CT Scanner
<number>
Spiral CT and Spiral multislice CT:
Volume acquisition may be preferred to serial CT
Advantages:
  dose reduction:
    reduction of single scan repetition (shorter examination times)
    replacement of overlapped thin slices (high quality 3D display) by the reconstruction of one helical scan volume data
    use of pitch > 1
  no data missing as in the case of inter-slice interval
  shorter examination time
    to acquire data during a single breath-holding period avoiding respiratory disturbances
    disturbances due to involuntary movements such as peristalsis and cardiovascular action are reduced

## Slide 22: Spiral (helical) CT
18: Optimization of Protection in CT Scanner
<number>
Drawbacks
  Increasing of dose:
    equipment performance may tempt the operator to extend the examination area
  Use of a pitch > 1.5 and an image reconstruction at intervals equal to the slice width results in lower diagnostic image quality due to reduced low contrast resolution
  Loss of spatial resolution in the z-axes unless special interpolation is performed
  Technique inherent artifact

## Slide 23: Part 18: Optimization of protection in CT scanner
Topic 2: Radiation protection rules and operational consideration
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 24: Contribution to collective dose (I)
18: Optimization of Protection in CT Scanner
<number>
As a result of such technological improvements, the number of examinations have markedly increased
Today CT procedures contribute for up to 40% of the collective dose from diagnostic radiology in all developed countries
Special protection measures are therefore required

## Slide 25: Contribution to collective dose (II)
18: Optimization of Protection in CT Scanner
<number>
0
100
200
300
400
500
70
75
80
85
90
95
Years
CT scanners in clinical use
in UK
3.3
Lumbar spine
7.1
Pelvis
7.2
Liver
7.6
Abdomen
7.8
Chest
2.6
Cervical spine
0.6
Orbits
0.7
Posterior fossa
1.8
Routine head
Mean effective dose (mSv)
Examination

## Slide 26: Justification of CT practice
18: Optimization of Protection in CT Scanner
<number>
Justification in CT is of particular importance for RP
CT examination is a “high dose” procedure
A series of clinical factors play a special part
  Adequate clinical information, including the records of previous imaging investigations, must be available
  In certain applications prior investigation of the patient by alternative imaging techniques might be required
Additional training in radiation protection is required for radiologists and radiographers
Guidelines of EU are available

## Slide 27: Optimization of CT practice
18: Optimization of Protection in CT Scanner
<number>
Once a CT examination has been clinically justified, the subsequent imaging process must be optimized
There is dosimetric evidence that procedures are not optimized from the patient radiation protection point of view
Examination
CTDIw (mGy)
Sample size
Mean
SD
Min
25%
Median
75%
Max
Head
102
50.0
14.6
21.0
41.9
49.6
57.8
130
Chest
88
20.3
7.6
4.0
15.2
18.6
26.8
46.4
Abdomen
91
25.6
8.4
6.8
18.8
24.8
32.8
46.4
Pelvis
82
26.4
9.6
6.8
18.5
26.0
33.1
55.2

## Slide 28: Optimization of CT practice
18: Optimization of Protection in CT Scanner
<number>
Optimal use of ionizing radiation involves the interplay of the imaging process:
  Diagnostic quality of the CT image
  Radiation dose to the patient
  Choice of radiological technique

## Slide 29: Optimization of CT practice
18: Optimization of Protection in CT Scanner
<number>
CT examinations should be performed under the responsibility of a radiologist according to the national regulations
Standard examination protocols should be available.
Effective supervision may aid radiation protection by terminating the examination when the clinical requirement has been satisfied
Quality Criteria can be adopted by radiologists, radiographers, and medical physicists as a check on the routine performance of the entire imaging process

## Slide 30: Part 18: Optimization of protection in CT scanner
Topic 3: Quality criteria for CT images
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 31: Quality criteria for CT images: Example of good imaging technique (brain general examination)
18: Optimization of Protection in CT Scanner
<number>
40 - 45 HU (supratentorial brain) 30 - 40 HU (brain in posterior fossa) 200 - 400 HU (bones)
Window level
0 - 90 HU (supratentorial brain) 140- 160 HU (brain in posterior fossa) 2000 - 3000 HU (bones)
Window width
Soft
Reconstruction algorithm
As low as consistent with required image quality
Tube current and exposure time product (mAs)
Standard
X Ray tube voltage (kV)
10-12 ° above the orbito-meatal (OM) line to reduce exposure of the eye lenses
Gantry tilt
Head dimension (about 24 cm)
FOV
Contiguous or a pitch = 1
Inter-slice distance/pitch
2 - 5 mm in posterior fossa; 5-10 mm in hemispheres
Nominal slice thickness
From foramen magnum to the skull vertex
Volume of investigation
Supine
Patient position

## Slide 32: Quality criteria for CT images: brain, general examination
18: Optimization of Protection in CT Scanner
<number>
Image criteria
  Visualization of
    Whole cerebrum, cerebellum, skull base and osseous basis
    Vessels after intravenous contrast media
  Critical reproduction
    Visually sharp reproduction of the
      border between white and grey matter
      basal ganglia
      ventricular system
      cerebrospinal fluid space around the mesencephalon
      cerebrospinal fluid space over the brain
      great vessels and the choroid plexuses after i.v. contrast
Criteria for radiation dose to the patient
    CTDIW 60 mGy
    DLP 1050 mGy cm

## Slide 33: Image criteria for CT images: brain, general examination (visualization of)
18: Optimization of Protection in CT Scanner
<number>
  Whole cerebrum, cerebellum, skull base and osseous basis
  Vessels after intravenous contrast media

## Slide 34: Image criteria for CT images: brain, general examination (critical reproduction)
18: Optimization of Protection in CT Scanner
<number>
Visually sharp reproduction of the:
  border between white and grey matter
  basal ganglia
  ventricular system
  cerebrospinal fluid space around the mesencephalon
  cerebrospinal fluid space over the brain
  great vessels and the choroid plexuses after i.v. contrast

## Slide 35: Quality criteria for CT images
18: Optimization of Protection in CT Scanner
<number>
A preliminary list of reference dose for the patient are given for some examinations expressed in term of:
  CTDIw for the single slice
  DLP for the whole examination
Examination
Reference doses
CTDIw (mGy)
DLP (mGy cm)
Routine head
60
1050
Routine chest
30
650
Routine abdomen
35
800
Routine pelvis
35
600

## Slide 36: Viewing conditions and film processing
18: Optimization of Protection in CT Scanner
<number>
Viewing conditions
It is recommended to read CT images on video display
Brightness and contrast control on the viewing monitor should give a uniform progression of the grey scale
Choice of window width dictates the visible contrast between tissues
Film Processing
Optimal processing of the film has important implications for the diagnostic quality
Film processors should be maintained at their optimum operating conditions by frequent (i.e., daily) quality control

## Slide 37: Summary
18: Optimization of Protection in CT Scanner
<number>
The CT scanner technology and the related radiation protection aspects
The ways of implementing the quality criteria system related to the image quality and to dosimetry
The importance of Quality Control
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Let’s summarize the main subjects we did cover in this session. (List the main subjects covered and stress again the important features of the session)

## Slide 38: Where to Get More Information (II)
18: Optimization of Protection in CT Scanner
<number>
Quality criteria for computed tomography, EUR 16262 report, (Luxembourg, EC), 1997. http://w3.tue.nl/fileadmin/sbd/Documenten/Leergang/BSM/European_Guidelines_Quality_Criteria_Computed_Tomography_Eur_16252.pdf
Radiation exposure in Computed Tomography; 4th revised Edition, December 2002, H.D.Nagel, CTB Publications, D-21073 Hamburg
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources