## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L12: Shielding and X Ray room design
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Introduction
12: Shielding and X Ray room design
<number>
Subject matter: the theory of shielding design and some related construction aspects.
The method used for shielding design and the basic shielding calculation procedure
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 3: Topics
12: Shielding and X Ray room design
<number>
  Equipment design and acceptable safety standards
  Use of dose constraints in X Ray room design
  Barriers and protective devices
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 4: Overview
12: Shielding and X Ray room design
<number>
To become familiar with the safety requirements for the design of X Ray systems and auxiliary equipment, shielding of facilities, and relevant international safety standards, e.g., IEC.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 5: Part 12: Shielding and X Ray room design
Topic 1: Equipment design and acceptable
safety standards
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 6: Purpose of Shielding
12: Shielding and X Ray room design
<number>
To protect:
  the patients (when not being examined)
  the X Ray department staff
  visitors and the public
  persons working adjacent to or near the X Ray facility
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 7: Radiation Shielding - Design Concepts
12: Shielding and X Ray room design
<number>
Data required include consideration of:
  Type of X Ray equipment
  Usage (workload)
  Positioning
  Whether multiple tubes/receptors are being used
  Primary beam access (vs. scatter only)
  Operator location
  Occupancy of Surrounding areas
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
A floor plan to a known scale, including not only the x-ray room, but also surrounding areas (including their function e.g. office, toilet, waiting room etc).
    The location of the x-ray table and the type and orientation of the equipment.
    The location of any upright bucky or chest stand (used to take X Rays of standing patients).
    Details of what lies above, below and adjacent to the X Ray room, and the nature of the floor, wall and ceiling construction.
    The distances from the X Ray tube and patient to points which are to be used in the calculations. Distance is denoted as d.
    The target, or design, weekly radiation dose at each calculation point. This is called P.

## Slide 8: Shielding Design (I)
12: Shielding and X Ray room design
<number>
Equipment
What equipment is to be used?
  General radiography
  Fluoroscopy (with or without radiography)
  Dental (oral, cephalometric, or OPG)
  Mammography
  CT
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 9: Shielding Design (II)
12: Shielding and X Ray room design
<number>
The type of equipment is very important for the following reasons:
  where the X Ray beam will be directed
  the number and type of procedures performed
  the location of the radiographer (operator)
  the energy (kVp) of the X Rays
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 10: Shielding Design (III)
12: Shielding and X Ray room design
<number>
Usage
Different X Ray equipment have very different usage.
For example, a dental unit uses low mAs and low (~70) kVp, and takes relatively few X Rays each week
A CT scanner uses high (~130) kVp, high mAs, and takes very many scans each week.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 11: Shielding Design (IV)
12: Shielding and X Ray room design
<number>
The total mAs used each week is an indication of the total X Ray dose administered
The kVp used is also related to dose, but also indicates the penetrating ability of the X Rays
High kVp and mAs means that more shielding is required.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 12: Shielding Design (V)
12: Shielding and X Ray room design
<number>
Positioning
The location and orientation of the X Ray unit is very important:
  distances are measured from the equipment (inverse square law will affect dose)
  the directions the direct (primary) X Ray beam will be used depend on the position and orientation
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 13: Radiation Shielding - Typical Room Layout
12: Shielding and X Ray room design
<number>
A to G are points
used to calculate
shielding
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 14: Shielding Design (VI)
12: Shielding and X Ray room design
<number>
Number of X Ray tubes
Some X Ray equipment may be fitted with more than one tube
Sometimes two tubes may be used simultaneously, and in different directions
This naturally complicates shielding calculation
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 15: Shielding Design (VII)
12: Shielding and X Ray room design
<number>
Surrounding areas
The X Ray room must be designed with knowledge of the location and use of all rooms which adjoin the X Ray room
Obviously a toilet will need less shielding than an office
Obtain a plan of the X Ray room and surroundings (including level above and below)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 16: Radiation Shielding - Design Detail
12: Shielding and X Ray room design
<number>
Must consider:
  appropriate calculation points, covering all critical locations
  design parameters such as workload, occupancy, use factor, leakage, target dose (see later)
  these must be either assumed or taken from actual data
  use a reasonable, worst case scenario (conservatively high estimates), since under-shielding is worse than over-shielding
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 17: Part 12: Shielding and X Ray room design
Topic 2: Use of dose constraints in
X Ray room design
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 18: Radiation Shielding - Calculation
12: Shielding and X Ray room design
<number>
Based on NCRP Report No 147, Structural Shielding Design for Medical X-Ray Imaging Facilities (2004)
Assumptions used are conservative, so over-shielding is common
Software is available, giving shielding in thickness of various materials
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
For many years, the method used for shielding design has been that described in the (American) National Council on Radiation Protection and Measurements (NCRP) Report 49. The data used in this report has usually resulted in overshielding, and the report is currently undergoing review. However, this document is currently the best reference available and for this reason, this module is based on NCRP 49

## Slide 19: Radiation Shielding Parameters (I)
12: Shielding and X Ray room design
<number>
P - design dose per week
usually based on 5 mSv per year for occupationally exposed persons (25% of dose limit), and 1 mSv for public
occupational dose must only be used in controlled areas, i.e., for radiographers, radiologists, and other radiation workers
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Each country has its own dose limits, but we will assume here that the values given in ICRP Report 60 (which are very widely used) apply. For occupationally exposed persons, the effective dose limit is 20 mSv per year. This averages to 0.4 mSv per week.
In addition, many countries are now applying an additional constraint in accordance with the ICRP 60 principle of optimisation of protection, on the basis that any one person could be exposed to more than one source of radiation

## Slide 20: Radiation Shielding Parameters (II)
12: Shielding and X Ray room design
<number>
Film storage areas (darkrooms) need special consideration
Long periods of exposure will affect film, but much shorter periods (i.e., lower doses) will fog film in cassettes
A simple rule is to allow 0.1 mGy for the period the film is in storage - if this is 1 month, the design dose is 0.025 mGy/week
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 21: Radiation Shielding Parameters (III)
12: Shielding and X Ray room design
<number>
Remember we must shield against three sources of radiation
In decreasing importance, these are:
  scattered radiation (from the patient)
  primary radiation (the X Ray beam)
  leakage radiation (from the X Ray tube)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Anything which separates one area from another is called a barrier. Any barrier which may be in the direct X Ray beam is called a primary barrier. If the X Ray beam will never be directed towards a barrier, it is called a secondary barrier. In practice, some barriers will have the primary beam directed at them part of the time only, and the rest of the time they will be a secondary barrier. This must be taken into account in the calculations.

## Slide 22: Radiation Shielding Parameters (IV)
12: Shielding and X Ray room design
<number>
U - use factor
fraction of time the primary beam is in a particular direction i.e.: the chosen calculation point
must allow for realistic use
for all points, sum may exceed 1
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Once the type of barrier has been decided, the next factor to be determined is the use factor (U) (i.e. the proportion of time the beam may be pointed at that barrier). Use factors are usually assumed, but may be calculated for a particular case, based on actual operational information.

## Slide 23: Radiation Shielding Parameters (V)
12: Shielding and X Ray room design
<number>
For some X Ray equipment, the X Ray beam is always stopped by the image receptor, thus the use factor is 0 in other directions, e.g., CT, fluoroscopy, mammography
For general radiographic and fluoroscopic equipment the primary beam is usually intercepted by the image detector
This reduces shielding requirements
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 24: Radiation Shielding Parameters (VI)
12: Shielding and X Ray room design
<number>
For radiography, there will be certain directions where the X Ray beam will be pointed:
  towards the floor
  across the patient, usually only in one direction
  toward the chest Bucky stand
The type of tube suspension will be important, e.g.: ceiling mounted, floor mounted, C-arm etc.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 25: Radiation Shielding Parameters (VII)
12: Shielding and X Ray room design
<number>
T - Occupancy
T = fraction of time a particular place is occupied by staff, patients or public
Has to be conservative
Ranges from 1 for adjacent offices and work areas, to 1/20 for public toilets and 1/40 for outdoor areas with transient traffic
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
The occupancy factor (T) is an indication of how long a particular place or room may be occupied by an individual person. Thus an occupancy of 1 implies that the same person will spend all their working week in that place. Occupancy factors vary greatly, depending on the category of the area.

## Slide 26: Occupancy (NCRP 147)
12: Shielding and X Ray room design
<number>
Area
Occupancy
Work areas, offices, staff rooms
1
Corridors
1/5
Toilets, unattended
waiting rooms
1/20
Outdoor areas with transient traffic
1/40
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
This table shows the standard NCRP 49 area categories and occupancy values.

## Slide 27: Radiation Shielding Parameters (VIII)
12: Shielding and X Ray room design
<number>
W - Workload
A measure of the radiation output in one week
Measured in mA-minutes
Varies greatly with assumed maximum kVp of X Ray unit
Usually a gross overestimation
Actual dose/mAs can be estimated
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
To calculate correct shielding, we need to know the amount of work an X Ray unit does in a week. This is known as the workload (W) of the unit.

## Slide 28: Workload (I)
12: Shielding and X Ray room design
<number>
For example: a general radiography room
The kVp used will be in the range 60-120 kVp
The exposure for each film will be between 5 mAs and 100 mAs
There may be 50 patients per day, and the room may be used 7 days a week
Each patient may have between 1 and 5 films
SO HOW DO WE ESTIMATE W ?
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 29: Workload (II)
12: Shielding and X Ray room design
<number>
Assume an average of 50 mAs per film, 3 films per patient
Thus W = 50 mAs x 3 films x 50 patients x 7 days
= 52,500 mAs per week
= 875 mA-min per week
We could also assume that all this work is performed at 100 kVp
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 30: Examples of Workloads
12: Shielding and X Ray room design
<number>
Weekly Workload (W)
mA-min at:
100
kVp
125
kVp
150
kVp
General Radiography
1,000
400
200
Fluoroscopy (including spot films)
750
300
150
Chiropractic
1,200
500
250
Mammography
700 at 30
kVp (1,500 for breast
screening)
Dental
6 at 70
kVp (conventional intra-oral
films)
For more realistic values and CT see NCRP 147
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
When considering workload, it is also important to know what kVp is used for the exposures. This is for two reasons: firstly, the mAs per exposure is lower for higher kVp, and secondly because the radiation is more penetrating as the kVp is increased.
NCRP 49 and many regulatory bodies have used quite high guidance values for usual workloads.
This table shows some of these values and the relevant kVp.It is now generally considered that these values are grossly inflated, particularly with modern radiographic film-screen systems which use very much less radiation than systems of 1976 when NCRP 49 was written. We also know, from recording actual workloads, that the kVp used is only occasionally greater than 100 kVp, and mostly around 90 kVp for a general radiography room. This concept of ‘workload spectrum’ will eventually be used in shielding calculations, but is not used here.

## Slide 31: Workload - CT
12: Shielding and X Ray room design
<number>
CT workloads are best calculated from NCRP 147
Remember that new spiral CT units, or multi-slice CT, could have higher workloads
A typical CT workload is about 28,000 mA-min per week
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 32: Tube Leakage
12: Shielding and X Ray room design
<number>
All X Ray tubes have some radiation leakage - there is only 2-3 mm lead in the housing
Leakage is limited in most countries to 1 mGy hr-1 at 1 meter, so this can be used as the actual leakage value for shielding calculations
Leakage is specified at the maximum rated continuous tube current, which is about 3-5 mA at 150 kVp for most radiographic X Ray tubes
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 33: Radiation Shielding Parameters
12: Shielding and X Ray room design
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
The distance from the X Ray tube to the scatterer (patient) is called dsca, the distance from the X Ray tube to a primary barrier is called dpri, and the distance from the scatterer to a secondary barrier is called dsec.

## Slide 34: Room Shielding - Multiple X Ray Tubes
12: Shielding and X Ray room design
<number>
Some rooms will be fitted with more than one X Ray tube (maybe a ceiling-mounted tube, and a floor-mounted tube)
Shielding calculations MUST consider the TOTAL radiation dose from the two tubes
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 35: Part 12: Shielding and X Ray room design
Topic 3: Barriers and protective devices
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 36: Shielding - Construction I
12: Shielding and X Ray room design
<number>
Materials available:
  Lead sheet brick
  gypsum or high Z plasterboard
  concrete block
  leaded glass or acrylic
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
The building materials available for shielding will vary according to the country. Some possibilities are:
    Lead sheet bonded onto a solid backing such as plywood, compressed cement fibre board, particle board or similar.
    Cement blocks - where used, they should preferably be solid, and care must be taken to ensure that the mortar joins carry through the full thickness of the blocks. As a rule of thumb, you can assume that a cement block is equivalent to at least 2/3 of its thickness in solid concrete.
    Bricks may be used provided that they will give sufficient attenuation. Mortar joints must carry through the full thickness of the brick. Bricks vary greatly in their attenuation, therefore you must be careful when using this type of shielding.
    Lead glass or lead acrylic for windows.
When advising on shielding materials it is often useful to know the comparative densities and lead equivalence of various materials, so that options can be considered.

## Slide 37: Shielding - Construction Problems
12: Shielding and X Ray room design
<number>
Some problems with shielding materials:
Brick walls - mortar joints
Use of lead sheets nailed to timber frame
Lead inadequately bonded to backing
Joints between sheets with no overlap
Use of hollow core brick or block
Use of plate glass where lead glass specified
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 38: Problems in shielding - Brick Walls & Mortar Joints
12: Shielding and X Ray room design
<number>
Bricks should be solid and not hollow
Bricks have very variable X Ray attenuation
Mortar is less attenuating than brick
Mortar is often not applied across the full thickness of the brick
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 39: Problems in shielding - Lead inadequately bonded to backing
12: Shielding and X Ray room design
<number>
Lead must be fully glued (bonded) to a backing such as wood or wallboard
If the lead is not properly bonded, it may peel off after a few years
Not all glues are suitable for lead (oxidization of the lead surface)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 40: Problems in shielding - Joins between sheets with no overlap
12: Shielding and X Ray room design
<number>
There must be 10 - 15 mm overlap between adjoining sheets of lead
Without an overlap, there may be relatively large gaps for the radiation to pass through
Corners are a particular problem
Penetrations for electrical boxes and ducts are of concern
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 41: Problems in shielding - Use of plate glass
12: Shielding and X Ray room design
<number>
Leaded glass or acrylic should be used for windows
Laminated layers of plate glass can be used where radiation levels are low, e.g., for the wall at the foot of a CT scanner
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 42: Radiation Shielding - Construction II
12: Shielding and X Ray room design
<number>
Continuity and integrity of shielding very important
Problem areas:
  joints
  penetrations in walls and floor
  window frames
  doors and frames
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 43: Penetrations
12: Shielding and X Ray room design
<number>
“Penetrations” means any hole cut into the lead for cables, electrical connectors, pipes etc.
Unless the penetration is small (~2-3 mm), there must be additional lead over the hole, Nails and screws used to fix bonded lead sheet to a wall do not require covering
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 44: Window frames
12: Shielding and X Ray room design
<number>
The lead sheet fixed to a wall must overlap any lead glass window fitted
It is common to find a gap of up to 5 cm, which is unacceptable
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 45: Shielding of Doors and Frames
12: Shielding and X Ray room design
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 46: Shielding - Verification I
12: Shielding and X Ray room design
<number>
Verification of appropriate thickness and proper installation is mandatory
Two choices - visual or measurement
Visual check (preferred) must be performed before shielding covered - the actual lead thickness can be measured easily
Radiation measurement necessary for window and door frames etc.
Isotope source simplifies measurements
Measurement for walls very slow
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
The installation of the shielding should be supervised by someone with the appropriate knowledge. Even a weekly visit to the building site, and good communications with the builder can avoid problems, delays, and expensive alterations.
You have two options when it comes to being satisfied that the shielding has been correctly constructed - you verify it as it is being built or you verify it after it has been built.
You should never just assume that shielding is correct. It must always be checked.
Of these two options, verification during construction is by far the easiest. All that is required is a visit to the site at each stage before the shielding material is covered up. That way, you can easily see that the shielding is free of holes, is the correct height and the correct thickness, with sufficient overlap of materials. Windows should be checked before the join to wall shielding is covered. A very common fault is that a gap in the shielding is left around the window, sometimes as wide as 3 cm.
Verification after completion is laborious, inexact, and difficult.

## Slide 47: Shielding Testing
12: Shielding and X Ray room design
<number>
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 48: Records
12: Shielding and X Ray room design
<number>
It is very important to keep records of shielding calculations, as well as details of inspections and corrective action taken to fix faults in the shielding
In 5 years, it might not be possible to find anyone who remembers what was done!
Shielding records should become a permanent part of the facility engineering records; a copy should be stored in the room (consider a permanent plaque on the wall in the room specifying the amount of shielding in each wall); and a copy should be retained by the medical physicist doing the calculations
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 49: Summary
12: Shielding and X Ray room design
<number>
The design of shielding for an X Ray room is a relatively complex task, but can be simplified by the use of some standard assumptions
Record keeping is essential to ensure traceability and constant improvement of shielding according to both practice and equipment modification
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Let’s summarize the main subjects we did cover in this session. (List the main subjects covered and stress again the important features of the session)

## Slide 50: Where to Get More Information (I)
12: Shielding and X Ray room design
<number>
Radiation shielding for diagnostic X Rays. BIR report (2000) Ed. D.G. Sutton & J.R. Williams
National Council on Radiation Protection and Measurements, Report No. 147, “Structural Shielding Design for Medical X-Ray Imaging Facilities” Bethesda, MD 2004.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources