## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L10: Patient dose assessment
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Introduction
10: Patient dose assessment
<number>
A review is made of:
The different parameters influencing the patient dose
The problems related to instrument calibration
The existing dosimetric methods applicable to diagnostic radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
When the audience is familiar with instrumentation and instrument calibration, slides 9-16 can be skipped
Instructions for the lecturer/trainer

## Slide 3: Topics
10: Patient dose assessment
<number>
  Parameters influencing patient exposure
  Dosimetry methods
  Instrument calibration
  Dose measurements
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 4: Overview
10: Patient dose assessment
<number>
To become familiar with the patient dose assessment and dosimetry instrument characteristics.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 5: Part 10: Patient dose assessment
Topic 1: Parameters influencing patient dose
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 6: Essential parameters influencing patient exposure
10: Patient dose assessment
<number>
Tube voltage
Tube current
Effective filtration
Exposure time
Field size
Kerma rate
[mGy/min]
[min]
Kerma
[Gy]
[m2]
Area exposure
product
[Gy m2 ]
}
}
}
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 7: Factors in conventional radiography: beam, collimation
10: Patient dose assessment
<number>
Beam energy
  Depending on peak kV and filtration
  Regulations require minimum total filtration to absorb lower energy photons
  Added filtration reduces dose
  Goal should be use of highest kV resulting in acceptable image contrast
Collimation
  Area exposed should be limited to area of CLINICAL interest to lower dose
  Additional benefit is less scatter, better contrast

## Slide 8: Factors in conventional radiography: grid,patient size
10: Patient dose assessment
<number>
Grids
  Reduce the amount of scatter reaching image receptor
  But at the cost of increased patient dose
  Improves image contrast significantly
  Typically 2-5 times: “Bucky factor”
  Patient size
  Thickness, volume irradiated…and dose increases with patient size
  Except for breast (compression): no control
  Technique charts with technique factors for various examinations and patient thickness essential to avoid retakes
  Also, patient thickness must be measured accurately to use technique charts properly

## Slide 9: Factors affecting dose in fluoroscopy
10: Patient dose assessment
<number>
Beam energy and filtration
Collimation
Source-to-skin distance
  Inverse square law: maintain max distance from patient
Patient-to-image intensifier
  Minimizing patient-to-image intensifier distance will lower dose and improve image sharpness

## Slide 10: Factors affecting dose in fluoroscopy
Image magnification
  Geometric and electronic magnification increase dose
Grid
  If small sized patient (less scatter) probably not needed
    No need for grids on pediatric patients
    Grids not necessary for high contrast studies, e.g., barium contrast studies
Beam-on time!
10: Patient dose assessment
<number>

## Slide 11: Factors affecting dose in CT
10: Patient dose assessment
<number>
Beam energy and filtration
  80-100 kV reduces dose for pediatric patients
  120-140 kV with additional filtration reduces adult doses (HVL can be increased to reduce dose)
Collimation or section thickness
  Post-patient collimator will reduce slice thickness imaged but not the irradiated thickness
Number and spacing of adjacent sections
Image quality and noise
  Like all modalities: dose increase=>noise decreases

## Slide 12: Factors affecting dose in spiral CT
10: Patient dose assessment
<number>
Factors for conventional CT also valid
Scan pitch
  Ratio of couch travel in 1 rotation dived by slice thickness
  If pitch = 1, doses are comparable to conventional CT
  Dose proportional to 1/pitch

## Slide 13: Part 10: Patient dose assessment
Topic 2: Patient dosimetry methods
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology

## Slide 14: Radiation Dose Measurement
Ionization chamber measurements
Thermoluminescent dosimeters (TLDs)
Optically stimulated luminescent (OSL) dosimeters
Solid state dosimeters
Film (silver halide or radiochromic)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
There are absolute methods which provide the information on dose but they need the knoledge of characteristic parameters. Some others are relative and should be calibrated against the refernce value

## Slide 15: Patient dosimetry
10: Patient dose assessment
<number>
Radiography: entrance surface dose ESD
  By TLD or OSL
  Output factor
Fluoroscopy: Dose Area Product (DAP) or using film
CT:
  Computed Tomography Dose Index (CTDI)
  Using pencil ion chamber, OSL, or TLD

## Slide 16: From ESD to organ and effective dose
10: Patient dose assessment
<number>
Except for invasive methods, no organ doses can be measured
The only way in radiography: measure the Entrance Surface Dose (ESD)
Use mathematical models based on Monte Carlo simulations: the history of thousands of photons is calculated
Dose to the organ tabulated as a fraction of the entrance dose for different projections
Since filtration, field size and projection play a role: long lists of tables (See NRPB R262 and NRPB SR262)

## Slide 17: From DAP to organ and effective dose
10: Patient dose assessment
<number>
In fluoroscopy: moving field, measurement of Dose-Area Product (DAP)
In similar way organ doses calculated by Monte Carlo modelling
Conversion coefficients were estimated as organ doses per unit dose-area product
Again numerous factors are to be taken into account as projection, filtration, …
Once organ doses are obtained, effective dose is calculated following ICRP 103

## Slide 18: Part 10: Patient dose assessment
Topic 3: Instrument calibration
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology

## Slide 19: Calibration of an instrument
10: Patient dose assessment
<number>
Establish Calibration Reference Conditions (CRC) [type and energy of radiation, SDD, rate, ...]
Compare response of your instrument with that of another instrument (absolute or calibrated)
Determine the calibration factor
[appropriate unit]
Response of the instrument to be calibrated
f the reference instrument
Response o
F


## Slide 20: Range of use
10: Patient dose assessment
<number>
Hypothesis: the instrument reading is a known monotonic function of the measured quantity (usually linear within a specified range)

1/F = tg 
Instrument
Reading
Measured Quantity
Response at
calibration
Calibration Value

## Slide 21: Use of a calibrated instrument
10: Patient dose assessment
<number>
Under the same conditions as the CRC
Within the range of use
Q (dosimetric quantity) = F x R (reading of the instrument)

## Slide 22: Correction factors for use other than under the CRC
10: Patient dose assessment
<number>
0.92
0.94
0.96
0.98
1
1.02
1.04
1.06
1
2
3
4
HVL(mm Al)
Correction
Factor
A. Energy correction factor

## Slide 23: Correction factors for use other than under the CRC
10: Patient dose assessment
<number>
B. Directional correction factor

## Slide 24: Correction factors for use other than under the CRC
10: Patient dose assessment
<number>
C. Air density correction factor
(for ionization chambers)
p
t
0
0
,
calibration values
)
273
(
)
273
(
0
0



t
p
t
p
K
D

## Slide 25: Accuracy and precision of a calibrated instrument (1)
10: Patient dose assessment
<number>
Curve A: Instrument both accurate and precise
Curve B: Instrument accurate but not precise
Curve C: Instrument precise but not accurate
True value
A
B
C
Readings

## Slide 26
10: Patient dose assessment
<number>
Accuracy and precision of a calibrated instrument (2)
Traceability
Accuracy
Primary standard
(absolute measurement)
Secondary standard
Field instrument
Calibration
decreases
Relative uncertainty associated to the dosimetric quantity Q:
Where:
rC is the relative uncertainty of the reading of
the calibrated instrument
rR is the relative uncertainty of the reading of
the reading instrument
Calibration
rQ2 ≥ rC2 + rR2

## Slide 27: Requirements on Diagnostic dosimeters
10: Patient dose assessment
<number>
Traceability
Accuracy
Well defined reference X Ray spectra
not available
At least 10 - 30 %
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
It is not so common to have a well defined reference X-ray spectrum available for the useful energy range to be used in diagnsotic radiology.

## Slide 28: Limits of error in the response of diagnostic dosimeters
10: Patient dose assessment
<number>
Parameter
Range of values
Reference condition
Deviation (%)
Radiation quality
According to manufacturer
70 kV
5-8
Dose rate
According to manufacturer
--
4
Direction of radiation incidence
±5°
Preference direction
3
Atmospheric pressure
80-106 hPa
101.3 hPa
3
Ambient temperature
15-30°
20° C
3

## Slide 29: Part 10: Patient dose assessment
Topic 4: Dose measurements: how to measure dose indicators ESD, DAP,CTDI…
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology

## Slide 30: What we want to measure
10: Patient dose assessment
<number>
The radiation output of X Ray tubes
The dose-area product
The computed-tomography dose index (CTDI)
Entrance surface dose
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 31: Measurements of Radiation Output
10: Patient dose assessment
<number>
X Ray tube
Filter
Ion. chamber
Lead slab
Table top
SDD
Phantom (PEP)

## Slide 32: Measurements of Radiation Output
10: Patient dose assessment
<number>
Operating conditions
Consistency check
The output as a function of kVp
The output as a function of mA
The output as a function of exposure time

## Slide 33: Dose Area Product (DAP)
10: Patient dose assessment
<number>
Transmission ionization chamber
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
DAP meter can be put anywhere between the tube and the patient since its reading is independent on the position

## Slide 34: Dose Area Product (DAP)
10: Patient dose assessment
<number>
0.5 m
1 m
2 m
Air Kerma:
Area:
Area exposure product
40*103 Gy
2.5*10-3 m2
100 Gy m2
10*103 Gy
10*10-3 m2
100 Gy m2
2.5*103 Gy
40*10-3 m2
100 Gy m2
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
As can be seen the measured DAP quantity is constant since the Air Kerma decreases according to the inverse square law and the area increases as the square of the distance

## Slide 35: Calibration of a Dose Area Product (DAP)
10: Patient dose assessment
<number>
Film cassette
10 cm
10 cm
Ionization
chamber

## Slide 36: Computed Tomography Dose Index (CTDI)
10: Patient dose assessment
<number>
0
10
20
30
40
50
1
2
3
4
5
6
7
8
9
10
11
12
TLD dose (mGy)
Nominal slice width
3 mm
CTDI=
 (ei di)
En
En: nominal slice width
ei : TLD thickness
CTDIn=
mAs
CTDI
Normalized CTDI:
CTDI = 41.4

## Slide 37: Computed Tomography Dose Index (CTDI)
10: Patient dose assessment
<number>
Dose
Nominal slice width
CTDI
Dose profile

## Slide 38: TLD arrangement for CTDI measurements
10: Patient dose assessment
<number>
Axis of
rotation
Support jig
X Ray beam
Capsule
Couch
Gantry
Gantry
Capsule
X Ray beam
axis
of
rotation
LiF -TLD

## Slide 39: CTDI in air with pencil-type ionization chamber
18: Optimization of protection in CT scanner
<number>
The Computed Tomography Dose Index (CTDI) in air can be measured using a 10cm pencil ionization chamber, bisected by the scan plane at the isocentre, supported from the patient table
The ion chamber can be supported using a retort stand and clamp, if a dedicated holder is not available

## Slide 40: CTDI in air with pencil-type ionization chamber
18: Optimization of protection in CT scanner
<number>
Ionization
chamber
Table

## Slide 41
18: Optimization of protection in CT scanner
<number>
Axial slice positions
Helical scan (pitch 1)
CTDI in air with pencil-type ionization chamber

## Slide 42: Measurement of entrance surface dose
10: Patient dose assessment
<number>
TLD or OSL
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
The entrance surface dose measured under these conditions includes the backscatter

## Slide 43: Summary
10: Patient dose assessment
<number>
In this lesson we learned the factors influencing patient dose, and how to determine the entrance dose, dose area product, and CT dose.

## Slide 44: Where to Get More Information
10: Patient dose assessment
<number>
The Essential Physics of Medical Imaging. JT Bushberg, JA Seibert, EM Leidholdt, JM Boone. Lippincott Williams & Wilkins, Philadelphia, 2011
The 2007 Recommendations of the International Commission on Radiological Protection, ICRP 103, Annals of the ICRP 37(2-4):1-332 (2007)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources