## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L 2: Radiation units and dose quantities
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Introduction
2: Radiation units and dose quantities
<number>
Subject matter: the basic dosimetric quantities
Several quantities and units are needed in the field of diagnostic radiology and related dosimetry
Some can be measured directly while others can only be calculated
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 3: Topics
2: Radiation units and dose quantities
<number>
Exposure and exposure rate
Absorbed dose and KERMA
Mean Absorbed Dose in a tissue
Equivalent dose H
Effective Dose
Related dosimetry quantities (surface and depth dose, backscatter factor…..)
Specific dosimetry quantities (Mammography, CT,…)

## Slide 4: Objective
2: Radiation units and dose quantities
<number>
To become familiar with dosimetric quantities and units, and to perform related calculations.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 5: Part 2: Radiation units and dose quantities
Topic 1: Exposure and exposure rate
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 6: Exposure
2: Radiation units and dose quantities
<number>
Exposure is a dosimetric quantity for ionizing radiation, based on the ability of the radiation to produce ionization in air.
This quantity is only defined for radiation producing interactions in air.

## Slide 7: Exposure
2: Radiation units and dose quantities
<number>
Before interacting with the patient
(direct beam) or with the staff (scattered radiation), X Rays interact with air
The quantity “exposure” gives an indication of the capacity of X Rays to produce a certain effect in air
The effect in tissue will be, in general, proportional to this effect in air

## Slide 8: Exposure
2: Radiation units and dose quantities
<number>
X = dQ/dm
The exposure is the absolute value of the total charge of the ions of one sign produced in air when all the electrons liberated by photons per unit mass of air are completely stopped in air.

## Slide 9: Exposure: X
2: Radiation units and dose quantities
<number>
The SI unit of exposure is Coulomb per kilogram [C kg-1]
The former special unit of exposure was Roentgen [R]
1 R = 2.58 x 10-4 C kg-1
C kg-1 = 3876 R

## Slide 10: Exposure rate: X/t
2: Radiation units and dose quantities
<number>
Exposure rate (and later, dose rate) is the exposure produced per unit of time.
The SI unit of exposure rate is the C/kg per second or R/s.
In radiation protection it is common to indicate these rate values “per hour” (e.g. R/h).

## Slide 11: Part 2: Radiation units and dose quantities
Topic 2: Absorbed dose and KERMA
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 12: Patient dosimetry quantities
2: Radiation units and dose quantities
<number>

## Slide 13: Absorbed dose, D
2: Radiation units and dose quantities
<number>
The absorbed dose D, is the energy absorbed per unit mass. This quantity is defined for all ionizing radiation (not only for electromagnetic radiation, as in the case of the “exposure”), and for any material.
D = dE/dm. The SI unit of D is the Gray [Gy].
1 Gy = 1 J/kg.
The former unit was the “rad”. 1 Gy = 100 rad.

## Slide 14: Absorbed dose, D and KERMA
2: Radiation units and dose quantities
<number>
The KERMA (kinetic energy released in a mass)
K = dEtrans/dm
  where dEtrans is the sum of the initial kinetic energies of all charged ionizing particles liberated by uncharged ionizing particles in a material of mass dm
The SI unit of kerma is the joule per kilogram (J/kg), termed Gray (Gy).
In diagnostic radiology, Kerma and D are equal.

## Slide 15: Relation between absorbed dose and exposure
2: Radiation units and dose quantities
<number>
It is possible to calculate the absorbed dose in a material if the exposure is known
D [Gy]. = f . X [C kg-1]
  f = conversion coefficient depending on medium
The absorbed energy in a quantity of air exposed to 1 [C kg-1] of X Rays is 0.869 [Gy]
  f(air) = 0.869

## Slide 16: Example of conversion coefficient: f
2: Radiation units and dose quantities
<number>
 | f values ([Gy] / Ckg-1]) |  | 
Photon energy | Water | Bone | Muscle
10 keV | 0.91 | 3.5 | 0.93
100 keV | 0.95 | 1.5 | 0.95

## Slide 17: Part 2: Radiation units and dose quantities
Topic 3: Mean Absorbed Dose in a tissue
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 18: Mean absorbed dose in a tissue or organ
2: Radiation units and dose quantities
<number>
The mean absorbed dose in a tissue or organ DT is the energy deposited in the organ divided by the mass of that organ.

## Slide 19: Exposure and absorbed dose or KERMA
2: Radiation units and dose quantities
<number>
Exposure can be linked to air dose or kerma by suitable conversion coefficients.
For example, 100 kV X Rays that produce an exposure of 1 R at a point will also give an air kerma of about 8.7 mGy (0.87 rad) and a tissue kerma of about 9.5 mGy (0.95 rad) at that point.

## Slide 20: Ratio of absorbed dose in soft tissue to that in air
2: Radiation units and dose quantities
<number>
Values of absorbed dose to tissue will vary by a few percent depending on the exact composition of the medium that is taken to represent soft tissue.
The following value is usually used for 80 kV and 2.5 mm Al:
Dose in soft tissue = 1.06 Dose in air

## Slide 21: Part 2: Radiation units and dose quantities
Topic 4: Equivalent dose H
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 22: Equivalent dose: H
2: Radiation units and dose quantities
<number>
The equivalent dose H is the absorbed dose multiplied by a dimensionless radiation weighting factor, wR which expresses the biological effectiveness of a given type of radiation
To avoid confusion with the absorbed dose, the SI unit of equivalent dose is called the sievert (Sv). The old unit was the “rem”
1 Sv = 100 rem

## Slide 23: Radiation weighting factor, wR
2: Radiation units and dose quantities
<number>
For most of the radiation used in medicine (X Rays, , e-) wR is = 1, so the absorbed dose and the equivalent dose are numerically equal
The exceptions are:
  alpha particles (wR = 20)
  neutrons (wR = 5 - 20).

## Slide 24: Part 2: Radiation units and dose quantities
Topic 5: Effective Dose
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 25: Detriment
2: Radiation units and dose quantities
<number>
Radiation exposure of the different organs and tissues in the body results in different probabilities of harm and different severity
The combination of probability and severity of harm is called “detriment”.

## Slide 26: Tissue weighting factor
2: Radiation units and dose quantities
<number>
To reflect the combined detriment from stochastic effects due to the equivalent doses in all the organs and tissues of the body, the equivalent dose in each organ and tissue is multiplied by a tissue weighting factor, WT, and the results are summed over the whole body to give the effective dose E

## Slide 27: Tissue weighting factors, wT
2: Radiation units and dose quantities
<number>
Organ/Tissue | WT | Organ/Tissue | WT
Bone marrow | 0.12 | Lung | 0.12
Bladder | 0.04 | Liver | 0.04
Bone surface | 0.01 | Oesophagus | 0.04
Brain | 0.01 | Salivary Glands | 0.01
Breast | 0.12 | Skin | 0.01
Colon | 0.12 | Stomach | 0.12
Gonads | 0.08 | Thyroid | 0.04
Liver | 0.05 | Remainder | 0.12

## Slide 28: Effective dose, E
2: Radiation units and dose quantities
<number>
E = T wT.HT
E: effective dose
wT: weighting factor for organ or tissue T
HT: equivalent dose in organ or tissue T

## Slide 29: Part 2: Radiation units and dose quantities
Topic 6: Related dosimetry quantities (surface and depth dose, backscatter factor…..)
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 30: Entrance surface dose (ESD)
2: Radiation units and dose quantities
<number>
Absorbed dose is a property of the absorbing medium as well as the radiation field, and the exact composition of the medium should be clearly stated.
Usually ESD refers to soft tissue (muscle) or water
Absorbed dose in muscle is related to absorbed dose in air by the ratio of the mass energy coefficients

## Slide 31: Entrance surface dose (ESD)
2: Radiation units and dose quantities
<number>
1.06
air
en
water
en
























µ

µ
The obtained value for all typical diagnostic X Ray qualities can be assumed to be 1.06 (± 1%)
F =
where (µen/) are the mass energy coefficients of water and air, respectively.
The obtained value for all typical diagnostic X Ray qualities can be assumed to be 1.06 (± 1%)
F =
where (µen/) are the mass energy coefficients of water and air, respectively.

## Slide 32: Entrance surface dose (ESD)
2: Radiation units and dose quantities
<number>
On the other hand, the ESD measured on the surface of the patient or phantom includes a contribution from photons scattered back from deeper tissues, which is not present for free air measurements
For this reason, correction factor (backscatter factor) must be introduced
If measurements are made at other distances than the true focus-to-skin distance, doses must be corrected by the inverse square law

## Slide 33: Backscatter factors (water)
2: Radiation units and dose quantities
<number>
HVL | Field size (cm x cm) |  |  |  | 
mm Al | 10 x 10 | 15 x 15 | 20 x 20 | 25 x 25 | 30 x 30
2.0 | 1.26 | 1.28 | 1.29 | 1.30 | 1.30
2.5 | 1.28 | 1.31 | 1.32 | 1.33 | 1.34
3.0 | 1.30 | 1.33 | 1.35 | 1.36 | 1.37
4.0 | 1.32 | 1.37 | 1.39 | 1.40 | 1.41

## Slide 34: Dose area product (I)
2: Radiation units and dose quantities
<number>
The dose-area product (DAP) quantity is defined as the dose in air in a plane, integrated over the area of interest
The DAP (cGy·cm2) is constant with distance since the cross section of the beam is a quadratic function which cancels the inverse quadratic dependence on dose
This is true neglecting absorption and scattering of radiation in air and even for X Ray housing near the couch table

## Slide 35: Inverse square law
2: Radiation units and dose quantities
<number>

## Slide 36: DAP-meter (Diamentor ®)
2: Radiation units and dose quantities
<number>

## Slide 37: Dose-area product meter
2: Radiation units and dose quantities
<number>

## Slide 38: Dose area product (II)
2: Radiation units and dose quantities
<number>
It is always necessary to calibrate and to check the transmission chamber for the X Ray installation in use
In some European countries, it is compulsory that new equipment is equipped with an integrated ionization transmission chamber or with automatic calculation methods
It is convenient, in this case, also to check the read-out as some systems overestimate the real DAP value

## Slide 39: Part 2: Radiation units and dose quantities
Topic 7: Specific dosimetry quantities (Mammography, CT,…)
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 40: The average glandular dose (AGD)
2: Radiation units and dose quantities
<number>
The Average Glandular Dose (AGD) is the dosimetry quantity generally recommended for risk assessment
The use of AGD is recommended by the ICRP, the British Institute of Physical Sciences in Medicine, the NCRP, the BSS and the Netherlands Commission on Radiation Dosimetry (NCS)

## Slide 41: The average glandular dose AGD (mammography)
2: Radiation units and dose quantities
<number>
The AGD cannot be measured directly but it is derived from measurements with the standard phantom for the actual technique set-up of the mammographic equipment
The Entrance Surface Air Kerma (ESAK) free-in-air (i.e., without backscatter) has become the most frequently used quantity for patient dosimetry in mammography
For other purposes (compliance with reference dose level) one may refer to ESD which includes backscatter

## Slide 42: The ESAK (mammography)
2: Radiation units and dose quantities
<number>
ESAK can be determined by:
  a TLD dosimeter calibrated in terms of air kerma free-in-air at a HVL as close as possible to 0.4 mm Al with a standard phantom
  a TLD dosimeter calibrated in terms of air kerma free-in-air at a HVL as close as possible to 0.4 mm Al stuck to the patient skin (appropriate backscatter factor should be applied to Entrance Surface Dose measured with the TLD to express ESAK)
  Note: due to low kV used the TLD is seen on the image
  a radiation dosimeter with a dynamic range covering at least 0.5 to 100 mGy (better than  10% accuracy)

## Slide 43: Dosimetric quantity for C.T.
2: Radiation units and dose quantities
<number>
CTDI (Computed Tomography Dose Index)
DLP (Dose-Length Product)
MSAD (Multiple Scan Average Dose)

## Slide 44: Computed tomography dose index (CTDI)
2: Radiation units and dose quantities
<number>
The CTDI is the integral along a line parallel to the axis of rotation (z) of the dose profile (D(z)) for a single slice, divided by the nominal slice thickness T
In practice, a convenient assessment of CTDI can be made using a pencil ionization chamber with an active length of 100 mm so as to provide a measurement of CTDI100 expressed in terms of absorbed dose to air (mGy).
D(z)dz
T
1
=
+
-
CTDI




## Slide 45: Computed tomography dose index (CTDI)
2: Radiation units and dose quantities
<number>
measurements of CTDI may be carried out free-in-air in parallel with the axis of rotation of the scanner (CTDI100, air)
or at the centre (CTDI100, c)
and 10 mm below the surface (CTDI100, p) of standard CT dosimetry phantoms
the subscript ‘n’ (nCTDI) is used to denote when these measurements have been normalised to unit mAs.

## Slide 46: Computed tomography dose index (CTDI)
2: Radiation units and dose quantities
<number>
On the assumption that dose in a particular phantom decreases linearly with radial position from the surface to the centre, then the normalised average dose to the slice is approximated by the (normalised) weighted CTDI: [mGy(mAs)-1]
where:
  C is the tube current x the exposure time (mAs)
  CTDI100,p represents an average of measurements at four different locations around the periphery of the phantom
On the assumption that dose in a particular phantom decreases linearly with radial position from the surface to the centre, then the normalised average dose to the slice is approximated by the (normalised) weighted CTDI: [mGy(mAs)-1]
where:
  C is the tube current x the exposure time (mAs)
  CTDI100,p represents an average of measurements at four different locations around the periphery of the phantom
)
(
CTDI
3
2
+
CTDI
3
1
C
1
=
CTDI
p
100,
c
100,
w
n

## Slide 47: Reference dose quantities
2: Radiation units and dose quantities
<number>
Two reference dose quantities are proposed for CT in order to promote the use of good technique:
  CTDIw in the standard head or body CT dosimetry phantom for a single slice in serial scanning or per rotation in helical scanning: [mGy]
where:
  nCTDIw is the normalised weighted CTDI in the head or body phantom for the settings of nominal slice thickness and applied potential used for an examination
  C is the tube current x the exposure time (mAs) for a single slice in serial scanning or per rotation in helical scanning.
C
CTDI
=
CTDI
w
n
w


## Slide 48: Reference dose quantities
2: Radiation units and dose quantities
<number>
DLP Dose-length product for a complete examination: [mGy • cm]
where:
  i represents each serial scan sequence forming part of an examination
  N is the number of slices, each of thickness T (cm) and radiographic exposure C (mAs), in a particular sequence.
  N.B.: Any variations in applied potential setting during the examination will require corresponding changes in the value of nCTDIw used.
C
N
T
CTDI
=
DLP
w
n
i





## Slide 49: Reference dose quantities
2: Radiation units and dose quantities
<number>
In the case of helical (spiral) scanning [mGy • cm]:
where, for each of i helical sequences forming part of an examination:
  T is the nominal irradiated slice thickness (cm)
  A is the tube current (mA)
  t is the total acquisition time (s) for the sequence.
N.B.: nCTDIw is determined for a single slice as in serial scanning.
t
A
T
CTDI
=
DLP
w
n
i





## Slide 50: Reference dose quantities
2: Radiation units and dose quantities
<number>
Multiple Scan Average Dose (MSAD): The average dose across the central slice from a series of N slices (each of thickness T) when there is a constant increment between successive slices:
where:
DN,I(z) is the multiple scan dose profile along a line parallel to the axis of rotation (z).
(z)dz
D
=
MSAD
I
N,
2
I
+
2
I
-
1

I

## Slide 51: Summary
2: Radiation units and dose quantities
<number>
Dosimetric quantities are useful to know the potential hazard from radiation and to determine radiation protection measures to be taken.
The old, non-S.I. quantities and units are mentioned, since these are still used in some countries, notably the United States of America.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Let’s summarize the main subjects we did cover in this session. (List the main subjects covered and stress again the important features of the session)

## Slide 52: Where to Get More Information
2: Radiation units and dose quantities
<number>
Gregg EC. Effects of ionizing radiation on humans. In Waggener RG and Kereikas JG., editors. Handbook of medical physics, Volume II. Boca Raton, CRC Press Inc., 1984.
Radiation Dosimetry. Volume 1. Ed: Attix F.H. and Roesch W.C. New York, Academic Press, 1968.
Radiation exposure in Computed Tomography; 4th revised Edition, December 2002, H.D.Nagel, CTB Publications, D-21073 Hamburg
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 53: References
2: Radiation units and dose quantities
<number>
Protection against ionizing radiation from external sources used in medicine. ICRP Publication 33. Pergamon Press 1982.
Radiological protection and safety in medicine. ICRP Publication 73. Pergamon 1996.
Quality Criteria for Computed Tomography. EUR 16262. Office for Official Publications of the European Communities. Luxembourg 1999

## Slide 54: References
2: Radiation units and dose quantities
<number>
Avoidance of radiation injuries from medical interventional procedures. ICRP Publication 85. Ann ICRP 2000;30 (2). Pergamon.
Quantities and Units in Radiation Protection Dosimetry. ICRU report 51. Bethesda, USA, 1993.