## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L 20: Optimization of Protection in Digital Radiography
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: Topics
20: Digital Radiology
<number>
Introduction
Basic concepts
Relation between diagnostic information and patient dose
Quality Control
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Explanation or/and additional information
Instructions for the lecturer/trainer

## Slide 3: Overview
20: Digital Radiology
<number>
To become familiar with the digital imaging techniques in projection radiography and fluoroscopy, to understand the basis of the DICOM standard, and the influence of the digital radiology on image quality and patient doses
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 4: Part 20: Digital Radiography
Topic 1: Introduction
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 5: Transition from conventional to digital radiography
20: Digital Radiology
<number>
Many conventional fluoroscopic and radiographic systems have been replaced with digital systems
Digital radiography has become a challenge which may have advantages as well as disadvantages
Changing from conventional to digital radiography requires additional training

## Slide 6: Transition from conventional to digital radiography
20: Digital Radiology
<number>
Digital images can be digitally processed This is not possible in conventional radiology!!.
Digital images can be easily transmitted through networks, archived, and rapidly retrieved
Attention should be paid to the potential increase of patient doses due to tendency of:
  producing more images than needed
  producing higher image quality not necessarily required for the clinical purpose (referred to as “dose creep”)

## Slide 7: Radiation dose in digital radiography
20: Digital Radiology
<number>
Conventional films quickly indicate if an incorrect radiographic technique is used: images are too white or too black
Digital technology provides user with a “good image” since its dynamic range and digital image processing compensates for incorrect techniques even if the dose is higher than necessary

## Slide 8: What is “dynamic range”?
20: Digital Radiology
<number>
Wide dose range to the detector, allows a “reasonable” image quality to be obtained
Flat panel detectors (discussed later) have a dynamic range of 104 (from 1 to 10,000) while a screen-film system has a range of approximately 101.5

## Slide 9: Characteristic curve of CR system
20: Digital Radiology
<number>
HR-III
CEA Film-Fuji Mammofine
CR response
Air Kerma (mGy)
0.001 0.01 0.1 1
3.5
3
2.5
2
1.5
1
0.5
0
Density

## Slide 10: Intrinsic digital techniques
20: Digital Radiology
<number>
Digital radiography and digital fluoroscopy are new imaging techniques, which replace film-based image acquisition

## Slide 11: Digitizing conventional films
20: Digital Radiology
<number>
Conventional radiographic images can be converted into digital information by a “digitizer”, and electronically stored
Such a conversion also allows some numerical post-processing
Such a technique cannot be considered as a “ digital radiography” technique.

## Slide 12: Part 20: Digital Radiography
Topic 2: Basic concepts
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 13: Analogue versus digital
20: Digital Radiology
<number>
Digital: A given output can only have discrete values
Analogue: A given output can have continuous values

## Slide 14: What is digital radiography?
20: Digital Radiology
<number>
In conventional radiographic images, spatial position and film density are analogue values
Digital radiography uses a matrix to represent an image
A matrix is a square area divided into rows and columns. The smallest element of a matrix is called ”pixel”
Each pixel of the matrix is used to store the individual gray levels of an image, which are represented by positive integer numbers
The location of each pixel in a matrix is encoded by its row and column number (x,y)

## Slide 15: Different number of pixels per image: original was 3732 x 3062 pixels x 256 grey levels (21.8 Mbytes). Here, resized at 1024 x 840 (1.6 MB).
20: Digital Radiology
<number>

## Slide 16: Different number of pixels per image: original was 3732 x 3062 pixels x 256 grey levels (21.8 Mbytes). Here, resized at 128 x 105 (26.2 kB).
20: Digital Radiology
<number>

## Slide 17: Different number of pixels per image: original was 3732 x 3062 pixels x 256 grey levels (21.8 Mbytes). Here, resized at 64 x 53 (6.6 kB)
20: Digital Radiology
<number>

## Slide 18: The digital radiology department
20: Digital Radiology
<number>
In addition to the X-ray rooms and imaging systems, a digital radiology department has two other components:
  A Radiology Information management System (RIS) that can be a subset of the hospital information system (HIS)
  A Picture Archiving and Communication System (PACS).

## Slide 19: DICOM
20: Digital Radiology
<number>
DICOM (Digital Imaging and Communications in Medicine) is the industry standard for transfer of radiological images and other medical information between different systems
All medical imaging products should be in compliance with the DICOM standard
However, due to the rapid development of new technologies and methods, the compatibility and connectivity of systems from different vendors is still a challenge

## Slide 20: DICOM format images:
20: Digital Radiology
<number>
Radiology images in DICOM format contain, in addition to the image, a header with an important set of additional data related with:
  the X ray system used to obtain the image
  the identification of the patient
  the radiographic technique, dosimetric details, etc.

## Slide 21: Digital radiography process
20: Digital Radiology
<number>
Image acquisition
Image processing
Image display
  Importance of viewing conditions
Image archiving (PACS)
Image retrieval
  Importance of time allocated to retrieve images

## Slide 22
20: Digital Radiology
<number>
Radiotherapy
Department
Outline of a basic PACS system

## Slide 23: Image acquisition (I):
20: Digital Radiology
<number>
Photostimulable phosphor plates (PSP).
  Called CR (computed radiography)
  Can be used with conventional
  X-ray systems
  Direct digital radiography (flat panel detectors)
  Direct conversion (selenium)
  Indirect conversion (scintillation)

## Slide 24: Computed Radiography (CR)
20: Digital Radiology
<number>
CR utilises photostimulable phosphor luminescence
Image plates made of a suitable phosphor material are exposed to X-rays in the same way as a conventional screen-film combination
Unlike a normal screens which release light spontaneously upon exposure to X-rays, the CR image plate retains most of the absorbed X-ray energy, in energy traps, forming a latent image

## Slide 25: Computed Radiography (CR)
20: Digital Radiology
<number>
A scanning laser is then used to release the stored energy producing photo-stimulated luminescence
The emitted light, which is linearly proportional to the locally incident X-ray intensity over at least four decades of exposure range, is detected by a photo multiplier-analogue to digital converter system and converted to a digital image
The resultant images consist of 2,370 x 1,770 pixels (for mammograms) with 1,024 grey levels (10 bits) and a pixel size of 100 µm corresponding to a 24 x 18 cm field size

## Slide 26: The principle of PSP
20: Digital Radiology
<number>
Excitation
Storage
Emission
CB
Trap
ADC
PMT

## Slide 27
20: Digital Radiology
<number>
PSP digitizer
(Images courtesy of Agfa)
Cassette and PSP
Workstation

## Slide 28
20: Digital Radiology
<number>
(Images courtesy of GE Medical Systems)
Digital detector

## Slide 29: Image acquisition (II)
20: Digital Radiology
<number>
Other alternatives are:
  Selenium cylinder detector (introduced for chest radiography with a vertical mounted rotating cylinder coated with selenium)
  Charge Coupled Devices (CCD)
  The image of a luminescent screen is recorded with CCD devices and converted into digital images

## Slide 30: Digital fluoroscopy
20: Digital Radiology
<number>
Digital fluoroscopic systems are based on the use of image intensifiers (I.I.) or flat panel imaging technology
In conventional systems the output screen of the I.I. is projected by an optical lens onto a film. In digital systems the output screen is projected onto a video camera system or a CCD camera. With flat panel technology the image is converted directly to a digital signal.
The output signals of the camera are converted into a digital image matrix (1024 x 1024 pixel in most systems).
Typical digital functions are “last image hold”, “virtual collimation”, etc.

## Slide 31: Part 20: Digital Radiography
Topic 3: Relation between diagnostic information and patient dose
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 32: Image quality and dose
20: Digital Radiology
<number>
Diagnostic information content in digital radiography is generally higher than in conventional radiology if equivalent doses are used
The wider dynamic range of the digital detectors and the capabilities of post processing provide more information from the digital radiographic images

## Slide 33: Tendency to increase dose ?
20: Digital Radiology
<number>
Image quality metrics, e.g., noise, correlate well with dose in digital radiography
For digital detectors, higher doses result in a better image quality (less “noisy” images)
Increasing dose results in a increase in signal-to-noise ratio
Consequently, there is a tendency to increase dose (“dose creep”) in order to produce more aesthetically pleasing images which usually do not contain any additional diagnostic information

## Slide 34: Digital radiography versus screen-film
20: Digital Radiology
<number>
In digital radiography (DR) the “image density” is automatically adjusted by the image processing, no matter of the applied dose.
This is one of the key advantages of DR which helps to significantly reduce the retake rate, but at the same time may hide occasional or systematic under- or over-exposures.
Under-exposures are easily corrected by radiographers but may result in noisy images.
Over-exposures cannot be detected unless patient dose measurements are monitored

## Slide 35
20: Digital Radiology
<number>
Under-exposure results in a “too noisy” image
Over-exposure yields good images with unnecessary high dose to the patient
Over range of digitiser may result in uniformly black area with potential loss of information
Exposure level 2,98
Exposure level 2,36

## Slide 36: An under-exposed image is “too noisy”
20: Digital Radiology
<number>
Exposure level 1,15
Exposure level 1,87

## Slide 37: Exposure level
20: Digital Radiology
<number>
Some digital systems provide the user with an “exposure level” index which expresses the dose level received at the digital detector
The relation between dose and exposure time is usually logarithmic: doubling the dose to the detector, will increase the “exposure level” to a factor of 0.3 = log(2).

## Slide 38: Risk of increased doses
20: Digital Radiology
<number>
The wide dynamic range of digital detectors results in good image quality while using high dose technique at the entrance of the detector and at the entrance of the patient
This is not possible with conventional screen-film systems since high dose techniques always result in an image which is too dark

## Slide 39: Digital fluoroscopy
20: Digital Radiology
<number>
In digital fluoroscopy there is a direct link between diagnostic information (number of images and quality of the images) and patient dose
Digital fluoroscopy allows for producing a large number of images (since there is no need to introduce cassettes or film changers as in the analog systems).
As a consequence, dose to the patient may increase without any benefit

## Slide 40: Difficulty in auditing the number of images per procedure
20: Digital Radiology
<number>
Deleting unnecessary images is very easy in digital imaging
This makes auditing the dose to the patient difficult
The same applies to digital radiography in terms of auditing the number of retakes

## Slide 41: Actions that can influence image quality and patient doses in digital radiology (1)
20: Digital Radiology
<number>
Avoid bad viewing conditions (e.g., low monitor brightness or contrast, poor spatial resolution, high ambient illuminance levels etc.)
Provide training on the workstation capabilities (window-level, inversion, magnification, etc.).

## Slide 42: Actions that can influence image quality and patient doses in digital radiology (2)
20: Digital Radiology
<number>
Eliminate post-processing problems, digitizer problems, local hard disk, fault in electrical power supply, network problems during image archiving, etc.
Avoid loss of images in the network or in the PACS due to improper image identification
Reduce artifacts due to incorrect digital post-processing (creation of false lesions or pathology)

## Slide 43: Actions that can influence image quality and patient doses in digital radiology (3)
20: Digital Radiology
<number>
Promote easy access to the PACS in order to retrieve previous images to avoid repeated examinations.
Display dose indication at the console of the X ray system.
Availability of a workstation for post-processing (also for radiographers) additional to hard copy to avoid some retakes.

## Slide 44: Influence of the different image compression levels
20: Digital Radiology
<number>
Image compression can:
  influence the image quality of stored images in the PACS
  modify the time necessary to have the images available (transmission speed on the intranet)
High levels of image compression may result in a loss of image quality and in a possible repetition of the examination (extra radiation dose to the patients)

## Slide 45: Digital radiography: initial pitfalls (1)
20: Digital Radiology
<number>
Lack of training (and people reluctant to use computers)
Lack of knowledge of the viewing possibilities on the monitors (and post-processing capabilities).
Changes in radiographic techniques or geometric parameters while disregarding patient doses (image quality is usually sufficient with the post-processing).

## Slide 46: Digital radiography: initial pitfalls (2)
20: Digital Radiology
<number>
Lack of a preliminary image visualization on the monitors (by the radiologist) may result in a loss of diagnostic information (inappropriate window and level selection made by the radiographer)
The quality of the digital image has to be adequately determined, in particular when re-processing is not available

## Slide 47: Part 20: Digital Radiography
Topic 4: Quality Control
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 48: Important aspects to be considered for the QC programs in digital radiography (1)
20: Digital Radiology
<number>
Availability of requirements for different digital systems (CR, digital fluoroscopy, etc).
Availability of procedures avoiding loss of images due to network problems or electric power supply
Information confidentiality
Compromise between image quality and compression level in the images
Recommended minimum time to archive the images

## Slide 49: Important aspects to be considered for the QC programs in digital radiography (2)
20: Digital Radiology
<number>
Measurement of patient dose and record keeping
Diagnostic reference levels
Inadvertent deletion of images (or full series in fluoroscopy systems)
Auditing patient doses

## Slide 50: Display of patient dose (1)
20: Digital Radiology
<number>
Imaging physicians should be aware of patient doses and monitor the dose display at the control panel (or inside the X-ray room, for interventional procedures)
Some digital systems offer a color code or a bar in the pre-visualization monitor. This code or bar indicates whether the dose received by the detector is in the normal range (green or blue) or too high (red).

## Slide 51
20: Digital Radiology
<number>
Example of bar a display indicating the level of exposure at the digital detector

## Slide 52: Display of patient dose
20: Digital Radiology
<number>
The radiographic and dose data from the DICOM header can be used to auditing patient doses
If radiographic techniques (kV, mA, time, distances, filters, field size, etc.) and dose data (entrance dose, dose area product, etc.) are part of the DICOM header, retrospective analysis of patient doses can be performed and assessed against the image quality. (Optimization)

## Slide 53: Diagnostic Reference Levels (DRLs)
20: Digital Radiology
<number>
In digital radiography, the evaluation of patient doses should be performed more frequently than in conventional radiology:
  Easy improvement of image quality
  Unknown use of high dose technique
Doses should be evaluated compared to DRLs when new digital equipment or techniques are introduced

## Slide 54: Initial basic quality control
20: Digital Radiology
<number>
Initial approach
  obtain images of a test object under different radiographic conditions and measure the corresponding doses
  decide the best compromise considering both image quality and patient dose

## Slide 55: Optimization
20: Digital Radiology
<number>
TOR(CDR) plus ANSI phantom to simulate chest and abdomen examinations and to evaluate image quality

## Slide 56: Optimization technique for Abdomen AP
20: Digital Radiology
<number>
Simulation with TOR(CDR) + ANSI phantom
81 kVp, 100 cm (focus-film distance)
1.6 mGy

## Slide 57: Optimisation technique for Chest PA
20: Digital Radiology
<number>
Simulation with TOR(CDR) + ANSI phantom
125 kVp, 180 cm (focus-film distance)
* Grid focalised at 130 cm
0.25 mGy

## Slide 58: Image quality comparison
20: Digital Radiology
<number>

## Slide 59: Routine QC programme
20: Digital Radiology
<number>
For conventional or digital radiography
  Patient dose evaluation (when optimised)
  Tube-generator controls
  Image receptors (screen-film, viewing...)
  Film processors
For digital radiography
  Image quality evaluation with test object
  Image quality evaluation with clinical criteria
  Image processing

## Slide 60: QC equipment
20: Digital Radiology
<number>
  TOR(CDR) image quality test
  Photometer
  Sensitometer and densitometer
  Dosimeter
  CR image quality test object
  SMPTE image test pattern

## Slide 61: Summary
20: Digital Radiology
<number>
Digital radiography requires training to benefit from the advantages of this technology.
Image quality and diagnostic information are closely related to patient dose.
The transmission, archiving and retrieval of images can influence the workflow
A quality control program is essential in digital radiography due to risk of increasing patient doses
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Let’s summarize the main subjects we did cover in this session. (List the main subjects covered and stress again the important features of the session)

## Slide 62: Where to Get More Information (1)
20: Digital Radiology
<number>
Balter S. Interventional fluoroscopy. Physics, technology and safety. Wiley-Liss, New York, 2001.
Managing Patient Dose in Digital Radiology, ICRP Publication 93, Ann. ICRP 34 (1), 2004, Elsevier
Vano E, Fernandez JM, Ten JI, Prieto C, et al. Transition from Screen-Film to Digital Radiography: Evolution of Patient Radiation Doses at Projection Radiography. Radiology 243:461-466
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 63: Where to Get More Information (2)
20: Digital Radiology
<number>
http://www.gemedicalsystems.com/rad/xr/education/dig_xray_intro.html (last access 22 August 2002).
http://www.agfa.com/healthcare/ (last access 22 August 2002).
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources