## Slide 1: RADIATION PROTECTION IN DIAGNOSTIC AND INTERVENTIONAL RADIOLOGY
L 9: Medical Exposure - BSS (Potential exposure and investigation of accidental medical exposures)
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 2: TOPICS
9: Medical Exposure - BSS
<number>
Institutions Involved
Regulatory aspects - General
Medical Exposure Responsibilities
Radiation Protection Requirements - Justification
Optimization of Protection for Medical Exposure
Diagnostic Reference Levels
Investigation of Accidental Medical Exposures
Today:
The medical exposure -BSS
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
The third part of this lectures is similar to what is teached in “Optimnization in Interventional Radiology”
The lecturers have to agree upon WHO is covering this topic extensively, to avoid overlap.

## Slide 3: Overview
9: Medical Exposure - BSS
<number>
To become familiar with the BSS requirement for medical exposures and investigation of unnecessary exposure.
Case studies reports and lessons learned.
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Lecture notes: ( about 100 words)
Instructions for the lecturer/trainer

## Slide 4: Part 9: Medical exposure - BSS (Potential exposure and investigation of accidental medical exposures)
Topic 1: The BSS requirements
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 5: National Regulatory Authority
9: Medical Exposure - BSS
<number>
How to frame regulations?….. Let these be consistent with UK…..No..No…with USA. … Oh no..it should be S. Africa…no…Europe…
Is there some Harmonised set of regulation?
Yes, ……the standards set by BSS.

## Slide 6: Main International Institutions involved in Regulatory aspects
9: Medical Exposure - BSS
<number>
IAEA (The International Atomic Energy Agency)
IEC (International Electro-technical Commission)
ICRP ( The International Commission on Radiological Protection)
OECD/NEA (The Nuclear Energy Agency)
CEC (The Commission of European Communities)
WHO (The World Health Organization)
ILO (The International Labor Organization)
ISO (International Organization for Standardization)

## Slide 7: Regulatory aspects (I)
9: Medical Exposure - BSS
<number>
Since March 1960 the IAEA has been involved in the establishment of safety standards for the protection of health and life.
A that time the Board of Governors of the IAEA first approved radiation protection and safety measures stating that “The Agency’s basic safety standards ….. will be based, to the extent possible, on the recommendations of the ICRP”

## Slide 8: Regulatory aspects (II)
9: Medical Exposure - BSS
<number>
Several revised versions of “The Basic Safety Standards” (BSS) published: 1962, 1967, 1982, 1994, 1996, 2011
The last version “Safety series N° 115” (2011) reflects knowledge and developments in radiation protection and safety and related fields at that time
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Several revised versions of “The Basic Safety Standards” (BSS) have been published in the IAEA Safety series (1962, 1967, 1982, 1994)
The last version “Safety series N° 115” (1996) supersedes the previous BSS and reflects knowledge gained subsequently and developments in radiation protection and safety and related fields

## Slide 9: The Basic Safety Standards
9: Medical Exposure - BSS
<number>

## Slide 10: Regulatory aspects (III)
9: Medical Exposure - BSS
<number>
Standards:
  based primarily on the ICRP recommendations (ICRP 103, 2007)
  also take account of the principles recommended by the International Nuclear Safety Advisory Group (INSAG)
The quantities and units used:
  those recommended by the International Commission on Radiation Units and Measurements (ICRU)
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
The “Standards” are based primarily on the ICRP recommendations (ICRP 60) but also take account of the principles recommended by the International Nuclear Safety Advisory Group (INSAG) which has been elaborating nuclear safety concepts since 1985 (Safety Principles for Nuclear Power Plants)
The quantities and units used in the “Standards” are primarily those recommended by the International Commission on Radiation Units and Measurements (ICRU)

## Slide 11: Structure of the BSS
9: Medical Exposure - BSS
<number>
Principal requirements
  General requirements, planned exposures situations, emergency exposures situations, existing exposure situations
Schedules
  Exemption and clearance
  Categorization of sealed sources
  Dose limits for planned exposure situations
  Criteria for use in emergency preparedness and response

## Slide 12: Application Fields
9: Medical Exposure - BSS
<number>
The practices to which the “Standards” apply include:
  the production of sources and the use of radiation or radioactive substances for medical, industrial, veterinary or agricultural purposes, or for education, training or research
  the generation of nuclear power
  practices involving exposure to natural sources specified by the Regulatory Authority as requiring control
  any other practice specified by the Regulatory Authority

## Slide 13
9: Medical Exposure - BSS
<number>
Responsibilities
Main responsibilities:
registrants and licensees
employers
Subsidiary include:
suppliers
workers
radiation protection officers
medical practitioners & health professionals
qualified experts, ethical review committees

## Slide 14: Medical exposure responsibilities (I)
9: Medical Exposure - BSS
<number>
REGISTRANTS AND LICENSEES SHALL ENSURE THAT:
  No patient be administrated a diagnostic or therapeutic medical exposure unless the exposure is prescribed by a medical practitioner
  Medical practitioners be assigned the primary task and obligation of ensuring overall patient protection and safety in the prescription of, and during the delivery of, medical exposure

## Slide 15: Medical exposure responsibilities (II)
9: Medical Exposure - BSS
<number>
REGISTRANTS AND LICENSEES SHALL ENSURE THAT:
  Medical and paramedical personnel be available as needed, and either be health professionals or have appropriate training adequately to discharge assigned tasks
  For therapeutic uses of radiation, the calibration, dosimetry and quality assurance requirements of the Standards be conducted by or under the supervision of a qualified expert in radiotherapy physics

## Slide 16: Medical exposure responsibilities (III)
9: Medical Exposure - BSS
<number>
REGISTRANTS AND LICENSEES SHALL ENSURE THAT:
  The exposure of individuals incurred knowingly while voluntarily helping in the care, visit, support or comfort of patients undergoing medical diagnosis or treatment be constrained so that it is unlikely that her or his dose will exceed 5 mSv during the period of a patient’s diagnostic examination or treatment.
  Training criteria be specified or be subject to approval, as appropriate, by the Regulatory Authority in consultation with relevant professional bodies

## Slide 17: Radiation Protection Requirements - JUSTIFICATION (I)
9: Medical Exposure - BSS
<number>
GENERIC MATTER
  No practice or source within a practice should be authorized unless the practice produces sufficient benefit to the exposed individuals or to society to offset the radiation harm that it might cause i.e.: unless the practice is justified, taking into account social, economic and other relevant factors
MEDICAL EXPOSURE
  Medical exposure should be justified by weighing the diagnostic or therapeutic benefits they produce against the radiation detriment they might cause, taking into account the benefits and risk of available alternative techniques that not involve medical exposure

## Slide 18: Radiation Protection Requirements - JUSTIFICATION (II)
9: Medical Exposure - BSS
<number>
MEDICAL EXPOSURE
  In justifying each type of diagnostic examination by radiography, fluoroscopy or nuclear medicine, relevant guidelines will be taken into account, such as those established by the WHO
  Any radiological examination for occupational, legal, or health insurance purposes undertaken without reference to clinical indications is deemed to be not justified unless it is expected to provide useful information on the health of the individual examined or unless the specific type of examination is justified by those requesting it in consultation with relevant professional bodies

## Slide 19: Radiation Protection Requirements - JUSTIFICATION (III)
9: Medical Exposure - BSS
<number>
MEDICAL EXPOSURE
  Mass screening of population groups involving medical exposure is deemed to be not justified unless the expected advantages for the individual examined or for the population as a whole are sufficient to compensate for the economic and social costs, including radiation detriment

## Slide 20: Radiation Protection Requirements - JUSTIFICATION (IV)
9: Medical Exposure - BSS
<number>
MEDICAL EXPOSURE
  The exposure of humans for medical research is deemed to be not justified unless it is:
    in accordance with the provisions of the Helsinki Declaration and follows the guidelines for its application prepared by Council for International Organization of Medical Sciences (CIOMS) and WHO
    subject to the advice of an Ethical Review Committee and to applicable national and local regulations

## Slide 21: OPTIMIZATION - DESIGN CONSIDERATIONS (I)
9: Medical Exposure - BSS
<number>
  Equipment used in medical exposure should be so designed that:
    failure of a single component of the system be promptly detectable so that any unplanned medical exposure of patients is minimized
    the incidence of human error in the delivery of unplanned medical exposure be minimized

## Slide 22: OPTIMIZATION - DESIGN CONSIDERATIONS (II)
9: Medical Exposure - BSS
<number>
Registrants and licensees should:
    Take into account information provided by suppliers, identify possible equipment failures and human errors that could result in unplanned medical exposure
    Take all reasonable measures to prevent failures and errors (qualified personnel, calibration, quality assurance, training,…)

## Slide 23: OPTIMIZATION - DESIGN CONSIDERATIONS (III)
9: Medical Exposure - BSS
<number>
Registrants and licensees should:
    Take all reasonable measures to minimize the consequences of failures and errors
    Develop appropriate contingency plans for responding to events that may occur, display plans prominently, and periodically conduct practice drills

## Slide 24: OPTIMIZATION - DESIGN CONSIDERATIONS (IV)
9: Medical Exposure - BSS
<number>
With regard to equipment consisting of radiation generators, registrants and licensees should ensure that:
  Whether imported into or manufactured in the country where it is used, the equipment conform to applicable standards (IEC, ISO)
  Performance specifications and operating and maintenance instructions be provided in a major world language understandable to the users and in compliance with the relevant IEC and ISO standards
  Radiation beam control mechanisms be provided (devices indicating clearly and in a fail-safe manner whether the beam is “on” or “off”)

## Slide 25: OPTIMIZATION - DESIGN CONSIDERATIONS (V)
9: Medical Exposure - BSS
<number>
With regard to equipment consisting of radiation generators, registrants and licensees should ensure that:
  As nearly as practicable, the exposure be limited to the area being examined by using collimating devices aligned with the radiation beam
  The radiation field within the examination area without any radiation beam modifiers (wedges) be as uniform as practicable and the non uniformity be stated by the supplier
  Exposure rate outside the examination area due to radiation leakage or scattering be kept as low as reasonably achievable

## Slide 26: OPTIMIZATION OF PROTECTION FOR MEDICAL EXPOSURES - DESIGN CONSIDERATIONS (VI)
9: Medical Exposure - BSS
<number>
With regard to equipment consisting of radiation generators...:
  Radiation generators and their accessories be designed and manufactured so as to facilitate the keeping of medical exposures to the minimum necessary to obtain adequate diagnostic information
  Operational parameters (kVp, filtration, focal spot position, source-image receptor distance, field size, either tube current and time or their product) be clearly indicated

## Slide 27: OPTIMIZATION - DESIGN CONSIDERATIONS (VII)
9: Medical Exposure - BSS
<number>
With regard to equipment consisting of radiation generators...:
  Radiographic equipment be provided with devices that automatically terminate the irradiation after a preset time, current-time product or dose
  Fluoroscopic equipment be provided with a device that energizes the X Ray tube only when continuously depressed (such as a “dead-man’s switch” and equipped with indicators of the elapsed time and/or entrance dose monitors

## Slide 28: OPTIMIZATION - OPERATIONAL CONSIDERATIONS (I)
9: Medical Exposure - BSS
<number>
Registrants and licensees should ensure for diagnostic radiology that:
  The medical practitioners who prescribe or conduct radiological examinations:
    ensure that the appropriate equipment be used
    ensure that the exposure of patients be the minimum necessary to achieve the required diagnostic objective, taking into account norms of acceptable image quality
    take into account relevant information from previous examinations in order to avoid unnecessary additional examinations

## Slide 29: OPTIMIZATION - OPERATIONAL CONSIDERATIONS (II)
9: Medical Exposure - BSS
<number>
Registrants and licensees shall ensure ... that:
  The medical practitioner, the technologist or other imaging staff select the following parameters such that their combination produce the minimum patient exposure consistent with acceptable image quality and the clinical purpose of the examination
    the area to be examined, the number and size of views per examination and the fluoroscopy time
    the type of image receptor (e.g. high v.s. low speed screens)
    the use of anti-scatter grids

## Slide 30: OPTIMIZATION - OPERATIONAL CONSIDERATIONS (III)
9: Medical Exposure - BSS
<number>
    proper collimation of the primary X Ray beam to minimize the volume of patient tissue being irradiated and to improve image quality
    appropriate values of operational parameters (kVp, mA…)
    appropriate image storage techniques in dynamic imaging (number of images per second)
    adequate image processing factors (chemicals, developer temperature, …)

## Slide 31: OPTIMIZATION - OPERATIONAL CONSIDERATIONS (IV)
9: Medical Exposure - BSS
<number>
Registrants and licensees should ensure ……. that:
  Portable and mobile radiological equipment be used only for examinations where it is impractical or not medically acceptable to transfer patients to a stationary radiological installation
  Radiological examinations causing exposure of the abdomen or pelvis of women who are pregnant or likely to be pregnant be avoided unless there are strong clinical reasons for such examination
  Whenever feasible, shielding of radiosensitive organs such as gonads, lens of the eye and thyroid be provided as appropriate

## Slide 32: OPTIMIZATION - CLINICAL DOSIMETRY
9: Medical Exposure - BSS
<number>
Registrants and licensees should ensure that in radiological examinations, representative values for typical sized adult patients of entrance surface dose, dose-area products, dose rates and exposure time, or organ doses be determined and documented

## Slide 33: OPTIMIZATION - QUALITY ASSURANCE (I)
9: Medical Exposure - BSS
<number>
Registrants and licensees should establish a comprehensive QA program with the participation of appropriate qualified experts in radiation physics taking into account the principles established by the WHO and the PAHO

## Slide 34: OPTIMIZATION - QUALITY ASSURANCE (II)
9: Medical Exposure - BSS
<number>
Quality Assurance programs should include:
  measurements of the physical parameters of the radiation generators, imaging devices at the time of commissioning and periodically thereafter
  verification of the appropriate physical and clinical factors used in patient diagnosis or treatment
  written records of relevant procedures and results
  verification of the appropriate calibration and conditions of operation of dosimetry and monitoring equipment

## Slide 35: GUIDANCE LEVELS
9: Medical Exposure - BSS
<number>
Registrants and licensees should ensure that guidance levels be determined as specified in the Standards, revised as technology improves and used as guidance by medical practitioners, in order that:
  corrective action be taken as necessary if doses fall substantially below the guidance levels and the exposures do not provide useful diagnostic information and do not yield the expected medical benefit to patient
  reviews be considered if doses exceed the guidance levels as an input to ensuring optimized protection of patients and maintaining appropriate levels of good practices
  for diagnostic radiology, including CT and pediatric examinations, the guidance levels be derived from the data from wide scale quality surveys for the most frequent examinations

## Slide 36: ACCEPTABLE AND INTERVENTION (or investigation) LEVELS
9: Medical Exposure - BSS
<number>
+ intervention level
+ tolerated level
guidance level
- tolerated level
- intervention level
time
(Immediate action required)
(Investigation recommended)
test value

## Slide 37: DOSE CONSTRAINTS FOR MEDICAL EXPOSURE
9: Medical Exposure - BSS
<number>
For medical exposure dose constraints should only be used in optimizing the protection of persons exposed for medical research purposes, or of persons, other than workers, who assist in the care, support or comfort of exposed patients.

## Slide 38: Part 9: Medical exposure - BSS (Potential exposure and investigation of accidental medical exposures)
Topic 2: Investigation of accidental medical exposure
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 39: Investigation of exposure (B.S.S. 3.46)
9: Medical Exposure - BSS
<number>
Registrants and licensees shall promptly investigate:
any diagnostic exposure substantially greater than intended or resulting in doses repeatedly and substantially exceeding the established guidance levels
any equipment failure, accident error, mishap or other unusual occurrence with the potential for causing a patient exposure significantly different from that intended.

## Slide 40: Investigation of exposure (B.S.S. 3.180)
9: Medical Exposure - BSS
<number>
Registrants and licensees shall:
a) calculate or estimate the doses received and their distribution within the patient
b) indicate the corrective measures required to prevent recurrence of such an incident
c) implement all the corrective measures that are under their own responsibility

## Slide 41: Investigation of exposure (B.S.S. 3.180)
9: Medical Exposure - BSS
<number>
Registrants and licensees shall:
d) submit to the Regulatory Authority, as soon as possible after the investigation or as otherwise specified by the Regulatory Authority, a written report which states the cause of the incident and includes the information specified in (a) to (c), as relevant, and any other information required by the Regulatory Authority
e) inform the patient and his or her doctor about the incident.

## Slide 42: Potential exposure (I)
9: Medical Exposure - BSS
<number>
The Basic Safety Standards (BSS) define potential exposure as an exposure that is not expected to be delivered with certainty but that may result from an accident at a source or owing to an event or sequence of events of a probabilistic nature, including equipment failures and operating errors.

## Slide 43: Potential exposure (II)
9: Medical Exposure - BSS
<number>
Accidental and unintended exposures (are not certain to occur)
They should be considered:
  In the design of equipment
  In the procedures
  In the planning of their application
Their control requires consideration of the probability of occurrence of the exposure as well as the magnitude of the resulting dose

## Slide 44: Reduction of the probability and magnitude of accidental exposures
9: Medical Exposure - BSS
<number>
Registrants and licensees shall ensure the reduction of the probability and magnitude of accidental or unintended doses.
If a potential exposure is greater than any level specified by the Regulatory Authority, have a safety assessment made and submitted to the Regulatory Authority as part of the application.
Main emphasis in radiotherapy but attention should be paid to accidents in diagnostic procedures.

## Slide 45: Potential exposures in the different stages of the installation
9: Medical Exposure - BSS
<number>
To the extent reasonable and practicable, the estimation of the probabilities and the magnitudes of potential exposures should be taken into account in the safety assessments related to protection and safety measures for X Ray installations at different stages, including:
  siting
  design
  manufacture
  construction
  assembly
  commissioning
  operation
  maintenance
  and decommissioning, as appropriate
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 46: Potential exposures are relevant for the classification of areas
9: Medical Exposure - BSS
<number>
Registrants and licensees shall designate as a controlled area any area in which specific protective measures or safety provisions are or could be required for:
  controlling normal exposures or preventing the spread of contamination during normal working conditions, and
  preventing or limiting the extent of potential exposures
In determining the boundaries of any controlled area, registrants and licensees shall take account of the magnitude of the expected normal exposures, the likelihood and magnitude of potential exposures, and the nature and extent of the required protection and safety procedures

## Slide 47: Potential exposures are relevant for occupational dose monitoring
9: Medical Exposure - BSS
<number>
The nature, frequency and precision of individual monitoring shall be determined with consideration of the magnitude and possible fluctuations of exposure levels and the likelihood and magnitude of potential exposures
Potential exposures should also be taken into account for the monitoring of the workplaces

## Slide 48: Practical advice (I)
9: Medical Exposure - BSS
<number>
OBJECTIVE: To ensure the reduction of the probability and magnitude of accidental or unintended doses.
STRATEGY: Record incidents, investigate causes, propose corrective actions, train people, disseminate the lessons learned

## Slide 49: Practical advice (II)
9: Medical Exposure - BSS
<number>
Working instructions, written protocols, QA programs and criteria of acceptability are relevant for the reduction of potential exposures.
Safety during equipment design, safety during installation, full QA program (with QC, clinical protocols and continuous training).

## Slide 50: Part 9: Medical exposure - BSS (Potential exposure and investigation of accidental medical exposures)
Topic 3: Accidental medical exposure
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 51
9: Medical Exposure - BSS
<number>
Coronary angioplasty twice in a day followed by bypass graft because of complication. Dose  20 Gy (ICRP 85)
(a) 6-8 weeks after multiple coronary angiography and angioplasty procedures.
(b) 16-21 weeks
(c) 18-21 months after the procedures showing tissue necrosis .
(d) Close-up photograph of the lesion shown in (c).
(e) Photograph after skin grafting. (Photographs courtesy of T. Shope & ICRP).
(d)
(e)
(a)
(c)
(b)

## Slide 52
9: Medical Exposure - BSS
<number>
Photograph showing temporary epilation of the right occipital
region of the skull 5-6 weeks following embolization (Courtesy W. Huda).
Regrowth (greyer than original) reported after 3 months.
Neuroradiology Trans-arterial embolization of para orbital AVM twice at a gap of 3 days
Total dose  8 Gy

## Slide 53
9: Medical Exposure - BSS
<number>
Trans-jugular Intrahepatic Portosystemic Shunt - TIPS -
a) Sclerotic depigmented plaque with surrounding hyperpigmentation on the midback of a patient following three TIPS procedures.
These changes were present 2 years after the procedures and were described as typical of chronic radiodermatitis.
(Photograph from Nahass and Cornelius (1998).
b) Ulcerating plaque with a rectangular area of surrounding
hyperpigmentation on the midback
(a)
(b)

## Slide 54: Interventional radiology (I)
9: Medical Exposure - BSS
<number>
Deterministic effects (skin injuries) are relevant.
Several cases have been reported in many countries.
Skin injuries in cardiac ablation in very young patients (reported by E. Vañó et al.)

## Slide 55: Interventional radiology (II) (occupational injuries)
9: Medical Exposure - BSS
<number>
Deterministic effects in lens of the medical specialists can be relevant if inappropriate X Ray systems for interventional procedures are used.

## Slide 56: Part 9: Medical exposure - BSS (Potential exposure and investigation of accidental medical exposures)
Topic 4: Lessons learned and preventive actions
IAEA Training Material on Radiation Protection in Diagnostic and Interventional Radiology
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources
Part …: (Add part number and title)
Module…: (Add module number and title)
Lesson …: (Add session number and title)
Learning objectives: Upon completion of this lesson, the students will be able to:
…
. (Add a list of what the students are expected to learn or be able to do upon completion of the session)
Activity: (Add the method used for presenting or conducting the lesson – lecture, demonstration, exercise, laboratory exercise, case study, simulation, etc.)
Duration: (Add presentation time or duration of the session – hrs)
Materials and equipment needed: (List materials and equipment needed to conduct the session, if appropriate)
References: (List the references for the session)

## Slide 57: I.R.: recommendations for avoiding potential exposures (I)
9: Medical Exposure - BSS
<number>
Equipment intended for interventional procedures should be specially designed for this purpose and it shall be installed in adequate rooms.
Medical doctors accomplishing interventional procedures should be especially qualified for this purpose.
X Ray systems used for interventional procedures should incorporate a patient dose meter and the results shall be recorded.
Patients should be informed of the risks from interventional x-ray procedures

## Slide 58: I.R.: recommendations for avoiding potential exposures (II)
9: Medical Exposure - BSS
<number>
If an incident occurs, it should be reported to the Health Authority together with the corrective actions adopted.
A specific level of training in radiation protection (additional to the one required for the specialists in diagnostic radiology) is required for medical doctors performing interventional procedures.
Quality assurance programs should include the provision of record of incidents, together with investigation of the causes and corrective actions undertaken.

## Slide 59: Discussion (I)
9: Medical Exposure - BSS
<number>
In interventional radiology, skin injuries could be considered as resulting from normal exposures when these can be anticipated and accepted in advance, i.e., when the patient pathology may demand long fluoroscopy time or repeated interventions in a short time interval (as the one required in the case of restenosis), and the best choice is to accept them.

## Slide 60: Discussion (II)
9: Medical Exposure - BSS
<number>
However, practically all the cases of deterministic effects described to date in the scientific literature for interventional procedures have the characteristics of (prior to intervention) potential exposures; causes are: wrong operation of the equipment (higher dose rate than necessary), or procedure mistakes (excessive time of high contrast fluoroscopy, collimator too close to the patient skin, etc), themselves also deriving in part from an incomplete training in radiation protection of the specialists performing the procedures.

## Slide 61: Conclusions
9: Medical Exposure - BSS
<number>
Recording of incidents
Notification of incidents
Corrective measures
Specific radiation protection training
Patient dose measurement and its recording.

## Slide 62: Summary
9: Medical Exposure - BSS
<number>
In order to reduce potential exposures:
Take preventive actions
  design of equipment and of procedures
  Specific radiation protection training
Benefit from feedback
  Record and review incidents
  Implement corrective measures
Monitoring and record doses

## Slide 63: References (I)
9: Medical Exposure - BSS
<number>
Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards, Revision of IAEA Safety Series No. 115, IAEA, Vienna Austria, 2011
ICRP, Radiological protection and safety in medicine, ICRP publication 73 (Oxford, Pergamon Press), 1996
European guidelines on quality criteria for diagnostic radiographic images, EUR 16260 report, (Luxembourg, EC), 1996. http://www.sprmn.pt/legislacao/ficheiros/EuropeanGuidelineseur16260.pdf
European guidelines for QA in breast cancer screening and diagnosis. (4th Edition,2006) http://ec.europa.eu/health/ph_projects/2002/cancer/fp_cancer_2002_ext_guid_01.pdf

## Slide 64: References (II)
9: Medical Exposure - BSS
<number>
European guidelines on quality criteria for diagnostic radiographic images in pediatrics, EUR 16261 report, (Luxembourg, EC), 1996. http://www.e-radiography.net/regsetc/European_guide_children_extract.pdf
Quality criteria for computed tomography, EUR 16262 report, (Luxembourg, EC), 1997. http://w3.tue.nl/fileadmin/sbd/Documenten/Leergang/BSM/European_Guidelines_Quality_Criteria_Computed_Tomography_Eur_16252.pdf
Notes: Part No...., Module No....Lesson No
Module title
IAEA Post Graduate Educational Course in Radiation Protection and Safe Use of Radiation Sources

## Slide 65: References (III)
Efficacy and Radiation Safety in Interventional Radiology. WHO 2001. Geneva.
Avoidance of radiation injuries from medical interventional procedures. ICRP Publication 85. Ann ICRP 2000; 30 (2). Elsevier.
9: Medical Exposure - BSS
<number>