## Slide 1: Module 6.2: Source not under control (Mexico)
IAEA Training Course

## Slide 2
Ciudad Juárez
Ciudad Juárez, México: An accident with 60Co

## Slide 3: Beginning of scenario
Nov. 1977
  A teletherapy unit was purchased and imported – 60Co unit
  This was an illegal import
Nov. 1977 – Nov. 1983
  Never reported to the authorities
  The unit was stored in a warehouse for 6 years
Typical Co unit

## Slide 4: Maintenance staff’s role
6 Dec. 1983
  Some maintenance staff became interested – scrap value should be high
  He dismounted the source
  Perforated the source container on the truck
  Drove to a junk yard and sold it together with some other “valuable” metal pieces
A dismantled Co treatment head

## Slide 5: The source
Typical 60Co source displaying the interior with a large amount of pellets
15 TBq or 430 Ci

## Slide 6: At the junkyard
We have now about 6000 pellets of 60Co
  About a 1 mm in size
  On the truck
  In the junkyard – everywhere since metal scrap is moved around by cranes, etc.
  Mixed with all other metal scrap
  Other trucks moving scrap out of the junkyard
Main purchaser of scrap constructs reinforcing rods, e.g. for motor vehicles, buildings
The first truck broke down and was parked for 40 d in the village + another 10 d at a second location

## Slide 7: At Los Alamos
Another company making table bases got metal scrap from the junkyard
A truck load of tables passing the Los Alamos Nuclear Center triggered the radiation monitors
The highway was monitored and the truck was identified
Two days later it was determined where the activity came from

## Slide 8: Chronology in summary
6 Dec. 1983
  Treatment unit dismantled
14 Dec. 1983 - 16 Jan. 1984
  Dissemination of radioactive substance
16-18 Jan. 1984
  Detection of contamination and its origin
19-22 Jan. 1984
  Actions of investigation
23 Jan. - 8 Feb. 1984
  Corrective actions

## Slide 9: Initial activities after the contamination was detected
Recognition of places with possible contamination
  The plant in Chihuahua
  The scrap yard in Juárez
  Ciudad Juárez
  The customs in Juárez
Determination of possible sequence of events on the basis of production record and negotiation
Confinement of contaminated material
Measures of radiological safety for workers and public
Estimation of dose to workers

## Slide 10: Range of the contamination
30,000 table bases produced
6,600,000 kg of rods produced
Aerial survey of 470 km2 identified 27 Cobalt pellets
17,636 buildings were visited to determine if radioactive material was used in the construction
Too high levels in 814 buildings
  Partly or completely demolished
Reinforcement rods

## Slide 11: Extent of the accident
Approx. 4000 persons exposed
5 persons with doses from 3 to 7 Sv in 2 months
80 persons with dose greater than 250 mSv
18% of the exposed public received doses of 5-25 mSv
Storage of 37,000,000 kg of rods, metallic bases, material in process, scrap iron, barrels with pellets and contaminated material, earth, etc.

## Slide 12: Management of the accident
To stop the dissemination of the contamination
Decontaminate contaminated areas
To avoid additional exposure of the public and workers and to determine received doses
Collect and confiscate contaminated materials
Extensive efforts to locate additional focuses of contamination

## Slide 13: Causes and contributing factors
A person dismantled and insecurely stored a cobalt source and broke the capsule
Non-compliance with regulations
  The unit was illegally imported
  Stored under unsafe conditions
A staff member did not recognize the potentially dangerous situation
Radioactive parts were sold as scrap

## Slide 14: Lessons to learn
The existence of an emergency infrastructure facilitates the operations and limits the extension of an accident
The identification of a coordinator of the emergency is important
The existence of regulations is not sufficient to prevent violations
The responsibility for the fulfillment of each regulation must be clear and specific
The initial measures for an accident are critical
  They require special effort to adapt the plans to the prevalente reality

## Slide 15
Reference
MINISTERIO DE ENERGIA Y MINAS. COMISION NACIONAL DE SEGURIDAD NUCLEAR Y SALVAGUARDIAS. Accidente de contaminación con 60Co. CNSN-IT-001. Mexico (1984)